EUROMET Project 509:

A comparison of platinum-iridium kilogram mass standards

Stuart Davidson

National Physical Laboratory, UK

Final Report

Jan 2007

Abstract

The kilogram is unique among the base SI units in being the only one defined by an artefact rather than with relation to naturally occurring physical constants [1]. This imposes a unique set of problems relating to the maintenance and traceability of the mass scale worldwide. All mass measurements must be ultimately traceable to the International Prototype of the Kilogram (K) held at the BIPM (Bureau International des Poids et Mesures). This traceability is achieved via copies of the Prototype held by National Measurement Institutes (NMIs) and maintained by the periodic re-verification of these copies by the BIPM [2]. A major problem with maintaining this traceability is the inherent instability of the kilogram artefacts. The national copies of K accrete surface contamination over time which leads to an increase in their mass values. Without reference to K it is difficult to assess the magnitude of this mass gain and the traceability of measurements is compromised to some extent.

The purpose of this project was to compare the values of the national copies of the prototype of 18 European NMIs. The comparison was piloted by NPL (UK) who circulated two platinum-iridium transfer standards using a petal type comparison, each petal consisting of three or four participants. While the results obtained can be used to demonstrate equivalence between the participants, the main objective of the work was to calculate the relative mass gains of the national copies, to analyse this data, and to draw conclusions about the relative stabilities of the kilograms. Each participating NMI uses an algorithm to predict the mass gain of their national standard. The results gained from this comparison have been used to evaluate the accuracy of these algorithms and also to recommend a best fit algorithm for the modelling of mass gain of national standard kilograms.

1. 	Introduction

The kilogram is unique among the base SI units in being the only one defined by an artefact rather than with relation to naturally occurring physical constants. This imposes a unique set of problems relating to the maintenance and traceability of the mass scale worldwide. All mass measurements must be ultimately traceable to the International Prototype of the Kilogram (K) held at the BIPM (Bureau International des Poids et Mesures). This traceability is achieved via copies of the Prototype held by National Measurement Institutes (NMIs) and maintained by the periodic re-verification of these copies by the BIPM. A major problem with maintaining this traceability is the inherent instability of the kilogram artefacts. The national copies of K accrete surface contamination over time which leads to an increase in their mass values. Without reference to K it is difficult to assess the magnitude of this mass gain and the traceability of measurements is compromised to some extent.

The purpose of this project was to compare the values of the national copies of the prototype of 18 European NMIs. While the results obtained can be used to demonstrate equivalence between the participants, the main objective of the work is to calculate the relative mass gains of the national copies, to analyse this data, and to draw conclusions about the relative stabilities of the kilograms. Each participating NMI uses an algorithm to predict the mass gain of their national standard. The results gained from this comparison have been used to evaluate the accuracy of these algorithms.

2.	Participants

The participating laboratories are listed in Table 1.

| Laboratory |  | Country |
| --- | --- | --- |
| Bundesamt fur Eich- und Vermessungswesen | BEV | Austria |
| Metrology Division, FPS Economy | SMD | Belgium |
| Czech Metrological Institute | CMI | Czech Republic |
| Danish Institute of Fundamental Metrology | DFM | Denmark |
| Mittatekniikan keskus | MIKES | Finland |
| Institut National de Métrologie | BNM-INM | France |
| Physikalisch-Technische Bundesanstalt | PTB | Germany |
| National Office of Measures | OMH | Hungary |
| Istituto di Metrologia "G Colonnetti" | IMGC | Italy |
| Nederlands Meetinstituut-Van Swinden Laboratorium | NMi VSL | Netherlands |
| Justervesenet | JV | Norway |
| Central Office of Measures | GUM | Poland |
| Slovak Institute of Metrology | SMU | Slovakia |
| Centro Español de Metrologia | CEM | Spain |
| SP Measurement Technology | SP | Sweden |
| Swiss Federal Office of Metrology and Accreditation | METAS | Switzerland |
| Tübitak Ulusal Metroloji Enstitüsü | UME | Turkey |
| National Physical Laboratory | NPL | United Kingdom |

Table 1: List of Participating Laboratories

The National Physical Laboratory acted as the pilot for this exercise. Additionally, the BIPM provided traceability for the reference and travelling standards both before and after the comparison.

3.	Mass standards

Two platinum-iridium kilograms were used as the travelling mass standards for this comparison; these had the designations 55 and 651. Kilogram 55 is an official copy of the International Prototype and was manufactured in 1950, calibrated as part of the second verification in 1953 allocated to the PTB in 1954. Kilogram 651 is an underweight copy produced in 1979 and calibrated and supplied to the NPL in 1982. NPL platinum-iridium kilogram A was used as a reference standard against which to monitor the stability of the travelling standards.

4.	Comparison Protocol

The protocol followed was similar to that which has been adopted for CCM regional key comparisons in the EUROMET area. The two transfer standards were circulated among the participating laboratories, returning to NPL for stability checks after each series of measurements. Each measurement series comprised either three or four laboratories. Participants had one month in which to complete their measurements.

While being transported the transfer standards were housed in a sealed enclosure vented to the air through a micropore filter. Figure 1 shows a sealed kilogram enclosure similar to those used to transport the masses. The mass standards were hand carried between laboratories.

## Figure 1: Kilogram enclosure

5.	Preparatory measurements

5.1	Stability

The stability of the transfer standards 55 and 651, and of the stability standard A, had been monitored at their respective laboratories since their allocation by the BIPM.  Figures 2 to 4 show the stabilities of the three platinum-iridium mass standards prior to the start of the comparison in November 2000.

Figure 2: Stability of kilogram 55 from September 1991 to project start

Figure 3: Stability of kilogram 651 from April 1995 to project start

Figure 4: Stability of kilogram A from September 1991 to project start

Kilogram 55 shows a mass gain of approximately 2.5 micrograms per year, which is typical of a platinum-iridium kilogram. Unusually, kilogram 651 shows a loss in mass equal to about 4.5 micrograms a year. Although this behaviour is unusual, and no obvious explanation can be given for the weight loss, kilogram 651 has shown this behaviour over a number of years. While not ideal, the fact that the mass change in 651 is predictable means it was suitable for use in this comparison. Kilogram A can be seen to be extremely stable over the period prior to the start of the comparison, making it suitable as a reference standard against which to monitor changes in the transfer standards.

5.2	Calibration of standards

Prior to the start of the comparison both of the transfer standards and the stability standard were calibrated by the BIPM. The weights were returned to NPL, compared with each other and then transported to the first participating laboratory.

Between each series of three or four measurements by the participating laboratories the two travelling standards were returned to NPL where their stability was checked against the reference standard.

After the completion of the comparison the travelling standards were checked against the reference standard at NPL and all three weights were again calibrated by the BIPM.

6.	Results.

The results of the participating laboratories, together with those of the pilot laboratory (NPL) and of the BIPM are given in Table 2.

|  | Date | Lab | Country | Measured mass value (1 kg +) / µg |
| --- | --- | --- | --- | --- |
|  |  |  |  | 651 | u (k=1) | 55(PTB) | u (k=1) |
| 2000 | Dec | BIPM | BIPM | -11395.0 | 5.0 | 269.0 | 5.0 |
| 2001 | Jan | NPL | UK | -11392.6 | 6.0 | 270.1 | 6.0 |
|  | Feb | SMD | Belgium | -11378.0 | 11.0 | 275.0 | 11.0 |
|  | Mar | NMi | Netherlands | -11379.0 | 7.0 | 279.0 | 7.0 |
|  | Apr | CEM | Spain | -11400.0 | 5.0 | 223.0 | 5.0 |
|  | May | NPL | UK | -11399.4 | 6.0 | 234.2 | 6.0 |
|  | June | GUM | Poland | -11391.0 | 3.1 | 240.0 | 3.1 |
|  | July | CMI | Czech Rep. | -11418.0 | 50.0 | 234.0 | 50.0 |
|  | Aug | PTB | Germany | -11404.6 | 5.0 | 234.0 | 5.0 |
|  | Sept | SMU | Slovakia | -11433.0 | 5.0 | 216.0 | 5.0 |
|  | Oct | NPL | UK | -11410.0 | 6.0 | 237.2 | 6.0 |
|  | Nov | OHM | Hungary | -11397.0 | 8.0 | 237.0 | 8.0 |
|  | Dec | BEV | Austria | -11411.0 | 11.0 | 237.0 | 11.0 |
| 2002 | Feb | IMGC | Italy | -11418.9 | 3.7 | 229.5 | 3.7 |
|  | Mar | metas | Switzerland | -11400.4 | 12.6 | 248.6 | 12.6 |
|  | Apr | NPL | UK | -11416.2 | 6.0 | 233.3 | 6.0 |
|  | May | JV | Norway | -11417.0 | 17.0 | 229.0 | 17.0 |
|  | June | DFM | Denmark | -11421.5 | 5.4 | 229.9 | 5.3 |
|  | July | SP | Sweden | -11418.0 | 7.0 | 245.0 | 7.0 |
|  | Aug | MIKES | Finland | -11417.0 | 13.0 | 236.0 | 13.0 |
|  | Sept | NPL | UK | -11422.8 | 6.0 | 234.7 | 6.0 |
|  | Oct | BMN-INM | France | -11404.5 | 7.2 | 229.0 | 7.2 |
|  | Nov | UME | Turkey | -11400.0 | 5.5 | 241.0 | 5.5 |
|  | Dec | NPL | UK | -11425.8 | 6.0 | 236.2 | 6.0 |
| 2003 | Jan | BIPM | BIPM | -11424.0 | 5.0 | 238.0 | 5.0 |

# Table 2: Comparison results reported by the participants

Results show reasonable consistency between participants except for a step change in the value of kilogram 55 between March and April 2001. On examination of the weight on return to NPL it was clear that the weight had been damaged.

7.	Analysis

7.1	Reference values

The results from the participants, from the NPL stability checks and from the BIPM calibration values were analysed. Three potential reference values were considered;

- Using only the BIPM data

- Using the BIPM and NPL data

- Using all consistent data

Consistent data was regarded as that data having a normalised deviation from the reference value of less than 2.0 (calculated using standard uncertainties).

The step change noted in the value of kilogram 55 was taken into account when calculating the reference vale for this weight.

Since the NPL reference standard A and the travelling standards were calibrated at the BIPM before and after the comparison there is a strong correlation between the stability measurements made at NPL and the calibration of the weights at the BIPM. For this reason a correlation coefficient of r=0.9 between the NPL and the BIPM measurements has been assumed. Additionally, since each participating laboratory is directly traceable to the BIPM there can be assumed to be some correlation between the measured values and the BIPM value (and between individual measured values). However, this correlation will depend on a number of factors, such as when the national standard of each participant was calibrated. For this reason no correlation coefficient for the participants’ data was included.

The results of these analysis methods in calculating reference values for the two travelling standards are shown in figures 5 and 6.

Figure 5: Potential reference values for kilogram 55

Figure 6: Potential reference values for kilogram 651

The three analysis methods give very consistent reference values both in terms of the absolute value and in terms of any calculated drift, the agreement being within the smallest uncertainty calculated. Apart from the step change in kilogram 55 previously discussed the value of the weight was stable during the period of the comparison. As expected the value of kilogram 651 showed some drift during the comparison. The drift calculated by the three analysis methods is almost identical and equal to about 1 microgram per month mass loss. Interestingly this is much higher than the drift previously measured in 651 (4.5 micrograms per year) but appears to have been consistent throughout the period of the comparison. It can be seen that in terms of the uncertainties in the reference values, as the number of results included increases the uncertainty in the value becomes less. Since there is little significant difference between the reference values, the values calculated from just the BIPM measurements were used. These reference values are shown in Figures 7 and 8.

Figure 7: Reference values for kilogram 55 based on BIPM calibration data

Figure 8: Reference values for kilogram 651 based on BIPM calibration data

7.2	Participant’s data

The participants’ results and uncertainties (k=1) with relation to the reference values are shown in figures 9 and 10.  Values have been “excluded” where the result had a normalised deviation from the reference value of greater than 2.0 (calculated using standard uncertainties).

Figure 9: Participants’ results relative to the reference value for kilogram 55

Figure 10: Participants’ results relative to the reference value for kilogram 651

It can be seen that the majority of the results are in agreement with the reference value (when the results are expressed at k=2 uncertainty). More of the measurements of kilogram 651 disagree with the reference value. The most obvious explanations for this are either that the relatively large discrepancy of the weight from nominal (over 11 mg) made it difficult for some laboratories to calibrate or that the model for the drift is in error. Figure 6 shows the reference value for 651 (including a linear drift coefficient) calculated in three ways. The fact that the reference values calculated by all three approaches (including one using all the measured values) agree so well suggests that the drift model is valid (providing a linear drift can be assumed). It is, however, interesting to note that the last three participant measurements, including one at NPL, show a marked discrepancy from the calculated reference value but relatively good agreement among themselves.  Table 3 summarises the results shown in the two graphs above.

| Participant | Cal. date | Kilogram 55 | Kilogram 55 |
| --- | --- | --- | --- |
|  |  | Mass Value (µg) | U (k=1) (µg) | Ref. Mass (µg) | U (k=1) (µg) | Diff (µg) | Mass Value (µg) | U (k=1) (µg) | Ref. Mass (µg) | U (k=1) (µg) | Diff (µg) |
| BIPM | Nov-00 | 269 | 4 | 268.5 | 1.7 | 0.5 | -11 395 | 4 | -11395.5 | 1.5 | 0.5 |
| NPL | Jan-01 | 271 | 6 | 268.5 | 1.7 | 2.5 | -11 392 | 6 | -11397.2 | 1.5 | 4.9 |
| SMD | Feb-01 | 275 | 11 | 268.5 | 1.7 | 6.5 | -11 378 | 11 | -11398.5 | 1.4 | 20.5 |
| Nmi | Mar-01 | 279 | 7 | 268.5 | 1.7 | 10.5 | -11 379 | 7 | -11400.1 | 1.4 | 21.1 |
| CEM | Apr-01 | 223 | 5 | 236.2 | 1.4 | -13.2 | -11 400 | 5 | -11400.6 | 1.4 | 0.6 |
| NPL | May-01 | 230 | 6 | 236.2 | 1.4 | -6.2 | -11 397 | 6 | -11401.7 | 1.4 | 4.5 |
| GUM | Jun-01 | 240 | 3 | 236.2 | 1.4 | 3.3 | -11 391 | 3 | -11402.9 | 1.4 | 11.5 |
| CMI | Jul-01 | 234 | 50 | 236.2 | 1.4 | -2.2 | -11 418 | 50 | -11404.2 | 1.4 | -13.8 |
| PTB | Aug-01 | 234 | 5 | 236.2 | 1.4 | -2.2 | -11 405 | 5 | -11405.4 | 1.4 | 0.8 |
| SMU | Sep-01 | 216 | 5 | 236.2 | 1.4 | -20.2 | -11 433 | 5 | -11405.7 | 1.4 | -27.3 |
| NPL | Oct-01 | 234 | 6 | 236.2 | 1.4 | -2.7 | -11 409 | 6 | -11406.8 | 1.4 | -2.5 |
| OMH | Oct-01 | 237 | 8 | 236.2 | 1.4 | 0.8 | -11 397 | 8 | -11407.8 | 1.4 | 10.8 |
| BEV | Feb-02 | 237 | 11 | 236.2 | 1.4 | 0.8 | -11 411 | 11 | -11411.5 | 1.4 | 0.5 |
| IMGC | May-02 | 230 | 4 | 236.2 | 1.4 | -6.7 | -11 419 | 4 | -11414.6 | 1.4 | -4.3 |
| METAS | Apr-02 | 249 | 13 | 236.2 | 1.4 | 12.4 | -11 400 | 13 | -11413.2 | 1.4 | 12.8 |
| NPL | May-02 | 232 | 6 | 236.2 | 1.4 | -4.5 | -11 416 | 6 | -11415.2 | 1.4 | -0.8 |
| JV | Jul-02 | 229 | 17 | 236.2 | 1.4 | -7.2 | -11 417 | 17 | -11417.0 | 1.5 | 0.0 |
| DFM | Jun-02 | 230 | 5 | 236.2 | 1.4 | -6.3 | -11 422 | 5 | -11415.6 | 1.5 | -5.9 |
| SP | Aug-02 | 245 | 7 | 236.2 | 1.4 | 8.8 | -11 418 | 7 | -11418.5 | 1.5 | 0.5 |
| MIKES | Sep-02 | 236 | 13 | 236.2 | 1.4 | -0.2 | -11 417 | 13 | -11419.3 | 1.6 | 2.3 |
| NPL | Oct-02 | 234 | 6 | 236.2 | 1.4 | -2.6 | -11 427 | 6 | -11420.1 | 1.6 | -7.0 |
| BNM-INM | Jan-03 | 229 | 7 | 236.2 | 1.4 | -7.1 | -11 405 | 8 | -11423.3 | 1.7 | 18.8 |
| UME | Feb-03 | 241 | 6 | 236.2 | 1.4 | 4.8 | -11 400 | 6 | -11424.9 | 1.7 | 24.9 |
| NPL | Mar-03 | 246 | 6 | 236.2 | 1.4 | 9.4 | -11 408 | 6 | -11425.7 | 1.8 | 17.7 |
| BIPM | Apr-03 | 238 | 5 | 236.2 | 1.4 | 1.8 | -11 424 | 5 | -11426.9 | 1.8 | 2.9 |

Table 3: Participant results and uncertainties compared with calculated reference values

Looking at the measurements as a whole only one laboratory can be said to have results that are systematically different to the reference values by more than the combined uncertainty. Table 4 summarises the deviations of the individual measurements and, were provided by the participants, indicates the algorithm which the individual laboratories use to account for drift in the values of their national standard.

| Participant | National std. No. | Diff. from ref. | U (k=1) (µg) | Date of cal. | Algorithm |
| --- | --- | --- | --- | --- | --- |
|  |  | 55 (µg) | 651. (µg) |  |  |  |
| BIPM | K | 0.5 | 0.5 | 4.0 |  |  |
| NPL | 18 | 2.5 | 4.9 | 6.0 | Oct-91 | m = m(1991) + 0.356 097 x t 0.511 678 µg (t in days) |
| SMD | 37 | 6.5 | 20.5 | 11.0 | Feb-91 | m=m(1991) + 3.0 μg/year |
| NMi | 53 | 10.5 | 21.1 | 7.0 | Jan-97 | Data fit - approximates m=m(1997) + 1.2 μg/year |
| CEM | 24 | -13.2 | 0.6 | 5.0 | May-93 | Note 1 |
| NPL | 18 | -6.2 | 4.5 | 6.0 | Oct-91 | As above |
| GUM | 51 | 3.3 | 11.5 | 3.1 | Jun-90 | m=m(1990) + 2.2632 μg/year |
| CMI | 67 | -2.2 | -13.8 | 50.0 | Feb-99 | none |
| PTB | 52 | -2.2 | 0.8 | 5.0 | 1996 | m=m(1996) + 3.064 μg/year |
| SMU | 65 | -20.2 | -27.3 | 5.0 | 1991 | none |
| NPL | 18 | -2.7 | -2.5 | 6.0 | Oct-91 | As above |
| OMH | 16 | 0.8 | 10.8 | 8.0 | Jun-90 | m=m(1990) + 1 μg/year + 0.0368 μg/day |
| BEV | 49 | 0.8 | 0.5 | 11.0 | Jun-90 | m=m(1990) + 10 μg + (n - 1) x 1.4 μg |
| IMGC | 62 76 | -6.7 | -4.3 | 3.7 | Feb-93 Feb-99 | Compromise between estimated drift and consistency between two standards |
| METAS | 38 | 12.4 | 12.8 | 12.6 | Feb-91 | m=m(1991) + 0.037 μg/day (first 18 months) + 1.205µg/year (subsequent) |
| NPL | 18 | -4.5 | -0.8 | 6.0 | Oct-91 | As above |
| JV | 36 | -7.2 | 0.0 | 17.0 | Jan-01 | m=m(2001) + 10 μg + 1 μg/year |
| DFM | 48 | -6.3 | -5.9 | 5.4 | Jan-01 | Note 2 |
| SP | 40 | 8.8 | 0.5 | 7.0 | Feb-91 | m=m(1991) + 3.2 μg + (n - 0.25) x 1 μg |
| MIKES | 23 | -0.2 | 2.3 | 13.0 | Feb-01 | m=m(2001) + 6 μg (year 1) + 5 μg/year (subsequent) |
| NPL | 18 | -2.6 | -7.0 | 6.0 | Oct-91 | As above |
| BNM-INM | 35 | -7.1 | 18.8 | 7.8 | Apr-02 | none |
| UME | 54 | 4.8 | 24.9 | 5.5 | May-93 | m=m(1993) + 10 μg |
| NPL | 18 | 9.4 | 17.7 | 6.0 | Oct-91 | As above |
| BIPM | K | 1.8 | 2.9 | 5.0 |  |  |

## Table 4: Participant results and uncertainties and mass gain algorithms used

Note 1: CEM algorithm

m(t)	- calculated mass

m	- mass at last calibration

m1 	- mass loss due to cleaning (last calibration)

m2 	- mass diff. between last two calibrations (after cleaning)

m3 	- 10 µg (first year mass gain)

n	- years since last calibration

n1 	- Years between last two calibrations

Note 2: DFM algorithm

m(t)	- calculated mass at time t

m(t0)	- mass at last calibration

Y1 	- 1.91 x 10-6 mg/day

Y2 	- 1.13 x 10-4 (mg/day

t - t0	- days since last calibration

7.3	Analysis of algorithms

In order to check the accuracy of the algorithms used by the individual laboratories to predict the mass gain of their national prototype, the results of the comparison were analysed with respect to the time elapsed since the national prototype was last calibrated. Table 5 summarises the participants’ results (with respect to the reference value), corrected using the individual drift algorithms, in temporal order.

| National prototype cal. date | Days since calibration | Laboratory | Dev. from ref. value | Unc. (k=1) (µg) |
| --- | --- | --- | --- | --- |
|  |  |  | 55 (µg) | 651 (µg) |  |
| Apr-02 | 272 | BNM-INM | -16.1 | 8.9 | 9.2 |
| Jul-01 | 374 | DFM | 3.5 | 9.0 | 17.6 |
| Jan-01 | 509 | JV | 6.7 | 8.1 | 7.1 |
| Feb-01 | 585 | MIKES | 0.0 | 2.0 | 13.8 |
| Feb-99 | 881 | CMI | 15.0 | 4.0 | 50.2 |
| Feb-99 | 1142 | IMGC | -12.2 | -10.6 | 13.4 |
| Jan-97 | 1518 | NMi | -21.1 | -10.0 | 8.1 |
| Jun-96 | 1909 | PTB | 0.5 | 4.0 | 6.5 |
| May-93 | 2879 | CEM | -0.6 | 15.0 | 6.5 |
| Oct-91 | 3358 | NPL | 0.8 | 0.4 | 7.2 |
| May-93 | 3571 | UME | -22.1 | -3.0 | 7.4 |
| Feb-91 | 3637 | SMD | -20.5 | -6.0 | 11.7 |
| Jun-91 | 3745 | SMU | 28.6 | 22.0 | 6.5 |
| Jun-90 | 4017 | GUM | -10.4 | -1.5 | 5.2 |
| Feb-91 | 4092 | METAS | 7.5 | 8.5 | 5.8 |
| Jun-90 | 4153 | OMH | -9.4 | 1.0 | 9.1 |
| Feb-91 | 4203 | SP | 1.8 | -7.0 | 8.4 |
| Jun-90 | 4263 | BEV | 1.2 | 1.0 | 11.8 |

## Table 5: Time between national prototype calibration and participation in the comparison

## (results corrected by participants for drift in national prototypes)

The results of the participants’ measurements on the two transfer standards, with respect to the time since their national prototype was calibrated, are plotted in Figures 11 and 12.

Figure 11: Results for kilogram 55 with respect to time since national prototype calibration (values corrected by participants for drift in national prototype)

## Figure 12: Results for kilogram 651 with respect to time since national prototype calibration (values corrected by participants for drift in national prototype)

The results broadly confirm that the algorithms used by the participants model the mass gain of their national prototypes correctly. While there is a reasonably large spread on the drift compensated results, fitting both sets of data gives a drift co-efficient of virtually zero, indicating that there is negligible error on the collective algorithms. Linear best-fit lines have been added to the graphs and these show negligible departure from the expected result (along the y axis).

Examining the results uncorrected for the drift in the national prototypes will give a measure of the (collective) change in the values of the national prototypes since cleaning. Table 6 summarises the uncorrected results of the participants arrange with respect to the time of the last calibration of the national prototype.

| National prototype cal. date | Days since calibration | Laboratory | Dev. from ref. value | Unc. (k=1) (µg) |
| --- | --- | --- | --- | --- |
|  |  |  | 55 (µg) | 651 (µg) |  |
| Apr-02 | 272 | BNM-INM | 8.9 | -16.1 | 9.2 |
| Jul-01 | 374 | DFM | 21.0 | 15.5 | 17.6 |
| Jan-01 | 509 | JV | 9.0 | 7.6 | 7.1 |
| Feb-01 | 585 | MIKES | 12.0 | 10.0 | 13.8 |
| Feb-99 | 881 | CMI | 4.0 | 15.0 | 50.2 |
| Feb-99 | 1142 | IMGC | 20.4 | 18.8 | 13.4 |
| Jan-97 | 1518 | NMi | -6.0 | -16.2 | 8.1 |
| Jun-96 | 1909 | PTB | 20.1 | 16.6 | 6.5 |
| May-93 | 2879 | CEM | 35.2 | 20.5 | 6.5 |
| Nov-00 | 3358 | NPL | 23.1 | 23.5 | 7.2 |
| May-93 | 3571 | UME | 7.0 | -12.1 | 7.4 |
| Feb-91 | 3637 | SMD | 24.0 | 10.2 | 11.7 |
| Jun-91 | 3745 | SMU | 14.0 | 20.6 | 6.5 |
| Jun-90 | 4017 | GUM | 28.3 | 19.4 | 5.2 |
| Feb-91 | 4092 | METAS | 26.0 | 25.0 | 5.8 |
| Jun-90 | 4153 | OMH | 16.0 | 5.6 | 9.1 |
| Feb-91 | 4203 | SP | 7.4 | 16.2 | 8.4 |
| Jun-90 | 4263 | BEV | 26.0 | 26.2 | 11.8 |

## Table 6: Time between national prototype calibration and participation in the comparison

## (results uncorrected for drift in national prototypes)

The results of the uncorrected participants’ measurements on the two transfer standards, with respect to the time since their national prototype was calibrated, are plotted in Figures 13 and 14.

Two forms of best fit lines have been added to the graphs. The first (green line) is of the form;

m(t) = m(t0) + A + B(t)……………………………..…………….(1)

where A is the initial rapid mass gain of the national standard and B is a constant representing the subsequent linear mass gain as a function of time t. This is the form of model used by the majority of participants.

The second (red line) is of the form;

m(t) = m(t0) + A(t)B……………………………………………….(2)

Where A and B are constants and t is the time since the calibration of the national prototype.

Figure 13: Results for kilogram 55 with respect to time since national prototype calibration (values uncorrected for drift in national prototype)

Figure 14: Results for kilogram 651 with respect to time since national prototype calibration (values uncorrected for drift in national prototype)

With the value of t in days the solutions for equations (1) and (2) are given in Table 7.  To allow comparison with the algorithms reported in Table 4 the solutions for equation 2 are also reported with the timescale in years. The root sum squares of the residual values (the differences between the modelled value and the measured value) have been calculated. These residual RSS values give an estimate of the goodness of fit of the algorithms shown.

| Transfer standard | Linear gain model (µg) | Residual RSS (µg) | Exponential gain model (µg) | Residual RSS (µg) | Linear gain model (years) |
| --- | --- | --- | --- | --- | --- |
| 55 | 7.9 + 0.003 09 t | 2.18 | 1.796 5 t 0.285 025 | 2.18 | +7.9 µg + 1.1 µg/year |
| 651 | 12.5 + 0.001 44 t | 3.12 | 1.205 1 t 0.338 054 | 3.03 | + 12.5 µg + 0.5 µg/year |

## Table 7: Best fit algorithms for mass gain of national prototype kilograms

Interestingly both forms of the model are significantly different for the two transfer standards.  Combining the data and fitting the two mass gain models gives the following solutions;

Mass gain = 3.567 6 t 0.206 563 (with t in days)………………….……….(3)

for the exponential model and

Mass gain = 8.4 µg + 1.1 µg/year……………………………………….(4)

for the linear model.

8.	Conclusions

Although the main aim of the comparison was not a demonstration of the participating laboratories ability to calibrate platinum-iridium kilogram mass standards there is generally good agreement between the results of the participants. Individual laboratories will be able to draw their own conclusions about their results.

Examining the algorithms used by the individual NMIs to correct for the mass gain of their national prototype kilograms shows that there is no bias to the data and that, taken as a whole, the various algorithms used represent good models of the mass gain of the respective national prototypes.  Again, individual laboratories can draw conclusions regarding their own algorithms. Looking at the data uncorrected for the changes in the national prototypes, the results can be modelled using the widely accepted algorithm of rapid mass gain, after cleaning and BIPM, over the first few months followed by a linear mass gain model. From the analysis of this data, the values calculated are 8.4 micrograms for the rapid mass gain after cleaning followed by 1.1 microgram per year linear mass gain.

Acknowledgements

Thanks to Richard Davis and the Mass Section at the BIPM for the initial and final calibrations of the transfer standards and to Lars Nielsen of DFM for help with the calculation of the reference values and analysis of the participants’ data.

The author acknowledges the National Measurement System Policy Unit of the UK Department of Trade and industry for providing the funding for some of the work described here.

References

[1]  Davis R. The SI unit of mass Metrologia, 2003, 40, n°6, 299-305

[2] Girard G. International Report: The Third Periodic Verification of National Prototypes of the Kilogram (1988-1992) Metrologia, 1994, 31, n°4, 317-336

© Crown copyright 2006. Reproduced by permission of the Controller of HMSO.

# Comments

Mass is a quantity, not a unit.

Mass is a quantity, not a unit.

Actually, 651 was traded to NPL for an ingot cast by Engelhard.