Thermometry

EUROMET Project No 714

Comparison of realization of the triple-point of water

Final Report

E. Renaot

Intermediate Temperature Laboratory

Phone : 33 1 40 27 20 21

Fax     : 33 1 42 71 37 36

Email : renaot@cnam.fr

LNE-INM/CNAM*

292 rue Saints Martin

75141 Paris Cedex

France

* BNM-INM/CNAM until 1 January 2005

1. Introduction	3

2. TPW and isothermal enclosure	4

3. Procedures	6

4. Description of the local facilities	7

5. Control of the transfer cell during the comparison	7

6. EUROMET Comparison Reference value (ECRV)	8

7. Uncertainties	9

8. Results	11

9. Conclusions	11

# 1. Introduction

A comparison of the different triple point of water (TPW) realizations in Europe has been organised under the auspices of EUROMET (projet N°714). This project is a re-activation of the :

- EUROMET Comparison 278 in which twelve countries took part between January 1994 and June 1997 [1].

- EUROMET Comparison 549 in which fifteen countries took part between February 2000 and October 2003 [2].

The re-activation answers to a request of four more countries. The goal of this project is not to deal with an extensive research on triple-point of water behaviour. The aim of this project based on the circulation of one cell and an adapted isothermal enclosure, was to assess the uncertainties associated to the practical realization of the triple point of water in the various European laboratories.

The LNE-INM/CNAM supplied the circulating TPW cell and the isothermal enclosure. It established the schedule and followed the progress of the comparison. Table 1 lists the participating laboratories and Table 2 the order in which the measurements were conducted.

The stability of the circulating TPW cell during the time of the comparison was studied by LNE-INM/CNAM.

The national metrological institutes of five countries: the LNE-INM/CNAM, METROSERT, MSA, LNMC, NSC (see Table 1 for acronyms), were involved in this work which lasted from June 2004 to October 2005. The comparison lasted longer than envisaged because six months were necessary to obtain the customs documents required for transfer the circulating cell and isothermal container from Europe to Ukraine.

Table 1. Participating laboratories

| Laboratory | Country | Participants |
| --- | --- | --- |
| (Laboratoire National d’Essais-Institut National de Métrologie/ Conservatoire National des Arts et Métiers (LNE-INM/CNAM), pilot laboratory. (AS METROSERT Ltd (METROSERT) (Malta Standard Authority, National Metroly Services (MSA) ( Latvian National Metrology Centre Ltd (LNMC ) (National Scientific Centre « Institute of Metrology (NSC) | France Estonia Malta Latvia Ukraine | E. Renaot (coordinator) M. Hoang R. Vendt N. Testa, J. Bartolo A. Klints, L. Nazarenko |

Table 2: Comparison programme

| Laboratory |  |
| --- | --- |
| LNE-INM/CNAM METROSERT MSA LNMC NSC LNE-INM/CNAM | June 2004 July - August 2004 September - November 2004 December 2004 - January 2005 July-August 2005 September 2005. |

# 2. TPW and isothermal enclosure

The circulating triple point of water cell is an NPL-made cell N°679. This cell was also the circulating cell for the EUROMET Projects 278 and 549. During the last twelve years this cell has been compared several times to one cell (N°673) belonging to the batch of water triple point cells that constitutes the LNE-INM/CNAM reference for this fixed point. LNE-INM/CNAM keeps a chronological accounts of these comparisons (see Figure 3).

The isothermal enclosure was designed and constructed at LNE-INM/CNAM (Figure 1). During the measurement, with a normal platinum resistance thermometer (length of the glass sheath: 48 cm) part of the head is inside the enclosure’s cover. This arrangement prevents visible and infrared radiation from penetrating the ice and reaching the thermometer sensor. The bottom of the SPRT is not in direct contact with the bottom of the thermometer well. The distance is approximately 5 mm. For thermometers longer than 48 cm, the laboratory had to adjust the position of the thermometer in order to position its bottom approximately 5 mm above the bottom of the well.  Several holes in the cell’s cover and in the lower part of the support allow the water resulting from melted ice to be drained away. This water is then pumped out from the bottom of the enclosure.

Figure 1: Isothermal enclosure

1: cover of the enclosure; 2: plastic cover of the cell full of crushed ice; 3: water extraction; 4: isothermal container; 5: cell holder; 6: crushed ice; 7: plastic container with foam rubber; 8: base containing drain holes; 9: water container

The LNE-INM/CNAM studied the immersion-temperature effects of cell N° 679 in the enclosure by using a Leeds & Northrup thermometer (N°1807664). The experimental measurements agree with the theoretical values calculated using dT/dh given by the ITS-90, see figure 2. All the thermometers used in this comparison were glass-sheathed SPRT. The behaviour relative to the depth of immersion was expected to be the same for all SPRTs.

Figure 2: Immersion temperature effects

# 3. Procedures

The aim of this project was to allow each participating laboratories to compare the temperature water triple-point realized by their local facilities (cell+enclosure+procedure) to the temperature materialised by the circulating instrument. The local preparation using the local cell and the local enclosure was performed according to the local procedure, whereas the realization with the circulating instrument was strictly defined by a precise procedure common to all the participating laboratories. This protocol is very close to the one used for the EUROMET Projects 278 and 549.

LNE-INM/CNAM tested, at the beginning and the end of the programme, the stability of the circulating instrument by measuring (T673-T679) (See part 5).

All the participants were required to use the same procedure for preparing the circulating cell. First, the isothermal enclosure is filled with crushed ice. When the isothermal enclosure is cooled, the circulating cell is introduced into the plastic container (7 in Figure 1). After, at least 3 hours alcohol is introduced into the thermometer well. In order to prepare the ice mantle a metal rod precooled in liquid nitrogen is inserted into the thermometer well. This operation is repeated several times in order to obtain an adequate mantle thickness (4 to 8 mm) which is uniform over its whole length. The alcohol is then removed and the well is washed out with pure water or alcohol precooled to a temperature of 0°C. Finally, precooled pure water or alcohol is once again poured into the well. The level of this liquid is adjusted to that of the free surface water in the cell when the thermometer is present. To avoid exothermic heat during measurements, special care must be taken to prevent mixing water and alcohol in the thermometer well. A second ice/water interface, immediately adjacent to the well surface, is formed by producing a layer of water (melted ice) by inserting a rod at room temperature into the well for about 30 seconds. The effectiveness of this layer should be verified prior to any measurement by checking the free rotation of the mantle.

The cell must left at rest for at least 20 hours in order to release any mechanical stress in the ice mantle. Three cycles of measurement using different freeze was prescribed by the comparison procedure. During one cycle the measurements must be perform at least the second , third, and fifth day after the realization of the mantle

The comparison was performed by measuring the difference in temperature between the circulating and local TPW cells. The difference in the observed resistances (corrected for the hydrostatic head effect and self-heating, and possible calibration of the measurement instrument) for the two cells was converted to a temperature difference using the dT/dR for the SPRT.

Tlocal  -  Tcirculating =[ Rlocal  -  Rcirculating] . [dT/dR]TPW

# 4. Description of the local facilities

Table 3 presents the cell features and the different local procedures used for this comparison (see page 14) . The measurements are performed using the local facilities given Table 4.

Table 4: Measurement facilities

| Laboratory | Bridge | SPRT |
| --- | --- | --- |
| LNE-INM/CNAM | Guildline 9975 / d.c. | Leeds & Northrup, Hart Scientific |
| METROSERT | MI6010B / d.c. | Isotech |
| MSA | ASL F18 / a.c. | Hart Scientific |
| LNMC | ASL F700 / a.c. | Isotech |
| NSC | CA300 / a.c. | VNIIM |

# 5. Control of the transfer cell during the comparison

The LNE-IN/CNAM checked the stability of the cell N° 679 during the comparison. It was compared with cell N°673 at the beginning and at the end of the comparison (Table 5). Variations in T673-T679 were analysed assuming that the cells do not change in the same way. Taking into account the combined standard uncertainty on T673-T679 (0.035 mK) the temperature realized in cell No 679 may be considered as stable during the entire comparison.

Table 5:  Stability test on cell No 679 during the comparison

| Dates |  | T673 - T679 |
| --- | --- | --- |
| June 2004 |  | - 0.006 mK |
| September 2005 |  | 0.022 mK |

In order to be able to compare the results obtained during the EUROMET Projects 278, 549 and 714 the Table 6 gives the means of the differences T673 - T679 measured by the LNE-INM/CNAM during the EUROMET Projects 278, 549 and 714.

Table 6: Means of the differences T673 - T679 measured during the EUROMET projects 278, 549 and 714

|  | T673 - T679 |
| --- | --- |
| EUROMET comparison 278 | 0.007 mK |
| EUROMET comparison 549 | 0.054 mK |
| EUROMET comparison 714 | 0.008 mK |

# 6. EUROMET Comparison Reference value (ECRV)

Only five laboratories are involved in this comparison, it is insufficient to be able to define a reliable reference value. So, it is suggested to use the reference value defined previously at the time of EUROMET comparison 549.

So, in this report, all the results obtained by the participants are analysed using the ECRV defined in the report of the EUROMET Comparison 549 [2]. This reference value corresponds to the mean of the twenty-six values of (Tlocal- Tcirculating) carried out by the participants in the EUROMET Comparison 549.

TECRV714= TECRV549

(TECRV549 -Tcirculating)  = 0.067 ( 0.030 mK [2].

The uncertainty contribution u(stability549,714) from the stability of the cell 679 between the EUROMET comparisons 549 and 714 has to be added to this uncertainty.

(TECRV714 -Tcirculating)  = 0.067 ( 0.058 mK

# 7. Uncertainties

In this comparison, we apply the general rules for expressing uncertainty in physical measurement as they are established in the ISO Guide [3].

The comparison consisted of measuring the difference in temperature between the circulating instrument and the local one, therefore the laboratories had to evaluate the uncertainty components related exclusively to this difference. So, the uncertainty budget doesn’t included the sources of uncertainty affecting the local and the circulating realizations of the triple point of water as: chemicals impurities, gas pressure, spurious heat flux,….

In this interlaboratory comparison we consider the contribution of some uncertainties to be negligible because these uncertainties are expected strongly positively correlated, for example:

- The electrical measurement on the circulating and the local cells are made with the same bridge and practically the same ratio.

- The same standard resistor stabilized in temperature is used.

- The thermal resistances have approximately the same magnitude in in circulating and local cells. The corresponding self-heating corrections are very close.

So, the following sources of uncertainty :

- Bridge accuracy

- Resistor calibration.

- Self heating-correction

are neglected.

Consequently, the standard uncertainty in the difference is smaller than the standard uncertainty in the value of the SPRT resistance.

A] Uncertainty on (Tlocal - Tcirculating)Lab  reported by the laboratory

Commentary on the uncertainty components

Type A evaluation

Component uA1

This corresponds to the repeatability of measurement results under the following conditions :

- Same measurement procedure

- Same observer

- Same Thermometer left in place

- Same bridge

- Same isothermal enclosure

- repetition over a short period of time ( i.e. without changing the ice mantle)

This component takes in account the noise affecting the measurements.

Component uA2

Through the different measurements performed in a given Laboratory on TPW cells the standard deviation of the measurements can be obtained. The standard deviation can be considered to be an estimate of the reproducibility of the measurements due to changes in the influencing quantities.

The observed temperature differences between the circulating and the local cells dependent on various factors including the following :

- crystal size

- age of the mantles

- different between mantles

- method of handling the cells before preparation of the mantle. If the cell is stored for a considerable time in the same position before preparation of the mantle, the impurities may not be uniformly distributed in the water and handling can then modify the local impurities concentration

- undesirable thermal reaction in thermal contact fluids (mixing alcohol + water)

- stability of SPRT (repeated measurements with reinsertion of  SPRT)

- different types of SPRT's (thermometer’s geometrical characteristics can influence the observed difference in temperature between the circulating and local TPW cells).

- influence of spurious heat flux linked with the temperature of the surrounding

Type B evaluation

Component uB1

The source of this standard uncertainty is electrical measurement. This component takes in account:

- the stability of bridge

- the stability of reference resistor (temperature effect)

during the time of a pair of the measurements (one on local cell, one on circulating cell)

Component uB2

This is the uncertainty of the correction related to the hydrostatic pressure. This component takes in account the uncertainties associated at once to the local and the circulating cells.

In this interlaboratory comparison we consider the contribution of some uncertainties to be negligible either because their values are very small, or because these uncertainties are strongly positively correlated. This applies, in particular to the uncertainty on the self-heating correction. . All the measurements are corrected for self-heating effect. As the thermal resistances have approximately the same magnitude in circulating and local cells the difference between the self-heating correction is very small. In addition the uncertainties on self-heating corrections in circulating and local cells are strongly correlated.

For each Laboratory the combined standard uncertainty, u(Tlocal - Tcirculating)Lab is calculated from the different components presented Table 6.

B] Final uncertainty on (Tlocal - Tcirculating)

The uncertainty contribution from the stability of the cell 679, during the EUROMET comparison 714 has to be added to the value of u (Tlocal - Tcirculating)Lab

So, finally

The corresponding expanded uncertainty (coverage factor=2) is given Table 3

# 8. Results

The aim of the comparison is to compare each local realization with that realized in the circulating one. The results of the temperature differences derived from the resistance measurement are given in Table 3. The differences correspond to averages.

The results (Tx- Tcirculating) with their related expanded uncertainties are presented Figure 3. Tx corresponds to Tlocal or TECRV714.

For each cell the value (Tlocal - TECRV714) was calculated.

(Tlocal - TECRV714) = (Tlocal - Tcirculating) - (TECRV714  - Tcirculating)

The values of (Tlocal-TECRV714) are given Table 7.

The uncertainty on (Tlocal-TECRV714) is obtained from:

The corresponding expanded uncertainty (coverage factor=2) is given Table 7

# 9. Conclusions

EUROMET project No 714 was based on the circulation of one TPW cell and an isothermal enclosure. Five European Laboratories were involved in the comparison of seven TPW cells. During the comparison, the stability of the circulating TPW was periodically checked by comparisons with an other LNE-INM/CNAM cell.

The results of this comparison are analysed using the Reference value of the EUROMET comparison 549.

Bibliographie

[1] E. Renaot, M. Elgourdou and G. Bonnier, Interlaboratory comparison of realization of the triple point of water, Metrologia 2000, 37, 693-699.

[2] E. Renaot and all, Comparison of realization of the triple-point of water (EUROMET Project No 549), Proceedings of TEMPMEKO 2004, 22-25 June Cavtat, Croatia, 1003-1008

[3] Guide to the expression of uncertainty in measurement, ISO/TAG 4/WG 3; June 1992

Figure 3: Temperature differences derived from the resistance measurement

Table 3 : cell features, local procedures, results and associated uncertainties

| Labo. |  | Cell |  | Preparation | Isothermal enclosure and bath | Tlocal-Tcirculating | U(Tlocal-Tcirculating)Lab | U(Tlocal-Tcirculating)end |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Type | N° | Delivery date |  |  | in mK | in mK | in mK |
| LNE-INM/CNAM | NPL | 673 | 1993 | Metal rod precooled in liquid nitrogen | ISOTECH bath | 0.008 | 0.070 | 0.078 |
| METROSERT | Jarret JA-11 ISOTECH B11 ISOTECH E11 | A11-2027 B11-320 E11-284 | 1999 2003 2003 | Metal rod precooled in liquid nitrogen | Hart Scientific 7012 bath | 0.112 0.087 0.101 | 0.148 0.146 0.150 | 0.152 0.150 0.154 |
| MSA | Hart Scientific 5901 | 1379 | 2004 | Metal rod precooled in liquid nitrogen | Ice bath, Dewar model 2028 | 0.040 | 0.240 | 0.242 |
| LNMC | ISOTECH E11 | E11-241 | 2002 | Immersion ftreezer Quick stick 2031 and dry ice | Hart Scientific 7312 | ( 0.001 | 0.311 | 0.313 |
| NSC | VNIFTRI | 9209 | 1996 | Metal rod precooled in liquid nitrogen | Ice bath NSC | -0.044 | 0.050 | 0.060 |

Table 6: Uncertainty budgets

| Laboratories | LNM-INM/CNAM | METROSERT | MSA | Latvia | Ukraine |
| --- | --- | --- | --- | --- | --- |
| Components |  |  |  |  |  |
| A1 | 0.015 | 0.007 | 0.060 | 0.050 | 0.008 |
| A2 | 0.025 | 0.033 | 0.100 | 0.068 | 0.014 |
| B1 | 0.020 | 0.067 | 0.021 | 0.111 | 0.017 |
| B2 | 0.005 | 0.004 | 0.015 | 0.069 | 0.004 |
| Combined standard uncertainty - | 0.035 | 0.075 | 0.120 | 0.156 | 0.023 |

Table 7 : Summary of results and associated uncertainties

| Laboratory | Cell | Tlocal-Tcirculating in mK | Tlocal-TECRV714 in mK | U(Tlocal-TECRV714) in mK |
| --- | --- | --- | --- | --- |
| LNM-INM/CNAM | 673 | 0.008 | -0.059 | 0.097 |
| METROSERT METROSERT METROSERT | 2027 320 284 | 0.112 0.087 0.101 | 0.045 0.020 0.034 | 0.162 0.161 0.164 |
| MSA | 1379 | 0.040 | -0.027 | 0.249 |
| LNMC | 241 | ( 0.001 | -0.067 | 0.318 |
| NSC | 9209 | -0.044 | -0.111 | 0.084 |