NPL Report MAC CP1 (10 pt) 


NPL REPORT

DEPC-EM 015

Calibration of two Spinning Rotor Gauges as a bilateral comparison between NPL and INRIM in the range 3 · 10-4 Pa to 9 · 10-2 Pa
P A Carroll , D Mari (INRIM), J C Greenwood, G Raiteri (INRIM), 

G. Rumiano (INRIM)

Not Restricted

February 2007 

Calibration of two Spinning Rotor Gauges as a bilateral comparison between NPL and INRIM in the range 3 · 10-4 Pa to 9 · 10-2  Pa

P A Carroll , D Mari (INRIM), J C Greenwood , G Raiteri (INRIM), G. Rumiano (INRIM)

DEPC

ABSTRACT

This report details measurements made at NPL and INRIM between July and August 2005 which form a comparison to determine the degree of equivalence of absolute pressure standards at these national metrology institutes.

( Crown copyright 2007

 Reproduced with the permission of  the Controller of HMSO

 and Queen's Printer for Scotland

ISSN 1744-0254

National Physical Laboratory

Hampton Road, Teddington, Middlesex, TW11 0LW

Extracts from this report may be reproduced provided the source is acknowledged and the extract is not taken out of context. 

Approved on behalf of the Managing Director, NPL

by Susan Evans, Director, DEPC

Content

Page number :

41. Introduction


52. Participating laboratories and their standards


52.1 The NPL series expansion system


62.2 The INRIM continuous expansion standard


73. Transfer standards


74. Calibration constant


95. Organisation of the comparison and schedule


96. Calibration procedure and reported results


107. Uncertainties of primary standards


118. Uncertainties of reported measured values


139. Reported results of the laboratories


1610. Stability of Transfer gauges


1711. Data reduction and evaluation of the degree of equivalence


2212. Discussion and conclusions


2413. References




1. Introduction

This report details measurements made at NPL and INRIM between July and August 2005 which form a comparison to determine the degree of equivalence of absolute pressure standards at these national metrology institutes.

The format and certain general portions of the text of the report follow the EUROMET comparison report (EUROMET.M.P-K1.b, Project 442B, July 2000 [1]) and a subsequent bilateral comparison between PTB and CENAM (key comparison SIM-EUROMET.M.P-BK3, August 2004 [2]).

The transfer standard package consisted of two spinning rotor gauges (SRGs). The gauges were supplied by PTB and are those used in both  EUROMET.M.P-K1b [1] and SIM-EUROMET.M.P-BK3 [2]. These have been shown to have excellent stability during comparisons, particularly due to the novel transportation mechanism which allows the rotors to travel in a fixed position under vacuum [3]. For the purposes of this bi-lateral comparison NPL acted as the pilot laboratory.

Comparison measurements at NPL were performed simultaneously for the two spinning rotor gauges, each being operated by a MKS SRG2 controller with a SH-700 sensing head. INRIM has only one MKS SRG2-CE control unit so their calibrations were performed independently.

Data was collected and analysed according to the protocol of the SIM-EUROMET.M.P-BK3 that was completed between April 2002 and June 2002 [2].

At both NPL and INRIM, the measurements were performed on consecutive days using the standards described below. The following target pressures were used for the calibration: 3·10-4 Pa, 9·10-4 Pa, 3·10-3 Pa, 9·10-3 Pa, 2·10-2 Pa and 9·10-2 Pa.

The report begins with a brief description of the vacuum standards used and their principle of operation. The measurements are then described and summarised. Finally, the evaluation of various uncertainty estimates is explained.

2. Participating laboratories and their standards

	Laboratory
	Standard
	Character of Standard
	Traceable to 
	CMC Listed

	NPL
	Static expansion system
	Primary
	independent
	yes

	INRIM
	Continuous expansion system
	Primary
	independent
	yes


Table 1. Details of participating standards.

2.1 The NPL series expansion system

The NPL series expansion system SEA2 used for this comparison is a CMC listed, independent primary standard which is shown schematically below.

[image: image63.wmf] 

A

 

buffer

 

inlet pressure

 

Q

 

P

 

L

 

O

 

K

 

F

 

E

 

D

 

C

 

B

 

N

 

S

 

R

 

M

 

H

 

G

 

V4

 

v4

 

V3

 

v3

 

V2

 

v2

 

V1

 

v1

 


Figure 1.    Schematic representation of the vacuum standard SEA2
Using this apparatus, calculable pressures between 5·10-7 Pa and 900 Pa may be generated using the series expansion method.

In operation, a sample of gas is contained in one of the small vessels (i.e. v1, v2, v3 or v4) and then expanded into the next large and small vessels, which have been previously evacuated to a high vacuum. This procedure is then repeated using subsequent expansion stages until the gas is expanded into the calibration vessel (V4). The pressure of the initial gas sample is measured using a calibrated resonant silicon gauge. By varying the initial pressure and the number of expansion stages, a range of pressures may be generated in the calibration vessel. The pressure generated is calculated from knowledge of the initial pressure, the gas temperature and the ratio of the relevant volumes e.g. (v3+V3+v4)/v3.

2.2 The INRIM continuous expansion standard

The INRIM continuous expansion system involved in this comparison is based on the expansion of a pure gas, which is pumped through a fixed conductance. 

Figure 2.    Schematic representation of the INRIM vacuum standard.
The fixed conductance was evaluated by both analytical and Monte Carlo methods. 

The effective pumping speed is computed from the values of calculated conductance and the ratio of the conductance to the pumping speed (periodically measured). 

The gas flow-rate device used to generate the pressures into the fixed conductance IMGC-CNR-D2 system is based on the constant pressure – variable volume method. The upper limit of the molecular range of the system is 10-1 Pa. 

3. Transfer standards

The spinning rotor gauge is widely accepted as a transfer standard due to its measurement accuracy and long-term and transport stability. Two devices were used in order to reduce the influence of transport instabilities, to provide redundancy and to increase the accuracy of the comparison.

There has still been no published evidence to indicate that a change of controller and sensing head has an influence on the calibration results of a rotor/finger combination. It was therefore permissible for each laboratory to use their own controllers.

The details of the two spinning rotor gauges used as transfer standards are shown in Table 2.

	Transfer Standard
	Rotor 1
	Rotor 2

	Identifier
	25 UM
	23 UM

	Material
	Etched Stainless Steel 
	Invar

	Nominal diameter
	4.762 mm
	4.50 mm

	Nominal density
	7.715  g /cm3
	8.2  g /cm3


Table 2 Transfer standards used for comparison.

Data exists from 1994 onwards for the calibration of each rotor at PTB. [2] 

4. Calibration constant

As reported elsewhere [1], [2], the value to be calibrated by each laboratory j for each rotor i was the effective accommodation coefficient (ij which is mainly determined by the tangential momentum accommodation coefficient of the gas molecules to the rotor, and partly by the energy accommodation factor and additionally by using nominal values for diameter and density of the rotors instead of the real ones.

(ij was determined by the following equation:

                           
[image: image1.wmf](

)

÷

÷

ø

ö

ç

ç

è

æ

-

÷

ø

ö

ç

è

æ

=

w

w

w

r

p

p

s

i

i

stj

i

i

j

ij

RD

p

d

m

kT

&

20

8

                                    (1)

where :


[image: image2.wmf]st

p

 = is the generated pressure in the standard


[image: image3.wmf]j

T

 = the temperature of the gas in the calibration vessel


[image: image4.wmf]i

d

= the nominal diameter of the rotor i


[image: image5.wmf]i

r

= the nominal density of the rotor i

[image: image6.wmf]m

= the molecular mass of nitrogen

(
[image: image7.wmf]w

&

/
[image: image8.wmf]w

)i = the relative deceleration rate of the rotor frequency 
[image: image9.wmf]w

 (DCR)


[image: image10.wmf]RD

i = pressure independent residual drag


[image: image11.wmf]RD

 is a function of 
[image: image12.wmf]w

, RD(
[image: image13.wmf]w

), so whenever a determination of (ij is made it was necessary to measure the value of 
[image: image14.wmf]w

 in order to subtract the correct RD(
[image: image15.wmf]w

) in Eq.(1).

 For measuring RD(
[image: image16.wmf]w

) two different approaches were used by the participants.

At INRIM : Before starting the calibrations the rotor frequency dependence of the residual drag (in units DCR = s-1) was measured over a long period of time. Furthermore, before each of the three calibration cycles of each rotor, the behaviour of the offset signal was checked to verify that it was comparable with the estimate of RD(
[image: image17.wmf]w

) obtained by the previous “long” acquisition. The rotor frequency covered the full range that may occur during the calibration. The residual pressure in the standard was about 1·10-6 Pa.

At NPL : RD(
[image: image18.wmf]w

) was determined during the course of the calibrations by pumping down the vacuum system to residual pressure conditions after each target pressure point.

In both cases a linear least square fit is applied to obtain the function RD(
[image: image19.wmf]w

).

5. Organisation of the comparison and schedule

Two spinning rotor gauges were used as transfer standards as described in Section 3. An initial calibration was completed at NPL before the standards were transported by courier to INRIM. The novel transport mechanism designed by PTB enabled the transportation of the rotors in fixed positions whilst under vacuum [3].  

Following calibration at INRIM, the standards were returned to NPL for a repeat calibration.

To monitor any transport instability, a third spinning rotor gauge (check standard) was calibrated by NPL at the same time as the transfer gauges. It was not removed from the NPL primary standard throughout the duration of the comparison.

The chronology of the calibrations is shown below in Table 3.

	Calibrating Laboratory
	Dates
	Notes

	NPL
	July 12, 13, 14 2005
	Afterwards, sent by courier to INRIM

	INRIM
	July 27, 28, 29 and August 2, 3, 4  2005
	Afterwards, sent by courier to NPL

	NPL
	August 24, 25, 26  2005
	Comparison Finished


Table 3. Chronology of measurements

6. Calibration procedure and reported results

The calibration procedure agreed upon before the comparison was similar to that of the SIM-EUROMET.M.P-BK3 bilateral comparison [2] : Each laboratory calibrated the two spinning rotor gauges at 6 nominal target pressures pt in ascending order: 3·10‑4 Pa, 9·10-4 Pa, 3·10-3 Pa, 9·10-3 Pa, 2·10-2 Pa and 9·10-2 Pa..

A tolerance of ( 10% in establishing the nominal pressure was accepted for pt < 9·10‑2 Pa and ( 5% for 9·10-2 Pa. Each target pressure had to be generated 3 times. After a measurement at the target point, the system was pumped down to residual pressure conditions and the same point re-generated. In total 18 measurements were made in each calibration sequence. It was required that this calibration sequence be repeated on three separate days.

The  number of readings and sample interval times for each point were as follows:

·  30s with 10 repeat readings for the target points 3·10-4 Pa, 9·10-4 Pa and 3·10-3 Pa.

·  30s with 5 repeat readings for 9·10-3 Pa, 2·10-2 Pa and 9·10-2 Pa

It was required that the offset be measured during the calibration no matter what method had been chosen for offset determination: 

·  at the beginning of the calibration sequence.

·  after the target pressures 3·10-4 Pa, 9·10-4 Pa, 3·10-3 Pa and 9·10-2 Pa (at the end of the calibration sequence).

It was agreed that no bake-out should be performed as part of the calibration sequence since it is both time consuming and may also affect the accommodation coefficient.

At the end of this calibration procedure, for each generated pstj near the respective target point and for each rotor i and for each of the calibration sequences a value for (ij existed. Together with the value of (ij ,  each laboratory also reported the standard uncertainty u(pstj) of pstj.

7. Uncertainties of primary standards

Table 4 shows the relative standard uncertainties due to Type B uncertainties for the two primary standards. Type A uncertainties of the standards will be incorporated in the scatter of data for repeat measurements.

	
[image: image20.wmf]st

p

 (Pa)
	NPL
	INRIM

	3.00E-04
	3.04E-03
	7.30E-03

	9.00E-04
	2.98E-03
	6.40E-03

	3.00E-03
	2.97E-03
	4.70E-03

	9.00E-03
	2.97E-03
	3.60E-03

	2.00E-02
	2.97E-03
	3.40E-03

	9.00E-02
	2.58E-03
	3.20E-03


Table 4 Relative standard uncertainties of generated pressures due to systematic effects (Type B) as reported by participants.

8. Uncertainties of reported measured values

Again, following the analysis of key comparison SIM-EUROMET.M.P-BK3 [2], we introduce:


[image: image21.wmf]20

15

.

296

8

i

i

i

d

m

k

K

r

p

p

×

=

   ,





 (2)

and


[image: image22.wmf]i

i

DCR

÷

ø

ö

ç

è

æ

=

w

w

&








 (3)

The effective accommodation coefficient has been measured n = 9 times (ijk , k=1 … n, therefore we can write :


[image: image23.wmf](

)

(

)

w

s

ik

ik

stjk

i

jk

ijk

RD

DCR

p

K

T

-

×

=

15

.

296





 (4)

and Eq. (1) can be replaced by taking the mean of the repeated measurements to give


[image: image24.wmf]å

=

=

n

k

ijk

ij

n

1

1

s

s







 (5)

The same values for k , m , di , and (i were used by each laboratory and the Ki were therefore fully correlated, so that in effect no uncertainty needed to be attributed to Ki (a systematic error in di for example, would be calibrated into (ij in the same way at each laboratory and would have no effect on the result of the comparison). All type A uncertainties of values on the right hand side of Eq. (4) will contribute to the scatter of (ijk and hence the standard deviation of (ij. Therefore since Type A uncertainties are accounted for by the standard deviation, for calculation of the overall uncertainty of (ij only the type B uncertainties of values on the right hand side of Eq. (4) have to be evaluated for inclusion. For DCRi there is no such uncertainty. RD(() is not determined at the same time as DCRi , but before or after the measurement, and so therefore has a Type B uncertainty. Also the gas temperature and the generated pressure pstj will have Type B uncertainties which are known before the measurements.

For this reason the variance in (ij is given by:


[image: image25.wmf]2

2

2

2

2

2

2

3

1

stj

j

i

ij

ij

p

stj

ij

T

j

ij

RD

i

ij

u

p

u

T

u

RD

s

n

n

u

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

+

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

+

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

+

-

-

=

s

s

s

s

s

  ,
             (6)

where 
[image: image26.wmf]2

ij

s

s

 is the square of the standard deviation of the mean of the repeat measurements (ijk and where we understand that all standard uncertainties u are due to systematic effects which do not contribute to the scatter of (ij. Since only 9 measurements were taken, 
[image: image27.wmf]ij

s

s

was multiplied by 
[image: image28.wmf](

)

(

)

3

/

1

-

-

n

n

 , i.e. 1.33 for n = 9  as suggested by Kacker and Jones [4]. The last term in Eq. (6) is due to the generated pressure as described in section 7, all other terms are due to uncertainties of the transfer standard.

The sensitivity coefficients are:


[image: image29.wmf]15

.

296

j

stj

i

i

ij

T

p

K

RD

×

-

=

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

s






 (7)


[image: image30.wmf](

)

i

i

j

stj

i

stj

ij

RD

DCR

T

p

K

p

-

×

-

=

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

15

.

296

2

s

                           (8)


[image: image31.wmf](

)

i

i

j

stj

i

j

ij

RD

DCR

T

p

K

T

-

×

×

=

÷

÷

ø

ö

ç

ç

è

æ

¶

¶

1

15

.

296

1

2

1

s



 (9)

u(
[image: image32.wmf]stj

p

) and u(Tj) were reported by each laboratory, where it is assumed that Tjk ( Tj (constant temperature of standard during 3 repeat measurements). The following effects may contribute to the uncertainty of the residual drag RDi: 

· the scatter of the measurement results,

· the imprecisely known frequency dependence of the offset,

· a possible drift of the offset values between it’s determination and the time of calibration,

9. Reported results of the laboratories

As NPL had carried out two calibration sequences and INRIM one, it was possible to check to see if a significant change could be observed between (ij for each sequence. 
No shift in the mean (ij value was seen at NPL between the results collected in July 2005 and August 2005 for both of the spinning rotor gauges used as transfer standards.

Therefore the mean value of all data for a single rotor at a single target pressure could be taken for data reduction:


[image: image33.wmf]å

=

=

n

k

ijk

ij

n

1

1

s

s

     
n = 9                          

             (5)

The results reported by each laboratory and the corresponding uncertainties according to Eq, (6) are shown in the following Tables and Figures.

	Pt /Pa
	NPL   1
	Jul 05
	INRIM 1
	Jul/Aug 05
	NPL  2
	Aug 05

	
	(1
	u((1)
	(1
	u((1)
	(1
	u((1)

	3.00E-04
	1.1467
	0.0145
	1.1657
	0.0088
	1.1461
	0.0171

	9.00E-04
	1.1461
	0.0072
	1.1661
	0.0075
	1.1455
	0.0079

	3.00E-03
	1.1494
	0.0039
	1.1645
	0.0055
	1.1469
	0.0040

	9.00E-03
	1.1499
	0.0036
	1.1626
	0.0042
	1.1485
	0.0036

	2.00E-02
	1.1504
	0.0036
	1.1617
	0.0040
	1.1491
	0.0036

	9.00E-02
	1.1505
	0.0031
	1.1509
	0.0037
	1.1484
	0.0031


Table 5. The mean values (2j of the reported results for Rotor 1 (25 UM) and the uncertainties as 

calculated by Eq.(6).


[image: image34.wmf]Rotor 1 (25 UM)

1.140

1.145

1.150

1.155

1.160

1.165

1.170

1.00E-04

1.00E-03

1.00E-02

1.00E-01

P ref (Pa)

Accomodation Factor

NPL July 2005

NPL August 2005

INRIM July / August 2005


Figure 3. Graphical presentation of the results of Rotor 1 (Table 6). Uncertainties are not shown for better visibility and can be taken from table 5.
	Pt /Pa
	NPL   1
	Jul 05
	INRIM 1
	Jul/Aug 05
	NPL  2
	Aug 05

	
	(2
	u((2)
	(2
	u((2)
	(2
	u((2)

	3.00E-04
	1.1061
	0.0346
	1.1191
	0.0084
	1.1062
	0.0441

	9.00E-04
	1.1112
	0.0120
	1.1231
	0.0072
	1.1108
	0.0148

	3.00E-03
	1.1131
	0.0051
	1.1245
	0.0053
	1.1125
	0.0056

	9.00E-03
	1.1136
	0.0038
	1.1228
	0.0041
	1.1129
	0.0037

	2.00E-02
	1.1138
	0.0036
	1.1217
	0.0039
	1.1133
	0.0035

	9.00E-02
	1.1141
	0.0030
	1.1121
	0.0036
	1.1126
	0.0030


Table 6. The mean values (1j of the reported results for Rotor 2 and the uncertainties as calculated by Eq.(6) 


[image: image35.wmf]Rotor 2 (23 UM)

1.1000

1.1050

1.1100

1.1150

1.1200

1.1250

1.1300

1.0000E-04

1.0000E-03

1.0000E-02

1.0000E-01

P Ref (Pa)

Accomodation Factor

NPL July 2005

NPL August 2005

INRIM July / August 2005


Figure 4. Graphical presentation of the results of Rotor 2 (23 UM) (Table 5). Uncertainties are not shown for better visibility and can be taken from table 6.

The results show that the values obtained at INRIM are systematically higher than those measured at NPL for all but the highest test pressure. 

Rotor 1 shows a difference in accommodation factor ranging from 1.8 % to 1.0 % in the molecular regime between NPL and INRIM values, except for the target pressure 9·10-2 Pa, characterized by a difference of about 0.1 % . Rotor 2 shows a maximum difference in accommodation factor of around 1.2 % in the molecular regime between NPL and INRIM values.

Good transport stability is shown by the rotors as the values of accommodation factor for NPL2 are within 0.25% of those recorded during NPL1.

10. Stability of Transfer gauges

In order to monitor the transport stability of the calibration constant of the three rotors during the course of the calibration, the mean values of (1 and (2 between 9·10-4 Pa and 2·10-2 Pa of the pilot laboratory were calculated. Rotor 3 was the check standard which did not travel.

	
	(1
	s((1)
	(2
	s((2)
	(3
	s((3)

	NPL1
	1.1489
	0.0010
	1.1129
	0.0006
	1.0028
	0.0008

	NPL2
	1.1475
	0.0016
	1.1124
	0.0006
	1.0011
	0.0010


Table 7 The mean values of  (1 and (2 between 9(10-4 Pa and 2(10-2 Pa for the two NPL calibrations and the respective scatter (standard deviation s) of the mean. Rotor 3 did not travel and served as a check standard for NPL.

As was argued in SIM-EUROMET.M.P-BK3 [2], the impossibility of being able to calibrate the transfer standards simultaneously on the different primary standards requires the mean value of the pilot laboratory calibrations to be calculated. This gives an approximation of a coincident time at which both participants values can be compared.

The measured values between NPL1 and NPL2 are within the combined standard deviation for the two rotors which travelled and also for the check standard which did not travel.

It is therefore reasonable to assume that both the effective accommodation factors did not change during this bilateral comparison.

11. Data reduction and evaluation of the degree of equivalence 

As outlined in the previous section, it is reasonable to assume that the transport instability of the two spinning rotor gauges was zero during this comparison. For this reason all values taken at the pilot laboratory NPL can be assumed to belong to the same parent population and the mean of all this data can be taken.





[image: image36.wmf]å

=

=

n

k

ijk

ij

n

1

1

s

s

        n = 18                                                   (10)

The situation is similar to the SIM-EUROMET.M.P-BK3 [2] therefore the analysis presented in that paper can be followed. Since one of the goals of this comparison is to compare the generated pressures of the two standards, a pressure value has to be generated from the 
[image: image37.wmf]ij

s

. 

For the data from each laboratory and for each spinning rotor gauge i a value of indicated pressure 
[image: image38.wmf]ij

p

 for a common hypothetical target pressure pt can be calculated with the following equation:





[image: image39.wmf]ij

t

ij

p

p

s

×

=

             
[image: image40.wmf]2

,

1

=

i

       
[image: image41.wmf]2

,

1

=

j

                           (11)


[image: image42.wmf]ij

s

 denotes the mean accommodation coefficient according to Eq.(5) of spinning rotor gauge i as determined by INRIM  (j=1) or according to Eq (10) of SRG i as determined by NPL (j=2).

The method adopted here is to use 
[image: image43.wmf]ij

s

 to predict gauge readings that would be observed when the different standards are set to generate pressures of the same value at a coincident time. The difference in the predicted gauge readings is taken as an indicator of the difference between the true pressures actually realised by the different standards (“generated pressures”) when the two laboratories state the same calculated value. This latter difference between calculated pressures, when the standards are set to produce exactly the same transfer gauge reading near the target pressure, is to a very good approximation (provided the differences are small) equal to the difference in the predicted gauge readings but of opposite sign.

It was found by Jousten [5] that (ij may be slightly temperature dependent. The size of this effect varies from rotor to rotor and is also dependent on the specific surface condition of a single rotor. Relative temperature changes of ( ((eff / (T)/ (eff = -1∙10-4 / K to -4∙10-4 / K were found around room temperature.

The mean temperatures during calibrations at NPL (293.08 K) and INRIM (293.65 K) differed by 0.57 K. Since the temperature dependence of (eff of the transfer standards was not measured and may also have changed during the comparison, the only possibility to consider this effect is to add an uncertainty for the pij in Eq. (11)

If we assume a mean temperature of 293.365 K so that each laboratory differs by 0.285 K from this temperature, and a possible temperature dependence of ( ((eff / (T)/ (eff = -2∙10-4 / K to calculate the standard uncertainty of the temperature effect for the pij in Eq. (11)




[image: image44.wmf](

)

ij

ij

T

p

p

u

4

10

57

.

0

-

×

=

s






           (12)

Since pt is simply a numerical value without uncertainty, the uncertainty of pij is calculated from the following equation:


[image: image45.wmf](

)

(

)

(

)

2

2

÷

÷

ø

ö

ç

ç

è

æ

+

÷

÷

ø

ö

ç

ç

è

æ

=

ij

ij

T

ij

ij

ij

ij

p

p

u

u

p

p

u

s

s

s

                                            (13)


[image: image46.wmf](

)

ij

u

s

 was given in Eq. (6) and 
[image: image47.wmf](

)

ij

T

p

u

s

 in Eq. (12)

The following tables show the results for the 
[image: image48.wmf]ij

p

 obtained from Eqs. (11) and (13).

	Pt (Pa)
	P1NPL (Pa)
	u(P1NPL)
	P1INRIM (Pa)
	u(P1INRIM)

	3.0E-04
	3.439E-04
	3.356E-06
	3.497E-04
	2.641E-06

	9.0E-04
	1.031E-03
	4.817E-06
	1.049E-03
	6.791E-06

	3.0E-03
	3.444E-03
	8.464E-06
	3.494E-03
	1.655E-05

	9.0E-03
	1.034E-02
	2.304E-05
	1.046E-02
	3.813E-05

	2.0E-02
	2.299E-02
	5.047E-05
	2.323E-02
	7.994E-05

	9.0E-02
	1.035E-01
	1.991E-04
	1.036E-01
	3.363E-04


Table 8 : The predicted gauge readings and their uncertainties as obtained from Eqs. (11) and (13) for rotor 1.
	Pt (Pa)
	P2NPL (Pa)
	u(P2NPL)
	P2INRIM (Pa)
	u(P2INRIM)

	3.0E-04
	3.319E-04
	8.403E-06
	3.357E-04
	2.517E-06

	9.0E-04
	9.999E-04
	8.595E-06
	1.011E-03
	6.504E-06

	3.0E-03
	3.338E-03
	1.128E-05
	3.373E-03
	1.594E-05

	9.0E-03
	1.002E-02
	2.395E-05
	1.011E-02
	3.674E-05

	2.0E-02
	2.227E-02
	5.006E-05
	2.243E-02
	7.725E-05

	9.0E-02
	1.002E-01
	1.941E-04
	1.001E-01
	3.238E-04


Table 9 : The predicted gauge readings and their uncertainties as obtained from Eqs. (11) and (13) for rotor 2.

The degree of equivalence between the two standards can be tested using [6] :


[image: image49.wmf]iNPL

iINRIM

i

p

p

d

-

=

                                                              (14)

for each target point and each transfer standard i. The uncertainty u(
[image: image50.wmf]i

d

) is given by


[image: image51.wmf])

(

)

(

)

(

2

2

iNPL

iINRIM

i

p

u

p

u

d

u

+

=

                                       (15)

However, since this comparison covered 3 orders of magnitude of pressure, the use of a ratio is more appropriate in order to get a clear visualisation of the results. This ratio is defined as:


[image: image52.wmf]iNPL

iINRIM

i

p

p

r

=

                                                                           (16)

with





[image: image53.wmf]2

2

)

(

)

(

)

(

÷

÷

ø

ö

ç

ç

è

æ

+

÷

÷

ø

ö

ç

ç

è

æ

=

iNPL

iNPL

iINRIM

iINRIM

i

p

p

u

p

p

u

r

u

                               (17)

	Pt (Pa)
	r1
	r2
	u(r1)
	u(r2)
	u’(r1)
	u’ (r2)

	3.0E-04
	1.0168
	1.0117
	0.0123
	0.0264
	0.0097
	0.0253

	9.0E-04
	1.0177
	1.0110
	0.0080
	0.0107
	0.0043
	0.0084

	3.0E-03
	1.0143
	1.0105
	0.0053
	0.0058
	0.0014
	0.0027

	9.0E-03
	1.0116
	1.0086
	0.0043
	0.0044
	0.0009
	0.0012

	2.0E-02
	1.0104
	1.0073
	0.0041
	0.0041
	0.0008
	0.0010

	9.0E-02
	1.0013
	0.9989
	0.0038
	0.0038
	0.0008
	0.0008


Table 10 The ratios ri = 
[image: image54.wmf]iNPL

iINRIM

p

p

/

 and their uncertainties as determined by Eq.(16) and (17). See below for u’ description.

If  r1 and r2 are of the approximate same value significantly different from 1, this would clearly indicate a difference in the true generated pressures of the two standards. Generally the values of r1 and r2 will be slightly different due to the scatter of the data. r1 and r2 are correlated, because the same standards j were used to determine (1j  and (2j. In the appendix of reference [2] is outlined a way to consider this correlation by omitting u(pstj) in Eq.(6) for the determination of 
[image: image55.wmf](

)

ij

p

u

 in Eq.(13). The respective value is called u’ which is also listed in Table 10. The weighted mean r of r1 and r2 is then calculated for each target point by


[image: image56.wmf]2

2

2

1

2

2

2

2

1

1

)

(

'

/

1

)

(

'

/

1

)

(

'

/

)

(

'

/

r

u

r

u

r

u

r

r

u

r

r

+

+

=

                            
           (18)

The standard uncertainty of r at the respective target pressure pt is











           
[image: image57.wmf](

)

(

)

(

)

(

)

1

2

2

1

2

2

2

2

2

,

,

2

2

,

,

'

1

'

1

2

)

(

)

(

)

(

-

÷

÷

ø

ö

ç

ç

è

æ

+

+

+

÷

÷

ø

ö

ç

ç

è

æ

+

÷

÷

ø

ö

ç

ç

è

æ

=

r

u

r

u

p

p

u

r

p

p

u

r

p

p

u

r

u

t

t

T

NPL

st

NPL

st

INRIM

st

INRIM

st

s


           (19)

The first two terms under the square root describe the influence of the uncertainty of the standard pressures pstj that correlate to r1 and r2, the third term is the uncertainty contribution due to the temperature dependence of (eff as described before Eq. (12), and the last term in the bracket is due to all other influences of Type A, which are due to the rotor instability, offset determination RD, temperature and scatter of data. Also for uT( it is conservatively assumed that r1 and r2 and respectively (1j and (2j are correlated by their temperature dependence of (eff.

The quantity





[image: image58.wmf]1

-

=

r

d







           (20)

describes the relative difference between the two primary standards in this comparison, where




[image: image59.wmf])

(

)

(

r

u

d

u

=







           (21)

Equivalence is generally assumed if





[image: image60.wmf])

(

)

(

2

d

U

d

u

d

=

£

                                                                (22)

or





[image: image61.wmf]1

£

n

E

                                                                                  (23)

with 





[image: image62.wmf])

(

d

U

d

E

n

=







           (24)

The values generated from Eqs. (20) to (24) are summarised in Table 16 below.

	Pt (Pa)
	r
	d
	U(d)
	En 

	3.0E-04
	1.0162
	0.0162
	0.0242
	0.67

	9.0E-04
	1.0163
	0.0163
	0.0163
	1.00

	3.0E-03
	1.0135
	0.0135
	0.0116
	1.16

	9.0E-03
	1.0105
	0.0105
	0.0096
	1.10

	2.0E-02
	1.0091
	0.0091
	0.0092
	0.99

	9.0E-02
	1.0001
	0.0001
	0.0083
	0.01


Table 11 Summary of data as a result of this comparison. r is the ratio of the true generated pressures in the two standards (Eq.(18)), d is the relative difference between the pressures in the two standards (Eq.(20)), U(d) is the expanded uncertainty (k=2) of d (Eqs. (20), (21) and (22)) and En is defined by Eq. (24).

12. Discussion and conclusions 

The comparison has tested the degree of equivalence between the pressures generated at the primary vacuum standards of NPL and INRIM. 

The results show that the equivalence between the pressures generated at the primary vacuum standards of NPL and INRIM cannot be demonstrated for the target pressures 3(10-3 Pa and 9(10-3 Pa at the k=2 level of expanded uncertainty; for the other pressures the equivalence is demonstrated, even if for the target pressures 9(10-4 Pa and 2(10-2 Pa the En values are high, being respectively equal to 1.00 and 0.99. 

In December 2005 INRIM started a series of investigations of its continuous expansion standard. The procedure followed in the comparison was globally considered. The effect of different geometries of the tubing connecting the rotor to the calibration chamber and the effect of the residual pressure in the calibration chamber were investigated. The whole system was shown to be leak-tight following extensive leak detection.

Possible errors associated with flow rate measurements by the flowmeter were assessed. The stability of the capacitance diaphragm gauges used for the inlet pressure measurement, the effect of the initial position of the flowmeter piston and the leak tightness of the piston’s Teflon seal were checked. The effects of the bypass valve closure between the reference and measurement chambers on the zero signal of the differential capacitance diaphragm gauge and the temperature drift between the two chambers of the flowmeter were also all investigated. The results obtained have shown that all these effects have a negligible influence on the generated standard pressure.

Finally the effect of non-molecular flow [7] at the highest generated standard pressure (9∙10-2 Pa) was considered. The implementation of a correction taking into account this effect, as suggested by Jousten in a private communication with INRIM, would increase the accommodation coefficient value by about 1%  for the target pressure 9∙10-2 Pa. Applying  this correction to all the INRIM accommodation coefficient values in the considered pressure range for both rotors would show a  pressure independent systematic shift compared to the corresponding values of NPL. This effect could be due to the conductance between the calibration and pumping chamber. Consequently, the geometry of the INRIM orifice conductance will be re-evaluated in the near future. 
As a consequence of this comparison, the expanded uncertainty associated to the INRIM CMC entry for the pressure range from 9∙10-4  Pa  to 9∙10-2 Pa has been increased from 1.5∙10-2 ·p / Pa  to 3∙10-2 ·p / Pa to include  the shift.

The INRIM CMC entries related to the continuous expansion system pressure range will be revaluated as soon as other results of internal and bilateral comparisons are available.

13. References

[1] K Jousten et al, Results of the regional key comparison EUROMET.M.P-K1.b in the pressure range from 3 × 10-4 Pa to 0.9 Pa, Metrologia 42

[2] K Jousten, Luis A Santander Romero and Jorge C. Torres Guzman, Final report on key Comparison SIM-EUROMET.M.P.BK3 (bilateral comparison) in the pressure range from 3 × 10-4 Pa to 0.9 Pa, Metrologia 42 No 1A (Technical Supplement 2005)

[3] P.Röhl and W.Jitschin, Performance of the SRG with a novel transport device as a transfer standard for high vacuum, Vacuum 38 (1988), 507…509.

[4] R.Kacker and A Jones, On use of Bayesian statistics to make the guide to the expression of uncertainty in measurement consistent, Metrologia 40 (2003), 235…248

[5] K. Jousten, Is the effective accommodation coefficient of the spinning rotor gauge temperature dependant?, J.Vac.Sci. Technol. A 21(1) (2003), 318…324.

[6] M.G.Cox, The evaluation of key comparison data, Metrologia 39 (2002). 589…595.

[7] K. Jousten, Wutz-Handbuch Vakuumtechnik, Vieweg Verlag, 2006




� EMBED Word.Picture.8  ���








[image: image64.wmf] 

A

 

buffer

 

inlet pressure

 

Q

 

P

 

L

 

O

 

K

 

F

 

E

 

D

 

C

 

B

 

N

 

S

 

R

 

M

 

H

 

G

 

V4

 

v4

 

V3

 

v3

 

V2

 

v2

 

V1

 

v1

 

_1161072840.unknown

_1191066345.unknown

_1192890094.unknown

_1205223394.xls
Rotor 2 Summary

		s/n MKS 90 792

		Mean Values : of 10 days data collected over 3 runs.

		P Nom (Pa)		P St (Pa)		U St		(ds / d pst)		U St ^2* (ds / d pst)^2		s_1		T		DCR1 Corrected				u _ s				u' _ s						u^2((sNPLk+sNPLk+1+sNPLk+2)/3)						u' ^2((sNPLk+sNPLk+1+sNPLk+1)/3)

		3.E-04		0.0002978777		0.0000009051		-3848.5734504359		0.0000121342		1.1464088504		293.05		0.0000001257				0.0111853655				0.0109107885						0						0

																														0						0

		9.E-04		0.0009052308		0.000002696		-1265.7186707587		0.0000116445		1.1457675704		293.05		0.0000003819				0.0053514779				0.0047766175						0						0

																														0						0

		3.E-03		0.0030208996		0.0000089767		-380.0678392667		0.00001164		1.1481461835		293.04		0.0000012771				0.0028207113				0.0014616401						0						0

																														0						0

		9.E-03		0.0090628034		0.000026925		-126.808603269		0.0000116576		1.1492412585		293.08		0.0000038347				0.002559057				0.0008485117

		2.E-02		0.0201448968		0.000059848		-57.0736641511		0.0000116673		1.1497430109		293.09		0.0000085274				0.0025228088				0.0007286326

		9.E-02		0.0883873457		0.0002281145		-13.0047311584		0.0000088005		1.1494518042		293.11		0.0000374038				0.0022115709				0.0007005659

		2730.2785640991		ball constant r_1

		7.715		ball density

		4.762		ball diameter



Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.



Jul-Sep Unc contribution

		

		NPL mean Jan		NPL mean Apr		NPL Jan/Apr mean						u^2((sNPLk+sNPLk+1)/2)				u' ^2((sNPLk+sNPLk+1)/2)

		1.0579589167		1.0387		1.0483						0.0001139946				0.0001018604

		1.0412221533		1.0239		1.0326						0.0000247052				0.0000130607

		1.0374134677		1.0195		1.0284						0.0000131766				0.0000015365

		1.0366077156		1.0176		1.0271						0.0000122518				0.0000005942

		1.0372641133		1.0175		1.0274						0.0000121596				0.0000004923

		1.0303697003		1.0186		1.0245						0.0000091766				0.0000003761





25 UM Plots

		0.0002971711		0.000296571		0.0002983716		0.0002972647		0.0002978355		0.0002981301

		0.0002969964		0.0002981444		0.0002979865		0.0002968979		0.0002981886		0.0002990624

		0.0002972344		0.0002978199		0.0002976556		0.0002964787		0.000297884		0.0002973803

		0.0009052709		0.000904939

		0.0009048374		0.0009037528		0.0009051038		0.0009030656		0.0009028369		0.0009037614

		0.0009046149		0.0009054913		0.000905018		0.0009037563		0.0009030588		0.0009036293

		0.0030172355		0.0030200352		0.0009055682		0.0009038786		0.0009034615		0.0009035374

		0.0030179128		0.0030181434

		0.0030184816		0.0030187682		0.0030227472		0.003014661		0.0030156696		0.0030147631

		0.0090601787		0.0090595228		0.0030211015		0.0030154766		0.0030141064		0.003015862

		0.0090589585		0.0090576162		0.0030239955		0.0030139904		0.0030155647		0.0030156979

		0.0090595627		0.0090579079

		0.0201331464		0.020137112		0.0090688471		0.0090416092		0.0090417855		0.0090469347

		0.0201363081		0.020144719		0.0090636205		0.0090413705		0.0090407565		0.0090460368

		0.0201328527		0.0201342034		0.0090666308		0.0090450009		0.0090436332		0.0090445829

		0.0888420244		0.0885673645

		0.088402816		0.0881772021		0.0201517294		0.0201061874		0.0201000888		0.020101097

		0.0884735305		0.0884624854		0.0201596103		0.0200970918		0.0201020578		0.0201049489

		0.2981617486		0.2990083559		0.0201611118		0.0200974858		0.0200949652		0.0201002257

		0.2982193997		0.2988603823

		0.2986955637		0.2989675766		0.0882988386		0.0883529051		0.0881190571		0.0878165098

		0.9006369523		0.9014778235		0.0890631432		0.0879028863		0.0880131927		0.0880619701

		0.9011340092		0.9017045514		0.0885654442		0.0881328277		0.0879171436		0.0879931728

		0.900960831		0.9018499962

						0.2991519582		0.2982168768		0.2983309005		0.2987115039

						0.29907786		0.298189282		0.2986609345		0.2988992193

						0.2990839826		0.2984327365		0.298083239		0.2983190094

						0.9013478553		0.8981329039		0.8982545352		0.8976439471

						0.9015690484		0.8980819151		0.8982664202		0.8979309523

						0.9016237647		0.8978296925		0.898117788		0.8976780195



12/07/2005

13/07/2005

14/07/2005

24/08/2005

25/08/2005

26/08/2005

P ref (Pa)

CF

25 UM Plots

1.1608261975

1.1336138154

1.1385521384

1.1455023397

1.1421239831

1.1483105386

1.1590985184

1.1409578203

1.1408864188

1.147385703

1.1474565514

1.1049246458

1.1565260516

1.1476319669

1.1425097032

1.1561839357

1.153949803

1.1100932777

1.1397940671

1.1512033149

1.1367871101

1.1515751638

1.1475476984

1.1539177661

1.1441031288

1.1436620815

1.1376017348

1.1519159992

1.149136189

1.1555766433

1.1424365943

1.1467257331

1.1464584861

1.1507374108

1.1489296905

1.1565257492

1.1419644412

1.147960605

1.1479697321

1.1510200023

1.148068414

1.1505659932

1.1494009917

1.1517207278

1.1483713956

1.1480291693

1.1486747607

1.1499136877

1.1498484334

1.1518011955

1.1489375912

1.1486295896

1.1487415655

1.1496046318

1.1507863401

1.1496276417

1.1490866008

1.1491474682

1.1501837362

1.1497014625

1.1496771334

1.1504793136

1.1507433508

1.1506056493

1.1508873623

1.1505023706

1.1500728452

1.1507433583

1.1508350932

1.1514407318

1.1511650433

1.1511341367

1.1503200792

1.1508266078

1.1510059538

1.1514657442

1.1518412968

1.1515065913

1.1505546473

1.1499580823

1.150735168

1.1499878019

1.1505215099

1.1523757385

1.151492352

1.151331504

1.1509822256

1.1496836752

1.1505463749

1.1521562681

1.1514672923

1.1513492994

1.1451052914

1.1435177069

1.1502192647

1.152136714

1.1518176772

1.1514628648

1.1448032398

1.1441283447

1.1448340797

1.1442816935

1.1511260123

1.1509535241

1.1515517956

1.1496138523

1.1286174919

1.1286107326

1.1508225263

1.1512866856

1.1519344478

1.1504668622

1.1283840482

1.1284484636

1.1510024818

1.1510690966

1.1517333655

1.1502769511

1.1290241995

1.1282652268

1.1449126907

1.1466065948

1.1470475782

1.1465502211

1.1446208866

1.146236387

1.1470880367

1.1464812318

1.1444149314

1.1462220717

1.1469992735

1.1463532571

1.1285602929

1.1311494358

1.131249103

1.1310956722

1.1286221451

1.1311728334

1.1313590733

1.1311719022

1.1285801127

1.1310127701

1.1312071796

1.1311030572



Mean plots 25UM

		0.0002975501		0.0002982052		0.0003083458

		0.0009049552		0.0009055064		0.0009007135

		0.0030198245		0.0030219747		0.0030596987

		0.0090614272		0.0090641796		0.0091162003

		0.0201434215		0.0201463722		0.019927673

		0.0885392054		0.0882354859		0.0897630163



NPL July 2005

NPL August 2005

INRIM July / August 2005

P ref (Pa)

Accomodation Factor

Rotor 1 (25 UM)

1.1467336256

1.1460840753

1.1656952405

1.146054552

1.1454805888

1.1660694647

1.1494284226

1.1468639443

1.1645144022

1.1499338047

1.1485487124

1.1625939097

1.1503784986

1.1491075231

1.1616840472

1.1505391801

1.1483644282

1.1509193263



July 2005

		

		7/12/05																		Mean Values																		Uncertainty Co-efficients								Uncertainty calculations

																																						Eq. 8		Eq. 9		Eq. 10				Eq. 7

		P Nom (Pa)		P St (Pa)		U St		s_2		T		var T		DCR2 Corrected						P St (Pa)		U St		s_2		ST Dev s_2		T		var T		DCR2 Corrected		U RD				(ds / dRDi)		(ds / d pst)		(ds / dT j)				term 1		term 2		term 3		term 4				u s ^2		u s				u ' s ^2		u ' s				u Jan/Apr contribution		s		(ds/dRD)		(ds/dT)		combined

		3.E-04		2.9717E-04		9.0307E-07		1.1608		293.33		0.2411		1.2695E-07						2.9755E-04		9.0417E-07		1.1467		0.0033		293.02		0.2406		1.2564E-07		1.2654E-09				-9127177.58518318		-3853.8872149277		1.9568E-03				0.0000631949		0.0001334005		9.2115E-07		1.2142E-05				2.0966E-04		0.0145				1.9752E-04		0.0140540593						1.073E-05		0.0001334005		9.2115E-07		3.626E-05

		3.E-04		2.9700E-04		9.0256E-07		1.1591		293.42		0.2412		1.2667E-07

		3.E-04		2.9723E-04		9.0325E-07		1.1565		293.44		0.2413		1.2649E-07

		9.E-04		9.0527E-04		2.6961E-06		1.1398		292.93		0.2404		3.7999E-07						9.0496E-04		2.6952E-06		1.1461		0.0021		292.96		0.2405		3.8193E-07		1.2654E-09				-3000715.68055118		-1266.4206324103		1.9560E-03				0.0000251803		0.000014419		9.2007E-07		1.1650E-05				5.2170E-05		0.0072				4.0519E-05		0.0063654826						4.276E-06		0.000014419		9.2007E-07		4.904E-06

		9.E-04		9.0484E-04		2.6949E-06		1.1368		292.80		0.2402		3.7889E-07

		9.E-04		9.0461E-04		2.6942E-06		1.1376		292.77		0.2402		3.7909E-07

		3.E-03		3.0172E-03		8.9658E-06		1.1465		292.82		0.2403		1.2741E-06						3.0198E-03		8.9735E-06		1.1494		0.0005		293.09		0.2407		1.2779E-06		1.2654E-09				-899431.81328029		-380.627623868		1.9609E-03				0.0000016652		0.0000012955		9.2550E-07		1.1666E-05				1.5552E-05		0.0039				3.8862E-06		0.001971337						2.828E-07		0.0000012955		9.2550E-07		6.259E-07

		3.E-03		3.0179E-03		8.9678E-06		1.1480		292.99		0.2405		1.2757E-06

		3.E-03		3.0185E-03		8.9695E-06		1.1481		293.02		0.2406		1.2760E-06

		9.E-03		9.0602E-03		2.6917E-05		1.1487		292.97		0.2405		3.8324E-06						9.0614E-03		2.6921E-05		1.1499		0.0003		293.06		0.2407		3.8365E-06		1.2654E-09				-299732.848019431		-126.9042827947		1.9619E-03				0.0000004784		0.0000001439		9.2631E-07		1.1672E-05				1.3220E-05		0.0036				1.5486E-06		0.0012444151						8.124E-08		0.0000001439		9.2631E-07		2.879E-07

		9.E-03		9.0590E-03		2.6914E-05		1.1487		292.95		0.2405		3.8323E-06

		9.E-03		9.0596E-03		2.6915E-05		1.1502		293.04		0.2406		3.8367E-06

		2.E-02		2.0133E-02		5.9813E-05		1.1497		292.99		0.2405		8.5234E-06						2.0143E-02		5.9844E-05		1.1504		0.0001		293.09		0.2407		8.5314E-06		1.2654E-09				-134839.912430986		-57.109387769		1.9625E-03				0.0000000823		0.0000000291		9.2703E-07		1.1680E-05				1.2719E-05		0.0036				1.0385E-06		0.0010190646						1.398E-08		0.0000000291		9.2703E-07		2.425E-07

		2.E-02		2.0136E-02		5.9822E-05		1.1501		292.95		0.2405		8.5283E-06

		2.E-02		2.0133E-02		5.9812E-05		1.1503		292.92		0.2404		8.5290E-06

		9.E-02		8.8842E-02		2.2912E-04		1.1506		292.98		0.1827		3.7641E-05						8.8539E-02		2.2845E-04		1.1505		0.0002		293.13		0.1829		3.7502E-05		1.2654E-09				-30679.4056168802		-12.994692044		1.9625E-03				0.0000001839		0.0000000015		7.0428E-07		8.8129E-06				9.7026E-06		0.0031				8.8964E-07		0.0009432084						3.122E-08		0.0000000015		7.0428E-07		1.843E-07

		9.E-02		8.8403E-02		2.2815E-04		1.1507		292.99		0.1827		3.7460E-05

		9.E-02		8.8474E-02		2.2831E-04		1.1510		293.07		0.1828		3.7492E-05

		3.E-01		2.9816E-01		7.1852E-04		1.1451		293.14		0.1829		1.2569E-04						2.9880E-01		7.2004E-04		1.1445		0.0002		293.24		0.1830		1.2588E-04		1.2654E-09

		3.E-01		2.9822E-01		7.1866E-04		1.1448		293.15		0.1829		1.2568E-04

		3.E-01		2.9870E-01		7.1979E-04		1.1448		293.20		0.1829		1.2588E-04

		9.E-01		9.0064E-01		2.1567E-03		1.1286		293.25		0.1830		3.7413E-04						9.0137E-01		2.1584E-03		1.1286		0.0001		293.32		0.1831		3.7438E-04		1.2654E-09

		9.E-01		9.0113E-01		2.1578E-03		1.1284		293.29		0.1831		3.7424E-04

		9.E-01		9.0096E-01		2.1574E-03		1.1290		293.28		0.1830		3.7439E-04

		7/13/05																				Molecular mean		1.1489

		P Nom (Pa)		P St (Pa)		U St		s_2		T		var T		DCR2 Corrected								st dev mol mean		0.0009840809

		3.E-04		2.9657E-04		9.0133E-07		1.1336		292.6513		0.2400		1.2387E-07

		3.E-04		2.9814E-04		9.0589E-07		1.1410		292.8071		0.2402		1.2530E-07

		3.E-04		2.9782E-04		9.0495E-07		1.1476		292.8121		0.2402		1.2590E-07

		9.E-04		9.0494E-04		2.6952E-06		1.1512		292.8700		0.2403		3.8369E-07

		9.E-04		9.0375E-04		2.6916E-06		1.1516		292.9856		0.2405		3.8324E-07

		9.E-04		9.0549E-04		2.6968E-06		1.1519		292.9831		0.2405		3.8409E-07

		3.E-03		3.0200E-03		8.9741E-06		1.1507		293.0661		0.2407		1.2795E-06

		3.E-03		3.0181E-03		8.9685E-06		1.1510		293.1037		0.2407		1.2790E-06

		3.E-03		3.0188E-03		8.9704E-06		1.1506		293.1238		0.2408		1.2787E-06

		9.E-03		9.0595E-03		2.6915E-05		1.1499		293.1011		0.2407		3.8354E-06

		9.E-03		9.0576E-03		2.6910E-05		1.1496		293.0232		0.2406		3.8341E-06

		9.E-03		9.0579E-03		2.6910E-05		1.1497		293.0232		0.2406		3.8345E-06

		2.E-02		2.0137E-02		5.9825E-05		1.1505		293.0335		0.2406		8.5303E-06

		2.E-02		2.0145E-02		5.9847E-05		1.1507		293.1441		0.2408		8.5339E-06

		2.E-02		2.0134E-02		5.9816E-05		1.1508		293.1541		0.2408		8.5299E-06

		9.E-02		8.8567E-02		2.2851E-04		1.1500		293.1818		0.1829		3.7492E-05

		9.E-02		8.8177E-02		2.2765E-04		1.1500		293.2095		0.1830		3.7326E-05

		9.E-02		8.8462E-02		2.2828E-04		1.1497		293.2545		0.1830		3.7434E-05

		3.E-01		2.9901E-01		7.2053E-04		1.1435		293.2520		0.1830		1.2585E-04

		3.E-01		2.9886E-01		7.2018E-04		1.1441		293.2645		0.1830		1.2585E-04

		3.E-01		2.9897E-01		7.2043E-04		1.1443		293.3670		0.1832		1.2589E-04

		9.E-01		9.0148E-01		2.1587E-03		1.1286		293.3720		0.1832		3.7440E-04

		9.E-01		9.0170E-01		2.1592E-03		1.1284		293.3820		0.1832		3.7444E-04

		9.E-01		9.0185E-01		2.1596E-03		1.1283		293.3970		0.1832		3.7443E-04

		7/14/05

		P Nom (Pa)		P St (Pa)		U St		s_2		T		var T		DCR2 Corrected

		3.E-04		2.9837E-04		9.0655E-07		1.1386		292.8148		0.2402		1.2513E-07

		3.E-04		2.9799E-04		9.0544E-07		1.1409		292.8977		0.2404		1.2521E-07

		3.E-04		2.9766E-04		9.0448E-07		1.1425		292.9655		0.2405		1.2523E-07

		9.E-04		9.0510E-04		2.6956E-06		1.1475		293.0234		0.2406		3.8244E-07

		9.E-04		9.0502E-04		2.6954E-06		1.1491		293.0988		0.2407		3.8289E-07

		9.E-04		9.0557E-04		2.6970E-06		1.1489		293.1364		0.2408		3.8303E-07

		3.E-03		3.0227E-03		8.9822E-06		1.1494		293.2018		0.2409		1.2789E-06

		3.E-03		3.0211E-03		8.9773E-06		1.1498		293.1917		0.2409		1.2787E-06

		3.E-03		3.0240E-03		8.9859E-06		1.1508		293.2668		0.2410		1.2808E-06

		9.E-03		9.0688E-03		2.6943E-05		1.1507		293.2244		0.2409		3.8413E-06

		9.E-03		9.0636E-03		2.6927E-05		1.1508		293.0713		0.2407		3.8404E-06

		9.E-03		9.0666E-03		2.6936E-05		1.1510		293.1667		0.2408		3.8416E-06

		2.E-02		2.0152E-02		5.9868E-05		1.1505		293.1868		0.2409		8.5346E-06

		2.E-02		2.0160E-02		5.9892E-05		1.1505		293.2119		0.2409		8.5378E-06

		2.E-02		2.0161E-02		5.9896E-05		1.1502		293.2269		0.2409		8.5358E-06

		9.E-02		8.8299E-02		2.2792E-04		1.1511		293.1743		0.1829		3.7417E-05

		9.E-02		8.9063E-02		2.2961E-04		1.1508		293.1241		0.1829		3.7734E-05

		9.E-02		8.8565E-02		2.2851E-04		1.1510		293.2070		0.1830		3.7523E-05

		3.E-01		2.9915E-01		7.2087E-04		1.1449		293.2595		0.1830		1.2606E-04

		3.E-01		2.9908E-01		7.2070E-04		1.1446		293.2644		0.1830		1.2600E-04

		3.E-01		2.9908E-01		7.2071E-04		1.1444		293.2445		0.1830		1.2598E-04

		9.E-01		9.0135E-01		2.1584E-03		1.1286		293.2770		0.1830		3.7439E-04

		9.E-01		9.0157E-01		2.1589E-03		1.1286		293.2745		0.1830		3.7451E-04

		9.E-01		9.0162E-01		2.1590E-03		1.1286		293.3194		0.1831		3.7449E-04





August 2005

		

		8/24/05																		Mean Values																		Uncertainty Co-efficients								Uncertainty calculations

		P Nom (Pa)		P St (Pa)		U St		s_2		T		Var T		DCR2 Corrected						P St (Pa)		U St		s_2		ST Dev s_1		T		var T		DCR1 Corrected		u RD				(ds / dRDi)		(ds / d pst)		(ds / dT j)				term 1		term 2		term 3		term 4				u s ^2		u s				u ' s ^2		u ' s						u loop contribution		s		(ds/dRD)		(ds/dT)		combined

		3.E-04		2.9794E-04		9.0531E-07		1.14289187		292.8550757576		0.2403129153		0.0000001254						2.9821E-04		9.0607E-07		1.1461		0.0018		293.08		0.2407		1.2583E-07		1.7640E-09				-9108197.17680087		-3843.259685944		1.9552E-03				0.000019605		0.0002581395		0.0000009201		1.2126E-05				0.0002907908		0.0171				0.0002786646		0.0166932512								3.329E-06		0.0002581395		9.2011E-07		6.560E-05

		3.E-04		2.9758E-04		9.0424E-07		1.1447709412		292.9681060606		0.2404984534		0.0000001254

		3.E-04		2.9716E-04		9.0302E-07		1.1535491238		292.9957828283		0.2405438954		0.0000001262

		9.E-04		9.0513E-04		2.6957E-06		1.1512881186		293.1541161616		0.2408039429		0.0000003836						9.0551E-04		2.6968E-06		1.1455		0.0019		293.15		0.2408		3.8184E-07		1.7640E-09				-2999893.55833142		-1265.016709107		1.9537E-03				0.000021823		0.0000280028		0.0000009192		0.0000116386				0.0000623836		0.0079				0.0000507449		0.0071235475								3.706E-06		0.0000280028		9.1915E-07		8.157E-06

		9.E-04		9.0582E-04		2.6978E-06		1.1529432154		293.2244949495		0.2409195785		0.0000003844

		9.E-04		9.0594E-04		2.6981E-06		1.1538901584		293.2771969697		0.2410061885		0.0000003847

		3.E-03		3.0215E-03		8.9786E-06		1.149096087		293.2596969697		0.2409774275		0.0000012779						3.0220E-03		8.9799E-06		1.1469		0.0005		293.00		0.2406		1.2762E-06		1.7640E-09				-898658.045749183		-379.5080546654		1.9571E-03				0.0000012251		0.0000025129		0.0000009214		0.000011614				0.0000162734		0.0040				0.0000046594		0.0021585639								2.080E-07		0.0000025129		9.2137E-07		9.106E-07

		3.E-03		3.0224E-03		8.9810E-06		1.1491763713		293.2897984798		0.2410269		0.0000012783

		3.E-03		3.0209E-03		8.9766E-06		1.1470077709		292.9506565657		0.2404698056		0.000001276

		9.E-03		9.0623E-03		2.6923E-05		1.1479835497		293.0084343434		0.2405646692		0.0000038307						9.0642E-03		2.6929E-05		1.1485		0.0001		293.09		0.2407		3.8329E-06		1.7640E-09				-299655.868863713		-126.7129237433		1.9594E-03				0.0000001278		0.0000002794		0.0000009241		0.0000116435				0.0000129749		0.0036				0.0000013313		0.001153828								2.171E-08		0.0000002794		9.2408E-07		3.063E-07

		9.E-03		9.0620E-03		2.6923E-05		1.1488167291		293.1088888889		0.2407296471		0.0000038327

		9.E-03		9.0657E-03		2.6933E-05		1.1488416845		293.2043686869		0.2408865073		0.0000038337

		2.E-02		2.0152E-02		5.9869E-05		1.149749605		293.0763636364		0.2406762242		0.0000085307						2.0146E-02		5.9852E-05		1.1491		0.0001		293.08		0.2407		8.5233E-06		1.7640E-09				-134818.951715168		-57.0379405332		1.9604E-03				0.0000001036		0.0000000566		0.000000925		0.0000116544				0.0000127396		0.0036				0.0000010851		0.0010416954								1.759E-08		0.0000000566		9.2498E-07		2.498E-07

		2.E-02		2.0143E-02		5.9842E-05		1.1495306348		293.101489899		0.2407174937		0.0000085248

		2.E-02		2.0143E-02		5.9844E-05		1.1495111253		293.0335606061		0.2406059291		0.0000085258

		9.E-02		8.8555E-02		2.2849E-04		1.1483306316		293.1868434343		0.1829338146		0.0000374331						8.8235E-02		2.2778E-04		1.1484		0.0003		293.09		0.1828		3.7306E-05		1.7640E-09				-30782.6024039824		-13.0147702728		1.9591E-03				0.000000369		0.0000000029		0.0000007016		0.0000087881				0.0000098616		0.0031				0.0000010735		0.0010361121								6.265E-08		0.0000000029		7.0162E-07		1.918E-07

		9.E-02		8.8104E-02		2.2749E-04		1.148663034		293.2269949495		0.1829839231		0.0000372506

		9.E-02		8.8334E-02		2.2800E-04		1.1484459408		293.216969697		0.1829714111		0.0000373416

		3.E-01		2.9890E-01		7.2027E-04		1.1439936085		293.1290151515		0.1828616579		0.0001258827						2.9911E-01		7.2077E-04		1.1440		0.0001		292.94		0.1826		1.2601E-04		1.7640E-09

		3.E-01		2.9887E-01		7.2020E-04		1.1436242444		293.0159343434		0.1827205994		0.0001258547

		3.E-01		2.9911E-01		7.2078E-04		1.1436099617		292.8425252525		0.1825043925		0.0001259932

		9.E-01		9.0018E-01		2.1556E-03		1.1285716746		292.975530303		0.1826702121		0.0003741053						9.0004E-01		2.1552E-03		1.1286		0.0000		293.03		0.1827		3.7402E-04		1.7640E-09

		9.E-01		9.0013E-01		2.1555E-03		1.1285950189		293.0359343434		0.1827455437		0.0003740532

		9.E-01		8.9988E-01		2.1548E-03		1.1284353204		293.0736111111		0.1827925394		0.0003738712

																						molecular mean		1.1475

		8/25/05																				st dev molecular mean		0.0016499007

		P Nom (Pa)		P St (Pa)		U St		s_2		T		Var T		DCR2 Corrected

		3.E-04		2.9852E-04		9.0697E-07		1.1395212122		293.053510101		0.2406386907		0.0000001252

		3.E-04		2.9887E-04		9.0800E-07		1.1448416282		293.2294444444		0.2409277118		0.0000001259

		3.E-04		2.9856E-04		9.0711E-07		1.1513200824		293.1641414141		0.2408204132		0.0000001265

		9.E-04		9.0490E-04		2.6950E-06		1.1414958476		293.0132575758		0.2405725892		0.0000003803

		9.E-04		9.0512E-04		2.6957E-06		1.139833111		293.0233333333		0.2405891344		0.0000003799

		9.E-04		9.0553E-04		2.6969E-06		1.1393620339		293.0735858586		0.240671662		0.0000003799

		3.E-03		3.0226E-03		8.9816E-06		1.1457543875		293.1364393939		0.2407749035		0.0000012749

		3.E-03		3.0210E-03		8.9770E-06		1.1463192929		292.8073232323		0.2402345516		0.0000012756

		3.E-03		3.0225E-03		8.9813E-06		1.1464679629		292.9605050505		0.2404859742		0.000001276

		9.E-03		9.0624E-03		2.6924E-05		1.1482646206		293.0032828283		0.2405562103		0.0000038318

		9.E-03		9.0614E-03		2.6921E-05		1.1485416689		293.1063131313		0.2407254162		0.0000038316

		9.E-03		9.0643E-03		2.6929E-05		1.1492163813		293.2143434343		0.2409028974		0.0000038344

		2.E-02		2.0146E-02		5.9851E-05		1.1488682317		293.2571212121		0.2409731944		0.0000085189

		2.E-02		2.0148E-02		5.9857E-05		1.1488432291		293.2871727173		0.2410225843		0.0000085191

		2.E-02		2.0141E-02		5.9836E-05		1.1491928155		293.0988383838		0.2407131385		0.0000085214

		9.E-02		8.8320E-02		2.2797E-04		1.1489275398		293.0962626263		0.1828207963		0.0000373592

		9.E-02		8.8214E-02		2.2773E-04		1.14930932		293.0837121212		0.1828051398		0.0000373275

		9.E-02		8.8118E-02		2.2752E-04		1.1491086959		293.1415151515		0.1828772538		0.0000372766

		3.E-01		2.9901E-01		7.2054E-04		1.144433587		292.8926767677		0.1825669084		0.0001260301

		3.E-01		2.9934E-01		7.2133E-04		1.1444739533		292.9379545455		0.1826233582		0.0001261642

		3.E-01		2.9876E-01		7.1995E-04		1.1443853924		292.9555808081		0.182645336		0.0001259067

		9.E-01		9.0031E-01		2.1559E-03		1.1286711148		292.9681818182		0.1826610487		0.0003741936

		9.E-01		9.0032E-01		2.1559E-03		1.1287808344		293.033510101		0.1827425201		0.0003741932

		9.E-01		9.0017E-01		2.1555E-03		1.1286292868		293.0863131313		0.1828083844		0.0003740474

		8/26/05

		P Nom (Pa)		P St (Pa)		U St		s_2		T		Var T		DCR2 Corrected

		3.E-04		2.9881E-04		9.0783E-07		1.1456936693		293.3266751675		0.2410875147		0.000000126

		3.E-04		2.9975E-04		9.1055E-07		1.102406648		292.7344191919		0.2401149378		0.0000001217

		3.E-04		2.9806E-04		9.0565E-07		1.1075635011		292.8424494949		0.2402921939		0.0000001216

		9.E-04		9.0583E-04		2.6978E-06		1.1410558055		293.1540909091		0.2408039014		0.0000003805

		9.E-04		9.0569E-04		2.6974E-06		1.1441124754		293.2018181818		0.2408823165		0.0000003814

		9.E-04		9.0560E-04		2.6971E-06		1.1453445331		293.2444459446		0.240952364		0.0000003818

		3.E-03		3.0216E-03		8.9789E-06		1.1454129411		292.8047979798		0.2402304079		0.0000012749

		3.E-03		3.0228E-03		8.9822E-06		1.1460119932		292.8575252525		0.2403169354		0.0000012759

		3.E-03		3.0226E-03		8.9817E-06		1.1465286916		292.937979798		0.2404489945		0.0000012762

		9.E-03		9.0676E-03		2.6939E-05		1.1478805063		293.0435606061		0.2406223511		0.0000038324

		9.E-03		9.0667E-03		2.6937E-05		1.1485108327		293.1339646465		0.2407708382		0.0000038335

		9.E-03		9.0652E-03		2.6932E-05		1.1488824385		292.9831818182		0.2405232055		0.0000038352

		2.E-02		2.0147E-02		5.9854E-05		1.1487077502		292.9028030303		0.2403912504		0.0000085233

		2.E-02		2.0151E-02		5.9866E-05		1.1487255051		293.0158080808		0.2405767773		0.0000085234

		2.E-02		2.0146E-02		5.9852E-05		1.1488388117		292.9907828283		0.2405356857		0.0000085226

		9.E-02		8.8017E-02		2.2729E-04		1.1469940128		293.0133838384		0.1827174185		0.0000371735

		9.E-02		8.8263E-02		2.2784E-04		1.1478450788		292.9731565657		0.1826672521		0.0000373076

		9.E-02		8.8194E-02		2.2769E-04		1.1476556005		292.835		0.1824950129		0.0000372811

		3.E-01		2.9939E-01		7.2145E-04		1.1439373634		292.8626767677		0.1825295109		0.0001261426

		3.E-01		2.9958E-01		7.2189E-04		1.1438685313		292.8702272727		0.1825389229		0.0001262127

		3.E-01		2.9900E-01		7.2051E-04		1.1437408481		292.920479798		0.1826015706		0.0001259428

		9.E-01		8.9969E-01		2.1544E-03		1.1285180335		292.9581565657		0.1826485477		0.0003738949

		9.E-01		8.9998E-01		2.1551E-03		1.1285940899		293.0561111111		0.1827707102		0.0003739772

		9.E-01		8.9973E-01		2.1545E-03		1.1285254018		293.0587121212		0.1827739546		0.0003738474






_1234093356.unknown

_1234254760.xls
Rotor 1 Summary

		s/n 23 UM

		Mean Values : of 6 days data collected over 2 runs.

		P Nom (Pa)		P St (Pa)		U St		(ds / d pst)		U St ^2* (ds / d pst)^2		s_1		T		DCR1 Corrected				u _ s				u' _ s						u^2((sNPLk+sNPLk+1+sNPLk+2)/3)						u' ^2((sNPLk+sNPLk+1+sNPLk+1)/3)

		3.E-04		0.0002979551		0.0000009053		-3712.61		0.0000112976		1.1062		293.0173		1.2083E-07				0.0280084436				0.0279074206						0						0

																														0						0

		9.E-04		0.0009052308		0.000002696		-1227.27		0.0000109479		1.1110		293.0536		3.6867E-07				0.0095492431				0.0092581915						0						0

																														0						0

		3.E-03		0.0030208996		0.0000089767		-368.36		0.0000109337		1.1128		293.0442		1.2323E-06				0.0037601728				0.0029448351						0						0

																														0						0

		9.E-03		0.0090628034		0.000026925		-122.84		0.000010939		1.1133		293.0759		3.6984E-06				0.0026601181				0.0012675619

		2.E-02		0.0201448968		0.000059848		-55.28		0.0000109441		1.1135		293.0875		8.2228E-06				0.0025021645				0.0008881277

		9.E-02		0.0883873457		0.0002281145		-12.60		0.0000082561		1.1133		293.1089		3.6070E-05				0.0021552823				0.0007191641

		2742.2560548405		ball constant r_1

		8.2		ball density

		4.5		ball diameter



Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.

Paul Carroll:
Combining the uncertainties - use formula pg.



23UM Plots

		0.0002983716		0.000296571		0.0002971711		0.0002988111		0.0002979437		0.0002985158		0.0003057308

		0.0002979865		0.0002981444		0.0002969964		0.0002997455		0.0002975761		0.0002988697		0.0008999502

		0.0002976556		0.0002978199		0.0002972344		0.0002980596		0.0002971559		0.0002985643		0.0030567689

						0.0009052709								0.00911423

		0.0009051038		0.000904939		0.0009048374		0.0009058257		0.0009051283		0.0009048991		0.0199370956

		0.000905018		0.0009037528		0.0009046149		0.0009056933		0.0009058206		0.0009051214		0.0897898371

		0.0009055682		0.0009054913		0.0030172355		0.0009056012		0.0009059431		0.0009055251

						0.0030179128

		0.0030227472		0.0030200352		0.0030184816		0.0030216491		0.0030215467		0.0030225577

		0.0030211015		0.0030181434		0.0090601787		0.0030227505		0.0030223643		0.0030209909

		0.0030239955		0.0030187682		0.0090589585		0.003022586		0.0030208746		0.0030224525

						0.0090595627

		0.0090688471		0.0090595228		0.0201331464		0.0090675987		0.0090622611		0.0090624378

		0.0090636205		0.0090576162		0.0201363081		0.0090666988		0.0090620219		0.0090614064

		0.0090666308		0.0090579079		0.0201328527		0.0090652415		0.0090656605		0.0090642897

						0.0888420244

		0.0201517294		0.020137112		0.088402816		0.0201470097		0.0201521118		0.0201459993

		0.0201596103		0.020144719		0.0884735305		0.0201508705		0.0201429954		0.0201479727

		0.0201611118		0.0201342034		0.2981617486		0.0201461365		0.0201433903		0.020140864

						0.2982193997

		0.0882988386		0.0885673645		0.2986955637		0.0880170908		0.0885547113		0.0883203291

		0.0890631432		0.0881772021		0.9006369523		0.0882631117		0.0881036645		0.0882142229

		0.0885654442		0.0884624854		0.9011340092		0.0881941572		0.0883341311		0.0881179544

						0.900960831

		0.2991519582		0.2990083559				0.2993937884		0.2988980315		0.2990123156

		0.29907786		0.2988603823				0.2995819325		0.2988703737		0.2993431034

		0.2990839826		0.2989675766				0.2990003973				0.2987640885

		0.9013478553		0.9014778235				0.8996942481				0.9003062309

		0.9015690484		0.9017045514				0.8999819089				0.9003181431

		0.9016237647		0.9018499962				0.8997283984				0.9001691714



14-Jul-05

13-Jul-05

12-Jul-05

26-Aug-05

24-Aug-05

25-Aug-05

P ref (Pa)

cf

23UM (rotor 1)

1.097832725

1.1162971507

1.1010162438

1.1104484577

1.0984934696

1.1109826915

1.1190836791

1.0973131755

1.1189914746

1.1016150631

1.1030308965

1.0992883022

1.1106444049

1.1231471828

1.0996735206

1.1198758277

1.1026963542

1.1061873723

1.1035653332

1.1133906422

1.1244937747

1.1073390542

1.1227810384

1.1093428222

1.1166903642

1.1070611451

1.1109999153

1.1088300861

1.1105789026

1.1217192833

1.1098569475

1.1171647174

1.1073806975

1.1111972233

1.109872783

1.1115465969

1.1121149322

1.1100015289

1.1155596981

1.11020458

1.1114948713

1.1105490244

1.1119425288

1.110951941

1.1126135374

1.1148956324

1.1104649285

1.1121421116

1.1113435358

1.1123075183

1.1132960666

1.1155085594

1.1112914199

1.1128161701

1.1118064655

1.11364121

1.1142783245

1.1154601595

1.1114729987

1.1129946577

1.1115989154

1.1134480624

1.1129572414

1.1147474922

1.1142980227

1.1121201739

1.1125886186

1.1121717483

1.1130467318

1.1147536389

1.1137467192

1.1125178801

1.1129946219

1.1128005203

1.1131095377

1.1150253364

1.1140813523

1.1127312076

1.1135481526

1.1125946665

1.1134262951

1.1131991457

1.1144547759

1.1144556955

1.1133709753

1.1131543494

1.1137361849

1.1129283625

1.1143927916

1.1147017319

1.1135732883

1.1131429003

1.1135410837

1.1128379606

1.1140823962

1.1146811133

1.1083083581

1.113288637

1.113562804

1.1133781316

1.1080233438

1.1151291296

1.113966413

1.1079454027

1.1112277282

1.1124738723

1.1131681886

1.1148650881

1.1139779916

1.0939449116

1.1120715751

1.1128325793

1.1135504031

1.1150102278

1.1136829495

1.0937441121

1.1119095156

1.1126260163

1.1133186576

1.0941178867

1.1094312584

1.1077593453

1.1087310156

1.1083136374

1.1092046652

1.109159169

1.1085337239

1.1086398785

1.1083012734

1.1092368733

1.10886566

1.1084369437

1.1085368727

1.1091250294

1.0947108795

1.094600937

1.0950362453

1.0949577372

1.0946340109

1.0946642488

1.0950402894

1.0949344427

1.0945239369

1.0942800266

1.0947937779

1.0949720391



Mean Plots 23UM

		0.0002975501		0.0002983602		0.0003057308

		0.0009049552		0.0009055064		0.0008999502

		0.0030198245		0.0030219747		0.0030567689

		0.0090614272		0.0090641796		0.00911423

		0.0201434215		0.0201463722		0.0199370956

		0.0885392054		0.0882354859		0.0897898371



NPL July 2005

NPL August 2005

INRIM July / August 2005

P Ref (Pa)

Accomodation Factor

Rotor 2 (23 UM)

1.1061457261

1.10622573

1.1190836791

1.1111552195

1.1107791035

1.1231471828

1.1130748588

1.1124554052

1.1244937747

1.1135971357

1.1129200992

1.1227810384

1.1137930851

1.1132856015

1.1217192833

1.1140861343

1.1125753929

1.1121149322



July 2005

		

		7/12/05																		Mean Values																		Uncertainty Co-efficients								Uncertainty calculations

																																						Eq. 8		Eq. 9		Eq. 10				Eq. 7

		P Nom (Pa)		P St (Pa)		U St		s_1		T		var T		DCR1 Corrected						P St (Pa)		U St		s_1		ST Dev s_1		T		var T		DCR1 Corrected		U RD				(ds / dRDi)		(ds / d pst)		(ds / dT j)				term 1		term 2		term 3		term 4				u s ^2		u s				u ' s ^2		u ' s				u Jul/Aug contribution		s		(ds/dRD)		(ds/dT)		combined

		3.E-04		2.9717E-04		9.0307E-07		1.1010		293.33		0.2411		1.1989E-07						2.9755E-04		9.0417E-07		1.1061		0.0031		293.02		0.2406		1.2066E-07		3.6575E-09				-9167217.70653174		-3717.5148986653		1.8875E-03				0.0000576069		0.0011242283		8.5711E-07		1.1298E-05				1.1940E-03		0.0346				1.1827E-03		0.0343902945						9.782E-06		0.0011242283		8.5711E-07		2.837E-04

		3.E-04		2.9700E-04		9.0256E-07		1.1016		293.42		0.2412		1.1986E-07

		3.E-04		2.9723E-04		9.0325E-07		1.1027		293.44		0.2413		1.2007E-07

		9.E-04		9.0527E-04		2.6961E-06		1.1073		292.93		0.2404		3.6756E-07						9.0496E-04		2.6952E-06		1.1112		0.0014		292.96		0.2405		3.6868E-07		3.6575E-09				-3013879.55501954		-1227.8558059657		1.8965E-03				0.000011289		0.0001215154		8.6489E-07		1.0952E-05				1.4462E-04		0.0120				1.3367E-04		0.0115615429						1.917E-06		0.0001215154		8.6489E-07		3.107E-05

		9.E-04		9.0484E-04		2.6949E-06		1.1071		292.80		0.2402		3.6737E-07

		9.E-04		9.0461E-04		2.6942E-06		1.1074		292.77		0.2402		3.6741E-07

		3.E-03		3.0172E-03		8.9658E-06		1.1102		292.82		0.2403		1.2285E-06						3.0198E-03		8.9735E-06		1.1131		0.0007		293.09		0.2407		1.2321E-06		3.6575E-09				-903377.541147682		-368.5892914358		1.8989E-03				0.0000029604		0.0000109174		8.6788E-07		1.0940E-05				2.5685E-05		0.0051				1.4746E-05		0.0038400044						5.027E-07		0.0000109174		8.6788E-07		3.072E-06

		3.E-03		3.0179E-03		8.9678E-06		1.1110		292.99		0.2405		1.2292E-06

		3.E-03		3.0185E-03		8.9695E-06		1.1105		293.02		0.2406		1.2288E-06

		9.E-03		9.0602E-03		2.6917E-05		1.1113		292.97		0.2405		3.6915E-06						9.0614E-03		2.6921E-05		1.1136		0.0005		293.06		0.2407		3.6991E-06		3.6575E-09				-301047.749531409		-122.8942443656		1.8999E-03				0.000001283		0.0000012124		8.6870E-07		1.0946E-05				1.4310E-05		0.0038				3.3641E-06		0.0018341487						2.179E-07		0.0000012124		8.6870E-07		5.747E-07

		9.E-03		9.0590E-03		2.6914E-05		1.1115		292.95		0.2405		3.6917E-06

		9.E-03		9.0596E-03		2.6915E-05		1.1130		293.04		0.2406		3.6964E-06

		2.E-02		2.0133E-02		5.9813E-05		1.1121		292.99		0.2405		8.2090E-06						2.0143E-02		5.9844E-05		1.1138		0.0003		293.09		0.2407		8.2240E-06		3.6575E-09				-135431.443208817		-55.293146851		1.9001E-03				0.0000006941		0.0000002454		8.6900E-07		1.0949E-05				1.2758E-05		0.0036				1.8085E-06		0.0013448083						1.179E-07		0.0000002454		8.6900E-07		3.081E-07

		2.E-02		2.0136E-02		5.9822E-05		1.1125		292.95		0.2405		8.2137E-06

		2.E-02		2.0133E-02		5.9812E-05		1.1127		292.92		0.2404		8.2142E-06

		9.E-02		8.8842E-02		2.2912E-04		1.1132		292.98		0.1827		3.6260E-05						8.8539E-02		2.2845E-04		1.1141		0.0002		293.13		0.1829		3.6155E-05		3.6575E-09				-30813.9934576817		-12.5829729708		1.9003E-03				0.0000003516		0.0000000127		6.6036E-07		8.2633E-06				9.2879E-06		0.0030				1.0246E-06		0.0010122392						5.970E-08		0.0000000127		6.6036E-07		1.832E-07

		9.E-02		8.8403E-02		2.2815E-04		1.1134		292.99		0.1827		3.6085E-05

		9.E-02		8.8474E-02		2.2831E-04		1.1136		293.07		0.1828		3.6115E-05

		3.E-01		2.9816E-01		7.1852E-04		1.1083		293.14		0.1829		1.2112E-04						2.9880E-01		7.2004E-04		1.1085		0.0002		293.24		0.1830		1.2138E-04		3.6575E-09

		3.E-01		2.9822E-01		7.1866E-04		1.1080		293.15		0.1829		1.2111E-04

		3.E-01		2.9870E-01		7.1979E-04		1.1079		293.20		0.1829		1.2129E-04

		9.E-01		9.0064E-01		2.1567E-03		1.0939		293.25		0.1830		3.6105E-04						9.0137E-01		2.1584E-03		1.0944		0.0001		293.32		0.1831		3.6144E-04		3.6575E-09

		9.E-01		9.0113E-01		2.1578E-03		1.0937		293.29		0.1831		3.6116E-04

		9.E-01		9.0096E-01		2.1574E-03		1.0941		293.28		0.1830		3.6123E-04

																												Unc. P st relative - July

		7/13/05																				Molecular mean		1.1129				0.0030387112

																												0.0029782717

		P Nom (Pa)		P St (Pa)		U St		s_1		T		var T		DCR1 Corrected								st dev mol mean		0.000602656				0.0029715282

																												0.002970936

		3.E-04		2.9657E-04		9.0133E-07		1.1163		292.6513		0.2400		1.2145E-07														0.002970877

		3.E-04		2.9814E-04		9.0589E-07		1.1190		292.8071		0.2402		1.2235E-07														0.0025802285

		3.E-04		2.9782E-04		9.0495E-07		1.1199		292.8121		0.2402		1.2231E-07														0.0024097559

																												2.3946E-03

		9.E-04		9.0494E-04		2.6952E-06		1.1167		292.8700		0.2403		3.7056E-07

		9.E-04		9.0375E-04		2.6916E-06		1.1172		292.9856		0.2405		3.7016E-07

		9.E-04		9.0549E-04		2.6968E-06		1.1156		292.9831		0.2405		3.7034E-07

		3.E-03		3.0200E-03		8.9741E-06		1.1149		293.0661		0.2407		1.2343E-06

		3.E-03		3.0181E-03		8.9685E-06		1.1155		293.1037		0.2407		1.2341E-06

		3.E-03		3.0188E-03		8.9704E-06		1.1155		293.1238		0.2408		1.2343E-06

		9.E-03		9.0595E-03		2.6915E-05		1.1143		293.1011		0.2407		3.7004E-06

		9.E-03		9.0576E-03		2.6910E-05		1.1137		293.0232		0.2406		3.6983E-06

		9.E-03		9.0579E-03		2.6910E-05		1.1141		293.0232		0.2406		3.6995E-06

		2.E-02		2.0137E-02		5.9825E-05		1.1145		293.0335		0.2406		8.2271E-06

		2.E-02		2.0145E-02		5.9847E-05		1.1147		293.1441		0.2408		8.2305E-06

		2.E-02		2.0134E-02		5.9816E-05		1.1147		293.1541		0.2408		8.2259E-06

		9.E-02		8.8567E-02		2.2851E-04		1.1140		293.1818		0.1829		3.6160E-05

		9.E-02		8.8177E-02		2.2765E-04		1.1140		293.2095		0.1830		3.5999E-05

		9.E-02		8.8462E-02		2.2828E-04		1.1137		293.2545		0.1830		3.6103E-05

		3.E-01		2.9901E-01		7.2053E-04		1.1078		293.2520		0.1830		1.2138E-04

		3.E-01		2.9886E-01		7.2018E-04		1.1085		293.2645		0.1830		1.2140E-04

		3.E-01		2.9897E-01		7.2043E-04		1.1084		293.3670		0.1832		1.2142E-04

		9.E-01		9.0148E-01		2.1587E-03		1.0946		293.3720		0.1832		3.6153E-04

		9.E-01		9.0170E-01		2.1592E-03		1.0947		293.3820		0.1832		3.6164E-04

		9.E-01		9.0185E-01		2.1596E-03		1.0943		293.3970		0.1832		3.6156E-04

		7/14/05

		P Nom (Pa)		P St (Pa)		U St		s_1		T		var T		DCR1 Corrected

		3.E-04		2.9837E-04		9.0655E-07		1.0978		292.8148		0.2402		1.2013E-07

		3.E-04		2.9799E-04		9.0544E-07		1.0973		292.8977		0.2404		1.1990E-07

		3.E-04		2.9766E-04		9.0448E-07		1.0997		292.9655		0.2405		1.2001E-07

		9.E-04		9.0510E-04		2.6956E-06		1.1093		293.0234		0.2406		3.6810E-07

		9.E-04		9.0502E-04		2.6954E-06		1.1099		293.0988		0.2407		3.6818E-07

		9.E-04		9.0557E-04		2.6970E-06		1.1100		293.1364		0.2408		3.6843E-07

		3.E-03		3.0227E-03		8.9822E-06		1.1126		293.2018		0.2409		1.2326E-06

		3.E-03		3.0211E-03		8.9773E-06		1.1133		293.1917		0.2409		1.2327E-06

		3.E-03		3.0240E-03		8.9859E-06		1.1143		293.2668		0.2410		1.2348E-06

		9.E-03		9.0688E-03		2.6943E-05		1.1147		293.2244		0.2409		3.7049E-06

		9.E-03		9.0636E-03		2.6927E-05		1.1148		293.0713		0.2407		3.7038E-06

		9.E-03		9.0666E-03		2.6936E-05		1.1150		293.1667		0.2408		3.7053E-06

		2.E-02		2.0152E-02		5.9868E-05		1.1145		293.1868		0.2409		8.2310E-06

		2.E-02		2.0160E-02		5.9892E-05		1.1144		293.2119		0.2409		8.2334E-06

		2.E-02		2.0161E-02		5.9896E-05		1.1141		293.2269		0.2409		8.2315E-06

		9.E-02		8.8299E-02		2.2792E-04		1.1151		293.1743		0.1829		3.6088E-05

		9.E-02		8.9063E-02		2.2961E-04		1.1149		293.1241		0.1829		3.6395E-05

		9.E-02		8.8565E-02		2.2851E-04		1.1150		293.2070		0.1830		3.6191E-05

		3.E-01		2.9915E-01		7.2087E-04		1.1094		293.2595		0.1830		1.2162E-04

		3.E-01		2.9908E-01		7.2070E-04		1.1092		293.2644		0.1830		1.2156E-04

		3.E-01		2.9908E-01		7.2071E-04		1.1089		293.2445		0.1830		1.2154E-04

		9.E-01		9.0135E-01		2.1584E-03		1.0947		293.2770		0.1830		3.6158E-04

		9.E-01		9.0157E-01		2.1589E-03		1.0946		293.2745		0.1830		3.6164E-04

		9.E-01		9.0162E-01		2.1590E-03		1.0945		293.3194		0.1831		3.6160E-04





Jul-Sep Unc contribution

		

		NPL mean Jan		NPL mean Apr		NPL Jan/Apr mean						u^2((sNPLk+sNPLk+1)/2)				u' ^2((sNPLk+sNPLk+1)/2)

		1.0579589167		1.0387		1.0483						0.0012520366				0.001240739

		1.0412221533		1.0239		1.0326						0.000145907				0.0001349591

		1.0374134677		1.0195		1.0284						0.0000233443				0.0000124106

		1.0366077156		1.0176		1.0271						0.0000125553				0.0000016162

		1.0372641133		1.0175		1.0274						0.0000114644				0.0000005203

		1.0303697003		1.0186		1.0245						0.0000084657				0.0000002096





August 2005

		

		8/24/05																		Mean Values																		Uncertainty Co-efficients								Uncertainty calculations

		P Nom (Pa)		P St (Pa)		U St		s_1		T		Var T		DCR1 Corrected						P St (Pa)		U St		s_1		ST Dev s_1		T		var T		DCR1 Corrected		u RD				(ds / dRDi)		(ds / d pst)		(ds / dT j)				term 1		term 2		term 3		term 4				u s ^2		u s				u ' s ^2		u ' s						u loop contribution		s		(ds/dRD)		(ds/dT)		combined

		3.E-04		2.9794E-04		9.0531E-07		1.0984934696		292.8550757576		0.2403129153		0.00000012						2.9836E-04		9.0652E-07		1.1062		0.0018124671		293.01884494		0.2406		1.2100E-07		4.7833E-09				-9142375.7716775		-3707.70		1.8876E-03				0.0000193452		0.0019124017		8.5724E-07		1.1297E-05				1.9439E-03		0.0441				0.0019326042		0.0439613939								3.285E-06		0.0019124017		1.9124E-03		9.570E-04

		3.E-04		2.9758E-04		9.0424E-07		1.0992883022		292.9681060606		0.2404984534		0.0000001199

		3.E-04		2.9716E-04		9.0302E-07		1.1035653332		292.9957828283		0.2405438954		0.0000001202

		9.E-04		9.0513E-04		2.6957E-06		1.1088300861		293.1541161616		0.2408039429		0.0000003679						9.0551E-04		2.6968E-06		1.1108		0.000320408		293.1518155427		0.2408		3.6866E-07		4.7833E-09				-3013053.82622227		-1226.69		1.8945E-03				0.0000006046		0.0002077183		8.6431E-07		0.0000109441				0.0002201313		0.0148				0.0002091872		0.0144633042								1.027E-07		0.0002077183		2.0772E-04		1.039E-04

		9.E-04		9.0582E-04		2.6978E-06		1.109872783		293.2244949495		0.2409195785		0.0000003684

		9.E-04		9.0594E-04		2.6981E-06		1.1105490244		293.2771969697		0.2410061885		0.0000003687

		3.E-03		3.0215E-03		8.9786E-06		1.1113435358		293.2596969697		0.2409774275		0.0000012306						3.0220E-03		8.9799E-06		1.1125		0.0002719125		293.0005247469		0.2406		1.2325E-06		4.7833E-09				-902600.379166807		-368.12		1.8984E-03				0.0000004354		0.0000186403		8.6692E-07		0.0000109276				0.0000308702		0.0056				0.0000199426		0.0044657118								7.394E-08		0.0000186403		1.8640E-05		9.339E-06

		3.E-03		3.0224E-03		8.9810E-06		1.1118064655		293.2897984798		0.2410269		0.0000012313

		3.E-03		3.0209E-03		8.9766E-06		1.1115989154		292.9506565657		0.2404698056		0.0000012312

		9.E-03		9.0623E-03		2.6923E-05		1.1121717483		293.0084343434		0.2405646692		0.000003695						9.0642E-03		2.6929E-05		1.1129		0.0001442546		293.0895931538		0.2407		3.6978E-06		4.7833E-09				-300970.432674939		-122.78		1.8986E-03				0.0000001225		0.0000020726		8.6764E-07		0.0000109324				0.0000139951		0.0037				0.0000030628		0.0017500719								2.081E-08		0.0000020726		2.0726E-06		1.041E-06

		9.E-03		9.0620E-03		2.6923E-05		1.1128005203		293.1088888889		0.2407296471		0.0000036964

		9.E-03		9.0657E-03		2.6933E-05		1.1125946665		293.2043686869		0.2408865073		0.0000036966

		2.E-02		2.0152E-02		5.9869E-05		1.1137361849		293.0763636364		0.2406762242		0.0000082274						2.0146E-02		5.9852E-05		1.1133		0.0000999515		293.084882266		0.2407		8.2216E-06		4.7833E-09				-135410.390540191		-55.26		1.8993E-03				0.0000000588		0.0000004195		8.6821E-07		0.0000109391				0.0000122857		0.0035				0.0000013466		0.0011604197								9.990E-09		0.0000004195		4.1953E-07		2.123E-07

		2.E-02		2.0143E-02		5.9842E-05		1.1135410837		293.101489899		0.2407174937		0.0000082218

		2.E-02		2.0143E-02		5.9844E-05		1.113562804		293.0335606061		0.2406059291		0.0000082231

		9.E-02		8.8555E-02		2.2849E-04		1.1124738723		293.1868434343		0.1829338146		0.0000361058						8.8235E-02		2.2778E-04		1.1126		0.0002485215		293.0859820426		0.1828		3.5985E-05		4.7833E-09				-30917.6429599677		-12.61		1.8980E-03				0.0000003637		0.0000000219		6.5857E-07		0.0000082489				0.000009293		0.0030				0.0000010442		0.0010218414								6.176E-08		0.0000000219		2.1871E-08		2.638E-08

		9.E-02		8.8104E-02		2.2749E-04		1.1128325793		293.2269949495		0.1829839231		0.000035931

		9.E-02		8.8334E-02		2.2800E-04		1.1126260163		293.216969697		0.1829714111		0.000036019

		3.E-01		2.9890E-01		7.2027E-04		1.1083136374		293.1290151515		0.1828616579		0.0001214239						2.9911E-01		7.2077E-04		1.1087		0.0001456553		292.9363411897		0.1826		1.2159E-04		4.7833E-09				-9118.212267644		-3.71		1.8924E-03				0.0000001249		0.0000000019		6.5397E-07		0.0000071375				0.0000079183		0.0028				0.0000007808		0.000883632								2.122E-08		0.0000000019		1.9023E-09		6.255E-09

		3.E-01		2.9887E-01		7.2020E-04		1.1083012734		293.0159343434		0.1827205994		0.0001214347

		3.E-01		2.9911E-01		7.2078E-04		1.1080153547		292.8425252525		0.1825043925		0.0001215385

		9.E-01		9.0018E-01		2.1556E-03		1.0950439012		292.975530303		0.1826702121		0.0003614058						9.0004E-01		2.1552E-03		1.0949		0.0000295586		2.9303E+02		0.1827		3.6129E-04		4.7833E-09				-3030.6966076421		-1.22		1.8683E-03				0.0000000051		0.0000000002		6.3787E-07		0.0000068746				0.0000075178		0.0027				0.0000006432		0.0008020118								8.737E-10		0.0000000002		2.1016E-10		3.235E-10

		9.E-01		9.0013E-01		2.1555E-03		1.0948755278		293.0359343434		0.1827455437		0.0003612925

		9.E-01		8.9988E-01		2.1548E-03		1.0948611426		293.0736111111		0.1827925394		0.0003611631

																												Unc. P st relative - August

		8/25/05																				Molecular mean		1.1124				0.0030383475

																						st dev mol. Mean		0.0005536835				0.0029782627

		P Nom (Pa)		P St (Pa)		U St		s_1		T		Var T		DCR1 Corrected														0.0029715272

		3.E-04		2.9852E-04		9.0697E-07		1.1110		293.0535		0.2406		0.0000001216														0.002970936

		3.E-04		2.9887E-04		9.0800E-07		1.1106		293.2294		0.2409		0.0000001216														0.002970877

		3.E-04		2.9856E-04		9.0711E-07		1.1134		293.1641		0.2408		0.0000001218														0.0025814735

																												0.0024097212

		9.E-04		9.0490E-04		2.6950E-06		1.1106		293.0133		0.2406		0.0000003684														2.3946E-03

		9.E-04		9.0512E-04		2.6957E-06		1.1115		293.0233		0.2406		0.0000003688

		9.E-04		9.0553E-04		2.6969E-06		1.1119		293.0736		0.2407		0.0000003691

		3.E-03		3.0226E-03		8.9816E-06		1.1123		293.1364		0.2408		0.0000012323

		3.E-03		3.0210E-03		8.9770E-06		1.1136		292.8073		0.2402		0.0000012338

		3.E-03		3.0225E-03		8.9813E-06		1.1134		292.9605		0.2405		0.0000012339

		9.E-03		9.0624E-03		2.6924E-05		1.1130		293.0033		0.2406		0.000003698

		9.E-03		9.0614E-03		2.6921E-05		1.1131		293.1063		0.2407		0.0000036972

		9.E-03		9.0643E-03		2.6929E-05		1.1134		293.2143		0.2409		0.0000036987

		2.E-02		2.0146E-02		5.9851E-05		1.1129		293.2571		0.2410		0.0000082164

		2.E-02		2.0148E-02		5.9857E-05		1.1128		293.2872		0.2410		0.0000082161

		2.E-02		2.0141E-02		5.9836E-05		1.1134		293.0988		0.2407		0.0000082198

		9.E-02		8.8320E-02		2.2797E-04		1.1132		293.0963		0.1828		0.0000360383

		9.E-02		8.8214E-02		2.2773E-04		1.1136		293.0837		0.1828		0.0000360081

		9.E-02		8.8118E-02		2.2752E-04		1.1133		293.1415		0.1829		0.0000359578

		3.E-01		2.9901E-01		7.2054E-04		1.1092		292.8927		0.1826		0.000121617

		3.E-01		2.9934E-01		7.2133E-04		1.1092		292.9380		0.1826		0.0001217457

		3.E-01		2.9876E-01		7.1995E-04		1.1091		292.9556		0.1826		0.0001214943

		9.E-01		9.0031E-01		2.1559E-03		1.0950		292.9682		0.1827		0.0003614309

		9.E-01		9.0032E-01		2.1559E-03		1.0949		293.0335		0.1827		0.0003613877

		9.E-01		9.0017E-01		2.1555E-03		1.0950		293.0863		0.1828		0.0003613077

		8/26/05

				2.9901E-01

		P Nom (Pa)		P St (Pa)		U St		s_1		T		Var T		DCR1 Corrected

		3.E-04		2.9881E-04		9.0783E-07		1.1104		293.3267		0.2411		0.0000001216

		3.E-04		2.9975E-04		9.1055E-07		1.1030		292.7344		0.2401		0.0000001213

		3.E-04		2.9806E-04		9.0565E-07		1.1062		292.8424		0.2403		0.0000001209

		9.E-04		9.0583E-04		2.6978E-06		1.1110		293.1541		0.2408		0.0000003689

		9.E-04		9.0569E-04		2.6974E-06		1.1112		293.2018		0.2409		0.0000003688

		9.E-04		9.0560E-04		2.6971E-06		1.1115		293.2444		0.2410		0.0000003689

		3.E-03		3.0216E-03		8.9789E-06		1.1121		292.8048		0.2402		0.0000012324

		3.E-03		3.0228E-03		8.9822E-06		1.1128		292.8575		0.2403		0.0000012335

		3.E-03		3.0226E-03		8.9817E-06		1.1130		292.9380		0.2404		0.0000012335

		9.E-03		9.0676E-03		2.6939E-05		1.1126		293.0436		0.2406		0.0000036984

		9.E-03		9.0667E-03		2.6937E-05		1.1130		293.1340		0.2408		0.0000036988

		9.E-03		9.0652E-03		2.6932E-05		1.1135		292.9832		0.2405		0.000003701

		2.E-02		2.0147E-02		5.9854E-05		1.1132		292.9028		0.2404		0.0000082234

		2.E-02		2.0151E-02		5.9866E-05		1.1131		293.0158		0.2406		0.0000082233

		2.E-02		2.0146E-02		5.9852E-05		1.1133		292.9908		0.2405		0.0000082228

		9.E-02		8.8017E-02		2.2729E-04		1.1112		293.0134		0.1827		0.000035857

		9.E-02		8.8263E-02		2.2784E-04		1.1121		292.9732		0.1827		0.000035987

		9.E-02		8.8194E-02		2.2769E-04		1.1119		292.8350		0.1825		0.0000359621

		3.E-01		2.9939E-01		7.2145E-04		1.1087		292.8627		0.1825		0.0001217264

		3.E-01		2.9958E-01		7.2189E-04		1.1086		292.8702		0.1825		0.0001217913

		3.E-01		2.9900E-01		7.2051E-04		1.1085		292.9205		0.1826		0.0001215332

		9.E-01		8.9969E-01		2.1544E-03		1.0950		292.9582		0.1826		0.0003612173

		9.E-01		8.9998E-01		2.1551E-03		1.0950		293.0561		0.1828		0.0003612737

		9.E-01		8.9973E-01		2.1545E-03		1.0948		293.0587		0.1828		0.000361089






_1193061489.doc


A







buffer







inlet pressure







Q







P







L







O







K







F







E







D







C







B







N







S







R







M















H







G







V4











v4







V3







v3







V2







v2







V1







v1




















































_1201344330.unknown

_1192957621.unknown

_1192889965.unknown

_1192890002.unknown

_1192890093.unknown

_1192889897.unknown

_1191067124.unknown

_1178698281.unknown

_1182599742.unknown

_1182599901.unknown

_1180791837.unknown

_1180523447.unknown

_1161074342.unknown

_1161093060.unknown

_1161093167.unknown

_1165232137.unknown

_1170571828.unknown

_1161093117.unknown

_1161090770.unknown

_1161092963.unknown

_1161085605.unknown

_1161074226.unknown

_1161074304.unknown

_1161073893.unknown

_1161074141.unknown

_1161073170.unknown

_1159104726.unknown

_1159339906.unknown

_1159703924.unknown

_1161072526.unknown

_1161072559.unknown

_1159703945.unknown

_1159339918.unknown

_1159356634.unknown

_1159193214.unknown

_1159193306.unknown

_1159193041.unknown

_1159104447.unknown

_1159104542.unknown

_1159104573.unknown

_1159104483.unknown

_1159104313.unknown

_1159104358.unknown

_1159104219.unknown

