Bilateral comparison EUROMET 900a

CO in nitrogen (5 µmol.mol-1)
Draft B
Angelique Botha1, Mellisa Janse van Rensburg1, James Tshilongo1, Ipeleng Mokgoro1, Nompumelelo Leshabane1, Rob M. Wessel2,

 1CSIR National Metrology Laboratory (CSIR NML), CSIR, Building 5, Meiring Naudé Road, Brummeria, Pretoria, 0001, South Africa

2NMi Van Swinden Laboratorium B.V. (NMi VSL), Thijsseweg 11, 2629 JA Delft, the Netherlands
Field

Amount of substance
Subject

Permanent gases (carbon monoxide in nitrogen)
Participants

NMi VSL, CSIR NML
Organising body

EUROMET

Table of contents
Field
1
Subject
1

Participants
1

Organising body
1

Table of contents
2

Introduction
3

Measurements standards
3

Measurement protocol
3

Schedule
4

Measurement methods
5

Degrees of equivalence
5

Computation of degrees of equivalence and associated uncertainties
5

Results
7
Discussion of results
12
Conclusions
13
References
14
Appendix: Measurement reports
15
Introduction
The first key comparison on carbon monoxide in nitrogen dates from 1992 (CCQM-K1a) [1]. It is in fact one of the first types of gas mixtures that was used in an international (key) comparison. Since then, numerous national metrology institutes (NMIs) have been setting up facilities for gas analysis, and have developed claims for their calibration and measurement capabilities (CMCs) for these mixtures. Furthermore, in the April 2005 meeting of the CCQM Gas Analysis Working Group, a policy was proposed to repeat key comparisons for stable mixtures every 10 years. Consequently, this comparison is consistent with the proposed policy and provides an opportunity for NMIs that could not participate in the previous comparison.
In preparation for a CCQM key comparison on CO in nitrogen (1-10 µmol.mol-1 range) a bilateral comparison is organised between the coordinating laboratory and another laboratory which participated in CCQM-K1. In this comparison CSIR NML will prepare two gas mixtures to be analysed by NMi VSL and NMi VSL will also prepare two mixtures to be analysed by CSIR NML. This would allow for the CSIR NML to validate its capabilities for this component in this range.
Measurement standards
Table 1 gives the nominal composition used throughout this comparison.
Table 1: Nominal composition for the comparison

	Component
	x

 (µmol.mol-1)

	Carbon monoxide
	5

	Nitrogen
	balance


Measurement protocol
A set of mixtures of carbon monoxide in nitrogen at approximately 5 µmol.mol-1 will be prepared gravimetrically. The mixtures will be verified against primary standard mixtures (PSMs). After verification of the composition of the mixtures, the gravimetric value will be adopted as key comparison reference value (KCRV).
The pressure in the cylinders will be approximately 100 bar; cylinders of 5 dm3 nominal volume will be used. The amount-of-substance fractions as obtained from gravimetry and purity verification of the parent gases will be used as reference values. Each cylinder will have its own reference value.
Participating laboratories are requested to specify in detail which analytical method(s) have been used and how the evaluation of the measurement uncertainty was performed.  Each participating laboratory will be responsible for the calibration of its own equipment. For a proper evaluation of the data, it is necessary that the calibration method, as well as the way in which the calibration mixtures have been prepared is reported to the co-ordinator. This information is needed for the evaluation of the preparation facilities (as part of this bilateral comparison).
After each calibration, the measurements on the gas mixture can be performed and recorded.  Each laboratory has to express the uncertainty on all results submitted, as expanded uncertainty. The evaluation of measurement uncertainty should be in accordance to the “Guide to the expression of uncertainty in measurement” (ISO GUM). The participant should provide a detailed description of the uncertainty budget, including

· method of evaluation (type A or type B)

· (assumed) probability distribution

· standard uncertainties and sensitivity coefficients

After the measurements, the participants have to return the cylinders with a sufficient amount of gas (pressure at least 30 bar) to the coordinating laboratory for re-analysis. 

Transport of the cylinders to the participating laboratories will be arranged by the coordinating laboratory and paid for by each participant. The return of the cylinders must be arranged and paid for by the participants.
The measurement report requires per cylinder at least three independent measurements, obtained under repeatability conditions with (at least) three independent calibrations, e.g. calibration ( measurement ( calibration ( measurement ( calibration ( measurement (etc.). This is a strict requirement to come to proper statistical analysis of the reported data.

One single measurement result is usually obtained from multiple readings (sub measurements), without recalibrations. Its standard deviation provides information about the performance of the measurement system.
Additional measurement reports and additional information can be submitted jointly with the report forms to the CSIR NML and will be taken into consideration during the evaluation. Please send all information in Word 2000 format
 (or compatible, such as RTF) by e-mail and /or by mail/fax.

Schedule
The schedule for this part of the comparison is presented in Table 2.
Table 2: Schedule followed for EUROMET 900a
	January 2006
	CSIR NML sent two 5 µmol.mol-1 cylinders to NMi VSL for measurement

	March 2006
	Measurement report received from NMi VSL

	April 2006
	NMi VSL sent a 10 µmol.mol-1 and a 8 µmol.mol-1 cylinder to CSIR NML for measurement

	June 2006
	Measurement report submitted by CSIR NML

	July to September 2006
	CSIR NML perform maintenance work on the primary standard mixtures for CO in nitrogen

	December 2006
	CSIR NML sent a second set of 5 µmol.mol-1 mixtures to NMi VSL

	January 2007
	Measurement report received from NMi VSL

	January 2007
	Draft A report

	April 2007
	Draft B report

	May to June 2007
	Another round of cylinder exchange between NMi VSL and CSIR NML


Measurement methods
Table 3: Summary of the measurement methods of the participants

	Laboratory code
	Calibration
	Traceability
	Matrix standards
	Measurement technique

	CSIR NML
	ISO 6143
	Own standards
	Nitrogen
	GC-FID

	NMi VSL
	ISO 6143
	Own standards
	Nitrogen
	NDIR


Degrees of equivalence [4]
Usually all participants perform analyses on the same artefact and the key comparison reference value is calculated from the mean of the individual results. In the current comparison on gas mixtures, measurements were performed on individually prepared gas mixtures with (slightly) different concentrations. Since the coordinating laboratory prepared these mixtures using the same methods and materials, the individual gravimetric values can be adopted as reference values, despite the small differences that exist. There is a problem that these small differences are of the same order of the differences found between the national metrology institutes, and will thus influence the outcome of the key comparison if it would be operated with a single reference value.
In order to evaluate the differences between the participating national metrology institutes, the value 
[image: image1.wmf]igrav

x

 is taken as the reference value for laboratory 
[image: image2.wmf]i

.

The degree of equivalence of each laboratory with respect to the reference value is given by a pair of numbers:


[image: image3.wmf](

)

igrav

i

i

x

x

D

-

=


(1)

and 
[image: image4.wmf]i

U

, its expanded uncertainty, 


[image: image5.wmf](

)

2

2

2

2

iprep

i

i

u

u

k

U

+

=



[image: image6.wmf](

)

2

=

k


(2)
Computation of degrees of equivalence and associated uncertainties

The model of CCQM-K1 has been adapted for this bilateral comparison. Although the uncertainty on the assigned value comes close to the reported uncertainties from some of the participants, the model of CCQM-K1 can still be used.
The degree of equivalence between two laboratories is defined as


[image: image7.wmf][

]

[

]

jprep

j

iprep

i

j

i

ij

x

x

x

x

D

D

D

-

-

-

=

-

=


(3)

The uncertainty in the degree of equivalence can be expressed as


[image: image8.wmf](

)

(

)

jprep

iprep

jprep

iprep

j

i

ij

x

x

u

u

u

u

u

k

U

,

2

2

2

2

2

2

2

-

+

+

+

=


(4)

Whereby


[image: image9.wmf]2

2

stab

grav

prep

u

u

u

+

=


(5)

Looking at the results, 
[image: image10.wmf]iprep

u

 can generally not be neglected compared to 
[image: image11.wmf]i

u

. The degree of correlation between 
[image: image12.wmf]iprep

x

 and 
[image: image13.wmf]prepi

x

 has been assessed and lead to the following simplification


[image: image14.wmf](

)

2

2

2

2

2

2

jgrav

igrav

j

i

ij

u

u

u

u

k

U

+

+

+

=


(6)

This simplification is the result of the establishment of 
[image: image15.wmf]stab

u

 (see equation (5)). From the stability data and the calibration of the equipment, it becomes clear that 
[image: image16.wmf]stab

u

 is a combined standard uncertainty from three main contributors: uncertainty of mixtures used, position of the “unknown(s)” on the curve, including lack of fit, and the repeatability of measurement. In formula, this expression reads as


[image: image17.wmf]2

2

2

2

r

lof

calibrants

stab

u

u

u

u

+

+

=


(7)

whereby 
[image: image18.wmf]calibrants

u

 denotes the contribution from the mixtures used for calibrating the equipment, 
[image: image19.wmf]lof

u

 denotes the contribution due to the fitting and interpolation procedure, and ur denotes the contribution due to repeatability of measurement. In this expression, 
[image: image20.wmf]lof

u

 is by far the dominant factor. It is also a factor, that is the same for all mixtures prepared for this comparison. As a result, the covariance between 
[image: image21.wmf](

)

A

ref

x

 and 
[image: image22.wmf](

)

B

ref

x

 can be approximated by


[image: image23.wmf](

)

2

2

,

calibrants

lof

gravj

gravi

u

u

x

x

u

+

»


(8)

Substitution of (5) and the approximation for the covariance in (8) leads then to (6).

Results
First comparison – March 2006
At the beginning of 2006 the CSIR NML prepared two mixtures of 5 µmol.mol-1 CO in nitrogen and sent it to NMi VSL for comparison.  The results are presented in Table 4 with the following data:


[image: image24.wmf]prep

x


amount of substance fraction, from preparation (10-6 mol.mol-1)

[image: image25.wmf]prep

u


uncertainty of 
[image: image26.wmf]prep

x

 (10-6 mol.mol-1)


[image: image27.wmf]ver

u


uncertainty from verification (10-6 mol.mol-1)


[image: image28.wmf]ref

u


uncertainty of reference value (10-6 mol.mol-1)


[image: image29.wmf]lab

x


result of laboratory (10-6 mol.mol-1)


[image: image30.wmf]lab

U


stated uncertainty of laboratory, at 95% level of confidence (10-6 mol.mol-1)


[image: image31.wmf]lab

k


stated coverage factor


[image: image32.wmf]x

D


difference between laboratory result and reference value (10-6 mol.mol-1)


[image: image33.wmf]k


assigned coverage factor for degrees of equivalence


[image: image34.wmf](

)

x

U

D


expanded uncertainty of difference, 
[image: image35.wmf]x

D

, at 95% level of confidence (10-6 mol.mol-1)
Table 4: A summary of the results for the first comparison.

	Laboratory code
	Cylinder


	 
[image: image36.wmf]prep

x


	 
[image: image37.wmf]prep

u


	
[image: image38.wmf]ver

u

 
	 
[image: image39.wmf]ref

u


	 
[image: image40.wmf]lab

x


	 
[image: image41.wmf]lab

U


	 
[image: image42.wmf]lab

k


	 
[image: image43.wmf]x

D


	
[image: image44.wmf]k

 
	 
[image: image45.wmf](

)

x

U

D



	NMi VSL
	MG3907
	5.041
	0.019
	0.019
	0.027
	5.006
	0.034
	2
	0.035
	2
	0.064

	 
	6664E
	5.188
	0.019
	0.019
	0.027
	5.166
	0.034
	2
	0.022
	2
	0.064


[image: image112.emf](


)


x


U


D




  x U


Figure 1: Degrees of equivalence for cylinder MG3907 and 6664E in the first comparison.
Second comparison – May 2006
In May 2006 NMi VSL sent two mixtures to the CSIR NML for measurement.  The results for this second comparison are presented in Table 5.

Table 5: A summary of the results for the second comparison.

	Laboratory code
	Cylinder


	 
[image: image46.wmf]prep

x


	 
[image: image47.wmf]prep

u


	
[image: image48.wmf]ver

u

 
	 
[image: image49.wmf]ref

u


	 
[image: image50.wmf]lab

x


	 
[image: image51.wmf]lab

U


	 
[image: image52.wmf]lab

k


	 
[image: image53.wmf]x

D


	
[image: image54.wmf]k

 
	 
[image: image55.wmf](

)

x

U

D



	CSIR NML
	M60344
	10.018
	
	0.017
	0.017
	10.021
	0.080
	2
	0.003
	2
	0.080

	
	0604E
	8.0019
	 
	0.017 
	0.017
	8.072
	0.125
	2
	0.070
	2
	0.125


[image: image113.emf](


)


x


U


D




  x U


Figure 2: Degrees of equivalence for cylinder M60344 and 0604E in the second comparison.

Third comparison – December 2006
During the period from July to October 2006 the CSIR NML performed maintenance work on the primary standard mixtures (PSMs) for the complete suite of CO in nitrogen standards.  The laboratory then sent another two mixtures of 5 µmol.mol-1 CO in nitrogen to the NMi VSL for analysis in December 2006.  The results from this third comparison are summarised in Table 6.
Table 6: A summary of the results for the third comparison.

	Laboratory code
	[image: image114.emf](


)


x


U


D




  x U

Cylinder


	 
[image: image56.wmf]prep

x


	 
[image: image57.wmf]prep

u


	
[image: image58.wmf]ver

u

 
	 
[image: image59.wmf]ref

u


	 
[image: image60.wmf]lab

x


	 
[image: image61.wmf]lab

U


	 
[image: image62.wmf]lab

k


	 
[image: image63.wmf]x

D


	
[image: image64.wmf]k

 
	 
[image: image65.wmf](

)

x

U

D



	NMi VSL
	D958384
	5.032
	0.026
	0.03
	0.04
	5.018
	0.025
	2
	0.014
	2
	0.084

	 
	M555703
	5.452
	0.026
	0.03
	0.04
	5.439
	0.027
	2
	0.013
	2
	0.084


[image: image115.emf](


)


x


U


D




  x U


Figure 3: Degrees of equivalence for cylinder D958384 and M5555703 in the third comparison.
Further work – March 2007
The draft A report was presented to the METCHEM meeting in February 2007.  The meeting proposed further work on the report to especially improve the gravimetric uncertainty of the 5 µmol.mol-1 CO in nitrogen results, but also attempt to improve the verification uncertainty.  The CSIR NML did some work on the purity analysis to optimise the GC-FID method for dedicated evaluation of the CO impurity in Nitrogen 6.0 instead of doing multi-component purity analysis for CO, CO2, O2 and argon in nitrogen with one method on PDHID (see Table 7).
Table 7: Comparison of the results for CO in BIP™ nitrogen by three different methods

	
	GC –FID

(NML-06-0010, 2006)
	GC-PDHID

(NML-06-0239, 2006)
	GC-FID

(NML-07-0038, 2007)

	LOD (x10-6 mol.mol-1)
	0.038
	0.037
	0.012

	LOQ (x10-6 mol.mol-1)
	0.125
	0.123
	0.042

	Repeatability (%CV)
	<1.4% (1.2ppm)
	<10%(0.2ppm)
	<2.3% (0.2ppm)

	Linearity 

(goodness of fit)
	<0.0949
	<1.2753
	<1.0126

	CO (x10-6 mol.mol-1)
	Not detected

 (<0.038)
	Not detected (<0.037)
	0.022

	Expanded uncertainty

(x10-6 mol.mol-1)
	
	
	0.012

	UCL of Result 

(x10-6 mol.mol-1)
	0.069
	0.068
	0.034

	LCL of Result 

(x10-6 mol.mol-1)
	0.026
	0.025
	0.010


The new gravimetric results with the new purity analysis data for CO in nitrogen 6.0 in comparison with the measured results by NMi VSL for cylinders D958384 and M555703 are summarised in Table 8.

Table 8: A comparison of the improved gravimetric results with the measurement results of NMi VSL during the third comparison.

	Laboratory code
	[image: image116.emf]EUROMET 900: CO in nitrogen

Cylinder nr: MG3907

4.92

4.97

5.02

5.07

5.12

CSIR NML NMi VSL

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: 6664E

5.05

5.1

5.15

5.2

5.25

5.3

CSIR NML NMi VSL

Concentration 

(µmol/mol)

Cylinder


	 
[image: image66.wmf]prep

x


	 
[image: image67.wmf]prep

u


	
[image: image68.wmf]ver

u

 
	 
[image: image69.wmf]ref

u


	 
[image: image70.wmf]lab

x


	 
[image: image71.wmf]lab

U


	 
[image: image72.wmf]lab

k


	 
[image: image73.wmf]x

D


	
[image: image74.wmf]k

 
	 
[image: image75.wmf](

)

x

U

D



	NMi VSL
	D958384
	5.029
	0.006
	0.025
	0.026
	5.018
	0.025
	2
	0.011
	2
	0.029

	 
	M555703
	5.449
	0.006
	0.025
	0.026
	5.439
	0.027
	2
	0.010
	2
	0.030


The latest results for the measurements performed by the CSIR NML with the improved verification uncertainty for the 5 µmol.mol-1 CO in nitrogen mixtures are summarised in Table 9.

Table 9: A summary of the results for the most recent measurements of 5 µmol.mol-1 CO in nitrogen by the CSIR NML.
	Laboratory code
	[image: image117.emf]EUROMET 900: CO in nitrogen

Cylinder nr: M60344

9.8

9.9

10

10.1

10.2

NMi VSL CSIR NML

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: 0604E

7.8

7.9

8

8.1

8.2

8.3

NMi VSL CSIR NML

Concentration 

(µmol/mol)

Cylinder


	 
[image: image76.wmf]prep

x


	 
[image: image77.wmf]prep

u


	
[image: image78.wmf]ver

u

 
	 
[image: image79.wmf]ref

u


	 
[image: image80.wmf]lab

x


	 
[image: image81.wmf]lab

U


	 
[image: image82.wmf]lab

k


	 
[image: image83.wmf]x

D


	
[image: image84.wmf]k

 
	 
[image: image85.wmf](

)

x

U

D



	CSIR NML
	D958384
	5.029
	0.006
	0.025
	0.026
	5.025
	0.058
	2
	0.007
	2
	0.078

	 
	M555703
	5.449
	0.006
	0.025
	0.026
	5.435
	0.063
	2
	0.017
	2
	0.082


[image: image118.emf]EUROMET 900: CO in nitrogen

Cylinder nr: D958384

4.8

4.9

5

5.1

5.2

CSIR NML NMi VSL

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: M555703

5.2

5.3

5.4

5.5

5.6

CSIR NML NMi VSL

Concentration 

(µmol/mol)


Figure 4: Latest results for cylinder D958384 and M5555703 for measurements performed by the CSIR NML.
Fourth comparison – May to August 2007
After the METCHEM and CCQM GAWG meetings in February and April 2007, respectively, the CSIR NML and NMi VSL decided to exchange another round of mixtures in an attempt to explain and eliminate the remaining bias between the two laboratories.  In May 2007 the CSIR NML shipped a 4, 5 and 6 µmol.mol-1 CO in nitrogen mixture to the NMi VSL and received the measurement report at the end of June.
Table 10: A summary of the results for the analysis performed by NMi VSL on the CSIR NML mixtures during the fourth comparison.
	 
	cyl. #
	xprep
	uprep
	uver
	uref
	xlab
	Ulab
	Δx
	U(Δx)
	klab

	NMi VSL
	D958306
	3.969
	0.006
	0.012
	0.013416
	3.962
	0.020
	0.007
	0.033
	2

	 
	D958363
	4.996
	0.006
	0.009
	0.010817
	4.983
	0.025
	0.013
	0.033
	2

	 
	D954870
	6.029
	0.006
	0.023
	0.02377
	6.027
	0.030
	0.002
	0.056
	2


[image: image119.emf]EUROMET 900: CO in nitrogen

Cylinder nr: D958384

4.8

4.9

5

5.1

5.2

CSIR NML NMi VSL CSIR NML CSIR

NML(grav)

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: M555703

5.2

5.3

5.4

5.5

5.6

CSIR NML NMi VSL CSIR NML CSIR

NML(grav)

Concentration 

(µmol/mol)


Figure 5: Results of the 4, 5 and 6 µmol.mol-1 mixtures prepared by CSIR NML and analysed by NMi-VSL in June 2007.

During the same period the CSIR NML received a 4, 5 and 6 µmol.mol-1 CO in nitrogen mixture from the NMi-VSL and the measurement report was submitted to the NMi-VSL in August 2007.

Table 11: A summary of the results for the analysis performed by CSIR NML on the NMi VSL mixtures during the fourth comparison.
	 
	cyl. #
	xprep
	uprep
	uver
	uref
	xlab
	Ulab
	Δx
	U(Δx)
	klab

	CSIR NML
	152995
	5.9956
	0.0028
	0.012
	0.012
	5.985
	0.024
	-0.0106
	0.027
	2

	 
	153048
	4.9986
	0.0023
	0.010
	0.010
	5.009
	0.017
	0.0104
	0.020
	2

	 
	153190
	3.9976
	0.0022
	0.010
	0.010
	4.015
	0.018
	0.0174
	0.021
	2


[image: image86.emf]EUROMET 900: CO in nitrogen

Cylinder nr: 152995

5.9

5.95

6

6.05

NMi VSL CSIR NML

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: 153048

4.95

5

5.05

NMi VSL CSIR NML

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: 153190

3.95

4

4.05

NMi VSL CSIR NML

Concentration 

(µmol/mol)


Figure 6: Results of the 4, 5 and 6 µmol.mol-1 mixtures prepared by NMi-VSL and analysed by CSIR NML in June 2007.

Discussion of results
In the first comparison there was an apparent bias of approximately 20 ppb between the two laboratories where the CSIR NML result was higher for both cylinders.  The uncertainties reported by the CSIR NML were also approximately double the uncertainties reported by the NMi-VSL.  The amount of CO in the nitrogen balance gas for cylinders MG3907 and 6664E was 32 ppb.
In the second comparison the CSIR NML experienced some problems with the calibration of the GC-FID results.  Closer inspection of the chromatogram for the 8 µmol.mol-1 CO in nitrogen mixture (Cylinder nr. 0604E) revealed that a small contamination peak of methane co-eluted with the CO peak, which resulted in a bias the measurement results.  The measurement method on the GC-FID was adopted to resolve the methane peak from the CO peak and the bias was eliminated.  However, for the effective resolution of the CO and methane peaks the temperature of the column oven was reduced in the GC method, which resulted in a broadening of the CO peak.  This resulted in deterioration in the precision of the method and the uncertainties of the measurement results increased.
[image: image87.png]et 03

Grpvenha

To

T
5





Figure 7: Chromatogram of the methane peak resolved from the CO peak in cylinder nr. 0604E
The CSIR NML embarked on an exercise to reduce the bias between the two laboratories by performing maintenance work on the primary standard mixtures for CO in nitrogen with improved purity analysis of the nitrogen balance gas.   The amount of CO in the nitrogen balance gas used for cylinders D958384 and M 555703 was 24 ppb.  The actual CO in the nitrogen balance gas was still below the detection limit of the GC-PDHID method.  The next step that will be investigated in an attempt to quantify the amount of CO in the nitrogen gas more accurately is to make use of packed columns that has a larger sample capacity than capillary columns.
The results from the third comparison (Figure 3) indicate a good comparison between the measurement results for the two mixtures, but the uncertainties reported by the CSIR NML is still significantly larger than those of the NMi VSL.
The results in Table 7 summarises the improvements the CSIR NML has made in the purity analysis of CO in nitrogen 6.0.  Figure 4 shows the effect of the improved purity analysis on the comparison between the NMi VSL measurement and the gravimetric value of the CSIR NML.  Due to the reduction in the effect of atmospheric pressure on the GC results the verification uncertainties have also improved.
Conclusions

The CSIR NML has purchased an NDIR monitor for the measurement of CO over the concentration range of 1 to 100 µmol/mol, but the new instrument arrived in the laboratory in April 2007 and is not yet fully commissioned.  It is expected that analysis of the CO in nitrogen by NDIR will reduce the uncertainty of the results by 50%.  In the meantime the laboratory is continuing to work on the GC-FID method in an attempt to improve the repeatability.
Further work on the purity analysis of CO in nitrogen has resulted in a reduction in the gravimetric uncertainty of the 5 µmol.mol-1 CO in nitrogen of the CSIR NML and a slight reduction in the bias between the gravimetric measurement result of the CSIR NML and the analytical measurement result of the NMi VSL.  The verification uncertainty of the CSIR NML has also improved.
During the fourth comparison between the CSIR NML and the NMi VSL it is shown with the uncertainties of the analytical capabilities of the two laboratories that they are now comparable.

References

[1]
Alink A., The first key comparison on Primary Standard gas Mixtures, Metrologia 37 (2000), pp. 35-49

[2]
International Organization for Standardization, “ISO 6142 -- Gas analysis -- Preparation of calibration gas mixtures -- Gravimetric method”, ISO, Geneva (CH), 2001

[3]
International Organization for Standardization, “ISO 6143 -- Gas analysis -- Comparison methods for determining and checking the composition of calibration gas mixtures”, ISO, Geneva (CH), 2001
[4]
Van der Veen A.M.H., Cox M.G., “Degrees of equivalence across key comparisons in gas analysis”, Metrologia 40 (2003), pp. 18-23
Coordinator

Angelique Botha

CSIR National Metrology Laboratory (CSIR NML)

CSIR, Building 5

Meiring Naudé Road

Brummeria, PRETORIA, 0001

South Africa

Tel: +27 12 841 3800

Fax: +27 12 841 2131

e-mail: abotha@csir.co.za

Project reference

EUROMET 900
Completion date

In progress
Appendix
Report form Euromet bilateral CSIR-NML and NMi-VSL

Laboratory

: NMi-VSL
Cylinder numbers
: MG 3907-  6664E

Measurement #1

	Cylinder
	Date
(dd/mm/yy)
	Result
((mol/mol)
	Standard uncertainty (% rel)

	MG3907
	13-06-14
	5.014
	0.21

	6664E
	13-06-14
	5.164
	0.20


Measurement #2 

	Cylinder
	Date
(dd/mm/yy)
	Result
((mol/mol)
	Standard uncertainty(% rel)

	MG3907
	13-06-14
	5.009
	0.18

	6664E
	13-06-14
	5.172
	0.19


Measurement #3 

	Cylinder
	Date
(dd/mm/yy)
	Result
((mol/mol)
	Standard uncertainty(% rel)

	MG3907
	13-06-14
	4.996
	0.19

	6664E
	13-06-14
	5.162
	0.20


Result

	Cylinder
	Assigned value( x) ((mol/mol)
	Standard uncertainty (u(x))

	MG3907
	5.006
	0.017

	6664E
	5.166
	0.017


Analysed using a set of 6 PSM’s ranging from 2 – 10 ((mol/mol) CO in N2. 

Analyser: ND-IR 

Analyses in accordance with ISO 6143. Quadratic fit type. Corrections for barometric drift.

Report Form EUROMET 900 Carbon monoxide in nitrogen

Laboratory name:
NML

Cylinder number:
M60344

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	05/05/2006
	10,031
	0,27
	15


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	22/05/2006
	10,043
	0,22
	15


Measurement #3
 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	23/05/2006
	9,9877
	0,33
	15


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty
	Coverage factor


	CO
	10,021
	0,080
	2


Report Form EUROMET 900 Carbon monoxide in nitrogen

Laboratory name:
NML

Cylinder number:
0604E

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	05/05/2006
	8,1373
	0,40
	15


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	22/05/2006
	8,0427
	0,30
	15


Measurement #3
 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	23/05/2006
	8,0373
	0,32
	15


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty
	Coverage factor


	CO
	8,072
	0,125
	2


Report Form Euromet project 900 Carbon monoxide in nitrogen

Laboratory name: NMi-VSL

Cylinder number: D 95 8384

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	19/12/06
	0,000005015
	0,23
	4*30


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	21/12/06
	0,000005021
	0,21
	3 * 30


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty (% relative)
	Coverage factor


	CO
	0,000005018
	0,5
	2


Cylinder number: M555703

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	19/12/06
	0,000005445
	0,21
	4*30


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	21/12/06
	0,000005433
	0,18
	3 * 30


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty (% relative)
	Coverage factor


	CO
	0,000005439
	0,5
	2


Reference Method:

For the determination of Carbon monoxide (CO):

ISO 6143;2001 was used and the generalized distance regression (GDR) method was applied using 7 NMi PSM’s and a cubic regression model.

Instrument Calibration:

The following NMi-VSL PSMs were used to calibrate the ND-IR analyser (ABB URAS 14). All 7 PSMs contained CO in a matrix of nitrogen similar to the sample mixture.

	Cylinder No
	Gravimetric composition
	Standard uncertainty

	VSL205201
	0,9998 · 10-6 mol/mol
	1,82 · 10-9 mol/mol

	VSL305260
	2,0012 · 10-6 mol/mol
	1,85 · 10-9 mol/mol

	VSL205149
	3,9915 · 10-6 mol/mol
	1,92 · 10-9 mol/mol

	VSL308613
	5,0013 · 10-6 mol/mol
	2,84 · 10-9 mol/mol

	VSL300990
	6,0014 · 10-6 mol/mol
	2,17 · 10-9 mol/mol

	VSL308615
	7,9997 · 10-6 mol/mol
	3,03 · 10-9 mol/mol

	VSL238362
	9,9971 · 10-6 mol/mol
	2,50 · 10-9 mol/mol


Sample handling:

Each cylinder was equipped with a stainless steel pressure regulator that was adequately purged.

A flow of approx. 350 mL/min was flushed for three minutes through the ND-IR analyser before the readings were taken.  

Calibration Standards:

The PSM’s used in calibration are prepared from pre-mixtures in accordance with ISO 6142: 2001
(Gas analysis - Preparation of calibration gas mixtures - Gravimetric method). 
After preparation the composition was verified. 

For nitrogen 6.0 Quality from Hoek Loos (Linde) and Scott Specialty Gases was used. The CO content was analysed using FT-IR. Five different nitrogen cylinders were analysed to contain less then detection limit level. This detection limit was estimated to be 3 and later 6 nmol/mol CO.

This value for the CO content and its associated uncertainty were taken into account when doing the gravimetric calculations.

After preparation the standards were verified by analytical comparisons against existing gravimetrically prepared standards. Only when no significant difference between the analysed and the calculated gravimetric composition is found, the “new prepared candidate ” is accepted as a PSM. Mixtures were prepared in 2002 (1), 2003 (3), 2004 (2) and 2006 (1). 

Evaluation of measurement uncertainty

The listed gravimetric uncertainty is a combined standard uncertainty and comes from:

· The uncertainty in the weighings (pooled estimate of 3 mg on minimum 20 g)

· The effect on the buoyancy correction due to cylinder expansion when pressurizing to 120 bar (approx. 20 mg on 600 g)

· The uncertainty on the purity analysis 

· The uncertainty in the molar masses

The listed standard deviations in the two measurements come from the GDR taking into account the gravimetric uncertainties on the PSMs and the standard deviation in the responses.

Both uncertainties can be combined in quadrature.

For D958384:

Standard uncertainty measurements: (((0,23/(4)2 + (0,21/(3)2)   = 0,17% rel.

Standard uncertainty in gravimetry: 0,18% rel.

Combined: 0.25 % rel.

Expanded: 0.5 % rel. (k=2)

For M555703:

Standard uncertainty measurements: (((0,21/(4)2 + (0,18/(3)2)   = 0,15% rel.

Standard uncertainty in gravimetry: 0,18% rel.

Combined: 0.23 % rel.

Expanded: 0.5 % rel. (k=2)

Report Form Euromet project 900 Carbon monoxide in nitrogen

Laboratory name: NMi-VSL

Cylinder number: D 95 8306 

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	07/06/2007
	0,000003962
	0,24
	4


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	13/06/2007
	0,000003960
	0,24
	4 


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty (% relative)
	Coverage factor


	CO
	0,000003961
	0,5
	2


Cylinder number: D 95 8363

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	07/06/2007
	0,000004984
	0,20
	4


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	13/06/2007
	0,000004982
	0,17
	4 


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty (% relative)
	Coverage factor


	CO
	0,000004983
	0,5
	2


Cylinder number: D 95 4870

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	07/06/2007
	0,000006027
	0,22
	4


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	13/06/2007
	0,000006027
	0,10
	4 


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty (% relative)
	Coverage factor


	CO
	0,000006027
	0,5
	2


Reference Method:

For the determination of Carbon monoxide (CO):

ISO 6143;2001 was used and the generalized distance regression (GDR) method was applied using 6 NMi PSM’s and a quadratric regression model.

Instrument Calibration:

The following NMi-VSL PSMs were used to calibrate the ND-IR analyser (ABB URAS 14). All 6 PSMs contained CO in a matrix of nitrogen similar to the sample mixture.

	Cylinder No
	Gravimetric composition
	Standard uncertainty

	VSL305260
	2,0012 · 10-6 mol/mol
	1,85 · 10-9 mol/mol

	VSL205149
	3,9915 · 10-6 mol/mol
	1,92 · 10-9 mol/mol

	VSL308613
	5,0013 · 10-6 mol/mol
	2,84 · 10-9 mol/mol

	VSL300990
	6,0014 · 10-6 mol/mol
	2,17 · 10-9 mol/mol

	VSL308615
	7,9997 · 10-6 mol/mol
	3,03 · 10-9 mol/mol

	VSL238362
	9,9971 · 10-6 mol/mol
	2,50 · 10-9 mol/mol


Sample handling:

Each cylinder was equipped with a stainless steel pressure regulator that was adequately purged.

A flow of approx. 350 mL/min was flushed for five minutes through the ND-IR analyser before the readings were taken.  

Calibration Standards:

The PSM’s used in calibration are prepared from pre-mixtures in accordance with ISO 6142: 2001
(Gas analysis - Preparation of calibration gas mixtures - Gravimetric method). 
After preparation the composition was verified. 

For nitrogen 6.0 Quality from Hoek Loos (Linde Benelux) and Scott Specialty Gases was used. The CO content was analysed using FT-IR. Five different nitrogen cylinders were analysed to contain less then detection limit level. This detection limit was estimated to be 3 and later 6 nmol/mol CO.

This value for the CO content and its associated uncertainty were taken into account when doing the gravimetric calculations.

After preparation the standards were verified by analytical comparisons against existing gravimetrically prepared standards. Only when no significant difference between the analysed and the calculated gravimetric composition is found, the “new prepared candidate ” is accepted as a PSM. Mixtures were prepared in 2002 (1), 2003 (3), 2004 (2) and 2006 (1). 

Evaluation of measurement uncertainty

The listed gravimetric uncertainty is a combined standard uncertainty and comes from:

· The uncertainty in the weighings (pooled estimate of 3 mg on minimum 20 g)

· The effect on the buoyancy correction due to cylinder expansion when pressurizing to 120 bar (approx. 20 mg on 600 g)

· The uncertainty on the purity analysis 

· The uncertainty in the molar masses

The listed standard deviations in the two measurements come from the GDR taking into account the gravimetric uncertainties on the PSMs and the standard deviation in the responses.

Both uncertainties can be combined in quadrature.

For D95 8384:

Standard uncertainty measurements: (((0,24/(4)2 + (0,24/(4)2)   = 0,17% rel.

Standard uncertainty in gravimetry: 0,18% rel.

Combined: 0.25 % rel.

Expanded: 0.5 % rel. (k=2)

For D95 8363:

Standard uncertainty measurements: (((0,20/(4)2 + (0,17/(4)2)   = 0,13% rel.

Standard uncertainty in gravimetry: 0,18% rel.

Combined: 0.23 % rel.

Expanded: 0.5 % rel. (k=2)

For D19 4870:

Standard uncertainty measurements: (((0,22/(4)2 + (0,10/(4)2)   = 0,12% rel.

Standard uncertainty in gravimetry: 0,18% rel.

Combined: 0.22 % rel.

Expanded: 0.5 % rel. (k=2)

Report Form CCQM-K51 Carbon monoxide in nitrogen

Laboratory name:
NML

Cylinder number:
NMi – VSL 152995

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	04/06/2007
	5,973636
	0,05
	10


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	05/06/2007
	5,973221
	0,07
	10


Measurement #3
 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	06/06/2007
	6,006877
	0,06
	10


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty
	Coverage factor


	CO
	5,985
	0,024
	2


Report Form EUROMET 900 Carbon monoxide in nitrogen

Laboratory name:
NMISA

Cylinder number:
NMi-VSL 153048

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	04/06/2007
	5,005869
	0,05
	10


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	05/06/2007
	4,998170
	0,10
	10


Measurement #3
 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	06/06/2007
	5,023040
	0,07
	10


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty
	Coverage factor


	CO
	5,009
	0,017
	2


Report Form EUROMET 900 Carbon monoxide in nitrogen

Laboratory name:
NMISA

Cylinder number:
NMi-VSL 153190

Measurement #1 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	04/06/2007
	4,027588
	0,10
	10


Measurement #2 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	05/06/2007
	3,999380
	0,10
	10


Measurement #3
 

	Component
	Date
(dd/mm/yy)
	Result
(mol/mol)
	Standard deviation
(% relative)
	number of replicates

	CO
	06/06/2007
	4,016804
	0,07
	10


Results

	Component
	Result
(mol/mol)
	Expanded Uncertainty
	Coverage factor


	CO
	4,015
	0,018
	2


Details of the measurement method used:

Gas Chromatography with flame ionisation detection after conversion of the CO peak to methane through a methaniser. The column used was a Molesieve 13X packed column, 60/80 packing size, length of 1,2m. Sample size was 2mℓ.

Detailed uncertainty budget:

Please include a list of the uncertainty contributions, the estimate of the standard uncertainty, probability distributions, sensitivity coefficients, etc.

The intermediate reproducibility uncertainty was calculated as below in Equation 1

[image: image88.wmf]3

)

(

stddev

x

u

ILITY

REPRODUCIB

=








(1)
Table 1: Verified concentrations of NMi VSL Cylinders 152995, 153048 and 153190.
	
	Cylinder NMi VSL 152995

Verified [CO]
	Cylinder NMi VSL 152995

Verification

uncertainty


[image: image89.wmf])

(

x

u


	Cylinder NMi VSL 153048

Verified [CO]
	Cylinder NMi VSL 153048

Verification

uncertainty


[image: image90.wmf])

(

x

u



	Day 1
	5,973636
	0,0027127
	5,005869
	0,0026179

	Day 2
	5,973221
	0,0040542
	4,998170
	0,0047525

	Day 3
	6,006877
	0,0033295
	5,023040
	0,0036697

	Average
	5,984578
	
	5,009026
	

	Standard deviation
	0,0193126
	
	0,0127321
	

	Reproducibility uncertainty
	0,011150
	
	0,007351
	

	
	Cylinder  NMi VSL 153190

Verified [CO]
	Cylinder 

Verification

uncertainty


[image: image91.wmf])

(

x

u


	
	

	Day 1
	4,027588
	0,0041342
	
	

	Day 2
	3,999380
	0,0038046
	
	

	Day 3
	4,016804
	0,0028855
	
	

	Average
	4,014591
	
	
	

	Standard deviation
	0,0142337
	
	
	

	Reproducibility uncertainty
	0,008218
	
	
	


The verification uncertainties from Table 1 were combined as in Equation 2, where the sensitivities are all equal to 1 since uncertainties in the same units are being combined.


[image: image92.wmf]3

)

(

)

(

)

(

)

(

2

3

2

2

2

1

DAY

DAY

DAY

ON

VERIFICATI

u

u

u

x

u

+

+

=






(2)

The reproducibility uncertainty and the combined standard verification uncertainty are combined as shown in Equation 3. The sensitivities are equal to 1 for the individual uncertainties since they are all in the same unit (µmol.mol-1).

[image: image93.wmf]2

2

)

(

)

(

)

(

ILITY

REPRODUCIB

VERIFIC

c

u

u

x

u

+

=






(3)
Table 2: Calculation of combined uncertainty
	
	Cylinder 

NMi VSL 152995
	Cylinder 

NMi VSL 153048
	Cylinder 

NMi VSL 153190

	
[image: image94.wmf])

(

x

u

VERIFIC


	0,0034098
	0,0037818
	0,0036466

	
[image: image95.wmf])

(

x

u

ILITY

REPRODUCIB


	0,0111501
	0,0073509
	0,0082178

	
[image: image96.wmf])

(

x

u

c


	0,0116599
	0,0082666
	0,0089905


The effective degrees of freedom (
[image: image97.wmf]eff

v

) can be evaluated by the Welch-Satterthwaite formula as shown in Equation 3, where 
[image: image98.wmf](

)

y

u

c

 is the combined standard uncertainty, 
[image: image99.wmf](

)

y

u

i

 is the contribution from each uncertainty component.


[image: image100.wmf]å

=

=

N

i

i

i

c

eff

v

y

u

y

u

v

1

4

4

)

(

)

(



(3)

The CSIR NML has chosen a level of confidence of 95,45% for all measurements made.  From t-distribution tables, the appropriate value of 
[image: image101.wmf]k

 (the factor by which the uncertainty is multiplied to ensure a 95,45% degree of certainty in the final result) for 
[image: image102.wmf]eff

v

 is chosen and the standard uncertainty converted to an expanded uncertainty for a 95,45% level of confidence.

The degrees of freedom for each day’s calculation is 
[image: image103.wmf]¥

, since the degrees of freedom for the standard concentrations are 
[image: image104.wmf]¥

 and 
[image: image105.wmf]=

n

10 measurements were made for each concentration standard. The degrees of freedom for the short term stability uncertainty is 2 because this uncertainty was calculated using the measurements from 3 days. Using the values calculated in Table 2, the effective degrees of freedom,
[image: image106.wmf]eff

v

, evaluates to 
[image: image107.wmf]¥

 and the corresponding value for 
[image: image108.wmf]k

is 2 for a 95,45% confidence level. The result and the expanded uncertainty are shown in Table 3.

Table 3: Calculation of expanded uncertainty
	
	Cylinder NMi VSL 152995
	Cylinder NMi VSL 153048
	Cylinder NMi VSL 153190

	
[image: image109.wmf])

(

x

u

c


	0,0116599
	0,0082666
	0,0089905

	
[image: image110.wmf]k


	2
	2
	2

	
[image: image111.wmf])

(

x

U


	0,024
	0,017
	0,018

	Verified concentration (µmol.mol-1)
	5,985
	5,009
	4,015

	% relative expanded uncertainty
	0,39
	0,33
	0,45








� EMBED Equation.3 ���





� EMBED Equation.3 ���





� EMBED Equation.3 ���











� Earlier formats of  Word are also acceptable


� If more than three measurements are taken, please copy and insert a table of the appropriate format as necessary


� The coverage factor shall be based on approximately 95% confidence.


� If more than three measurements are taken, please copy and insert a table of the appropriate format as necessary


� The coverage factor shall be based on approximately 95% confidence.


� The coverage factor shall be based on approximately 95% confidence.











� The coverage factor shall be based on approximately 95% confidence.











� The coverage factor shall be based on approximately 95% confidence.











� The coverage factor shall be based on approximately 95% confidence.











� The coverage factor shall be based on approximately 95% confidence.











� If more than three measurements are taken, please copy and insert a table of the appropriate format as necessary


� The coverage factor shall be based on approximately 95% confidence.


� If more than three measurements are taken, please copy and insert a table of the appropriate format as necessary


� The coverage factor shall be based on approximately 95% confidence.


� If more than three measurements are taken, please copy and insert a table of the appropriate format as necessary


� The coverage factor shall be based on approximately 95% confidence.





PAGE  
30

[image: image120.emf]EUROMET 900: CO in nitrogen

Cylinder nr: D958306

3.9

3.95

4

CSIR NML NMi VSL

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: D958363

4.95

5

5.05

CSIR NML NMi VSL

Concentration 

(µmol/mol)

EUROMET 900: CO in nitrogen

Cylinder nr: D954870

5.95

6

6.05

6.1

CSIR NML NMi VSL

Concentration 

(µmol/mol)

_1231842537.unknown

_1231917855.unknown

_1231918158.unknown

_1231918257.unknown

_1231918297.unknown

_1248690020.unknown

_1248690069.unknown

_1248690113.unknown

_1248690047.unknown

_1231918323.unknown

_1231918272.unknown

_1231918198.unknown

_1231918236.unknown

_1231918180.unknown

_1231918121.unknown

_1231918139.unknown

_1231918097.unknown

_1231917664.unknown

_1231917767.unknown

_1231917800.unknown

_1231917692.unknown

_1231842573.unknown

_1231842657.unknown

_1231917636.unknown

_1231842699.unknown

_1231842712.unknown

_1231842674.unknown

_1231842617.unknown

_1231842636.unknown

_1231842590.unknown

_1231842553.unknown

_1188197196.unknown

_1188197200.unknown

_1188197206.unknown

_1202207308.unknown

_1209281425.unknown

_1231842458.unknown

_1202207324.unknown

_1202124676.unknown

_1202125095.unknown

_1202125112.unknown

_1188197207.unknown

_1188197202.unknown

_1188197205.unknown

_1188197201.unknown

_1188197198.unknown

_1188197199.unknown

_1188197197.unknown

_1037607810.unknown

_1188197194.unknown

_1188197195.unknown

_1037607833.unknown

_1035982512.unknown

_1035983126.unknown

_1035983149.unknown

_1035981924.unknown

