Comparison of material characterisation for semiconductors using resonators, free-space, and on-wafer techniques

TECHNICAL PROTOCOL

draft version

23IND10

OnMicro

April 2025

CONTENTS

1.	Introduction	3

2.	Travelling Standards	3

3.	Participating Laboratories	4

4.	Time Schedule	5

5.	Transport Case	7

6.	Transportation of Travelling Standard	7

6.1.	Damage or loss of samples (travelling standards)	7

6.2.	Financial aspects	7

7.	Measurement Quantities and Points	7

8.	Measurement Instructions	8

8.1.	Precautions	8

8.2.	Before the Measurements	8

8.3.	Environmental Conditions	8

8.4.	Method of measurement	8

9.	Measurement Uncertainty	8

10.	Reporting of Results	8

11.	Final Report of the Comparison	9

12.	References	10

# Introduction

This pilot study is part of the European partnership for metrology (EPM) project 23IND10 OnMicro [1], where different methods for accurate measurement of material properties have been developed (i.e., permittivity and loss tangent) for a variety of materials including semiconductors, thin-films, and novel 2D materials. Precise knowledge of the properties of such materials is crucial for the development of circuits and systems for emerging applications such as 6G communications, automotive radar sensors, and wearable electronics. To meet the demand of such applications exploiting various ranges of the electromagnetic spectrum, these material properties (which are typically frequency-dependent) require characterisation at microwave, millimetre-wave, and terahertz frequencies. A set of reference materials will be measured by different VNA-based setups. The measurements will be analysed with different material parameter extraction methods to study and compare different algorithms.

# Travelling Standards

The list of samples to be used is given in Tab. 1. These samples correspond to dielectric samples deposited to a transmission line.

Tab. 1   List of samples to be used for the comparison

| sample | material | nominal ε | shape/size (mm) | nominal thickness (μm) |
| --- | --- | --- | --- | --- |
| 1 | High-Resistivity Silicon | 11.7 | ø76.2 | 275 |
| 2 | Gallium Arsenide | 12.9 | ø76.4 | 625 |

# Participating Laboratories

The pilot institute for this comparison (pilot study) is NPL. The contact details of the coordinator are given below.

| Pilot Institute : | NPL |  |
| --- | --- | --- |
| Coordinator : | Xiaobang Shang +44 (0) 20 8943 6643 xiaobang.shang@npl.co.uk | National Physical Laboratory Hampton Road Teddington Middlesex, UK, TW11 0LW |

The participating institutes and contact persons with their addresses are given in Tab. 2.

Tab. 2   Participating laboratories

| Country | Institute | Acronym | Shipping Address | Contact Person |
| --- | --- | --- | --- | --- |
| UK | National Physical Laboratory | NPL | National Physical Laboratory Hampton Road Teddington Middlesex, UK TW11 0LW | Xiaobang Shang +44 (0) 20 8943 6643 xiaobang.shang@npl.co.uk |
| Poland | Central Office of Measures | GUM | GUM Laboratory of Electricity and Magnetism Elektoralna 2 00-139 Warszawa Poland | Marcin Wojciechowski +48 (0) 22 581 9303 marcin.wojciechowski@gum.gov.pl |
| France | Laboratoire national de métrologie et d'essais | LNE | LNE 29 avenue Roger Hennequin 78197 Trappes Cedex France | Djamel Allal +33 (0) 130 692 150 djamel.allal@lne.fr |
| Belgium | Keysight Technologies Belgium | Keysight BE | Wingepark 51 3110 Rotselaar Belgium | Robin Schmidt robin.schmidt@keysight.com |
| Germany | Physikalisch-Technische Bundesanstalt | PTB | Physikalisch-Technische Bundesanstalt (PTB) Bundesallee 100 38116 Braunschweig Germany | Gia Ngoc Phung +49 531 592 2226 gia.phung@ptb.de |
| USA | Cornell University | Cornell | 424 Philips Hall School of Electrical and Computer Engineering Cornell University, Ithaca, NY 14853 | Jim Hwang jch263@cornell.edu |
| France | Centre National de la Recherche Scientifique | CNRS | Cité Scientifique Avenue Henri Poincaré CS 60069 59 652 Villeneuve d'Ascq Cedex, France | Kamel Haddadi kamel.haddadi@univ-lille.fr |

# Time Schedule

The time schedule for the comparison is given in Tab. 3. The circulation of travelling standard will be organized as a ring style. Each laboratory will have 6 weeks to carry out the measurements. Any deviation in the agreed plan should be approved by the pilot institute.

Once a participant is ready with the measurement of the travelling standard, each laboratory will arrange the delivery of the equipment to the next participant in the schedule list.

The travelling standard should be checked for damage immediately after reception. In case some damage may be detected, the pilot laboratory should be informed immediately.

Tab. 3   Time schedule of the comparison

| Acronym of Institute | Country | Measurements | Transportation to the next participant due (not later than): |
| --- | --- | --- | --- |
| UK | NPL | April 2025 | 30 April 2025 |
| France | LNE | May 2025 | 31 May 2025 |
| France | CNRS | June/July 2025 | 31 July 2025 |
| Germany | PTB | August 2025 | 31 August 2025 |
| Poland | GUM | September 2025 | 30 September 2025 |
| Belgium | Keysight | October 2025 | 30 October 2025 |
| USA | Cornell | November 2025 / January 2026 | 31 January 2026 |

# Transport Case

The travelling standard and miscellaneous equipment will travel in a well-protected package whose dimensions are approximately 30 cm length x 30 cm width x 20 cm height) and a weight, including the travelling standards, of approx. 2 kg. The transport case can easily be opened for customs inspection.

# Transportation of Travelling Standard

The comparison will be organised in one ring loop.

Each participant laboratory is responsible for the transportation of the travelling standard to the next laboratory. The cost of transportation will be paid by the laboratory which sends the travelling standard. Each participant will be responsible for the transportation expenses to the next participant laboratory in the list.

The estimated cost of the travelling standard and miscellaneous equipment (for the carrier company to include a transportation insurance covering any damage caused to the equipment) is 1000 EUR.

After arrival in the participant’s laboratory, the standard and accompanying instrumentation should be checked for any damage and allowed to stabilise in a temperature and, possibly, humidity-controlled room.

Each participant will have one week to allow for successful transportation from the preceding participant in the list and one week available for measurement. This includes the measurements and the stabilisation of the standard.

## Damage or loss of samples (travelling standards)

In case of any damage of the samples, the participating institutes will decide further steps (e.g., limited number of samples, limited number of participating institutes).

## Financial aspects

Each participant laboratory is responsible for costs for the transportation expenses to the following laboratory, measurements as well as any damage that may occur within its country.

The overall costs for the organisation of the comparison are covered by the pilot laboratory. The pilot laboratory has no insurance for any loss or damage of the travelling standard, except for the insurance related to the delivery to the first participant in the list (first transportation of the standards).

# Measurement Quantities and Points

The quantity to be measured is the complex permittivity (real and imaginary part) or alternatively the real part of the dielectric constant and loss tangent. Each participant should characterise as many samples as possible depending on the used measurement setup. The frequency range should be chosen based on the capabilities of particular setups.

Also the following information should be given:

- Ambient temperature,

- Ambient relative humidity.

No correction will be applied for the ambient temperature and relative humidity.

# Measurement Instructions

## Precautions

- Any precautions or useful information that should be taken into account (information from the sample owner).

- Avoid extreme temperature, humidity or pressure changes as well as violent impacts.

## Before the Measurements

- It should be allowed to stabilize in a temperature and humidity-controlled environment for at least 1 day before commencing measurements.

## Environmental Conditions

- The ambient temperature and humidity should be measured. No corrections will be performed for temperature and humidity effects.

- Preferably, the measurements should be carried out at the ambient conditions given below:

Temperature	: (23 ± 1) °C

Relative humidity	: (45 ± 15) %

## Method of measurement

Each participant institute may use its own measurement method and participant should give its measurement method in the comparison report.

# Measurement Uncertainty

The uncertainty of measurement must be calculated according to Ref. [2] or Ref [3] for a confidence level of 95 % (in the case of infinite degrees of freedom or an assumed Gaussian probability density function, this corresponds to a coverage factor k=2).

All contributions to the measurement uncertainty, or at least the most relevant ones, should be listed in the measurement report submitted by each participant.

# Reporting of Results

The results should be sent to the pilot laboratory within 20 days of completing the measurements.

The report must contain at least:

- Details of participating laboratory,

- The date of the measurements,

- A detailed description of the measurement method and system used,

- Calibration of the system (if applicable),

- The samples used in the comparison measurements,

- Information about the samples’ thickness used for calculations (i.e., nominal value, measured independently),

- Software used in the comparison measurements (if used),

- The environmental conditions during the measurements (if applicable):

  - Ambient Temperature

  - Relative Humidity

- Results of measurement;

- Model function of measurement with explanations of the symbols,

- Expanded measurement uncertainty, estimated for a confidence level of 95 % (if possible).

# Final Report of the Comparison

CMI and METAS with help of the pilot laboratory is responsible for the preparation of a comparison report.

The draft version of the comparison report will be issued by CMI with cooperation with the pilot laboratory of all measurement reports from the participants. The draft report will be sent to the participants for discussion, amendment if applicable, and eventually for approval. Due to the complexity of measurement setups and relative novelty of the material properties measurement methods, this comparison is considered as a pilot study and as such will be submitted to EURAMET. No reference value from participants’ data will be calculated.

# References

- On-wafer microwave metrology for future industrial applications (OnMicro), https://www.ptb.de/epm2023/onmicro/home [online] Accessed 20 February 2025

- BIPM, IEC, IFCC, ILAC, ISO, IUPAC, IUPAP, and OIML. Evaluation of measurement data — Guide to the expression of uncertainty in measurement. Joint Committee for Guides in Metrology, JCGM 100:2008. URL: https://www.bipm.org/documents/20126/2071204/JCGM_100_2008_E.pdf

- BIPM, IEC, IFCC, ILAC, ISO, IUPAC, IUPAP, and OIML. Evaluation of measurement data — Supplement 1 to the “Guide to the expression of uncertainty in measurement” — Propagation of distributions using a Monte Carlo method. Joint Committee for Guides in Metrology, JCGM 101:2008. URL: https://www.bipm.org/documents/20126/2071204/JCGM_101_2008_E.pdf