EURAMET.QM-S10/Euramet project 1274

Supplementary comparison of preparative capabilities for automotive gas mixtures

Bratislava, SMU

2016

Final Report

M. Vaľková1, Z. Ďurišová1, Zs. Nagyné Szilágyi2, T. Büki2, J. Fükő2

1Slovak Institute of Metrology (SMU), Karloveská 63, SK-842 55, Bratislava, Slovakia.

2Magyar Kereskedelmi Engedélyezési Hivatal (Hungarian Trade Licensing Office) MKEH, Németvölgyi út 37-39, H-1124 Budapest, Hungary

Field

Amount of substance: gas analysis

Subject

Comparison in the field of automotive gas mixtures

Table of contents

Field.......................................................................................................................................2

Subject...................................................................................................................................2

Table of contents...................................................................................................................2

Introduction...........................................................................................................................3

Participants............................................................................................................................3

Measurement standards.........................................................................................................3

Schedule................................................................................................................................4

Measurement protocol...........................................................................................................4

Measurement equation.......................................................................................................... 6

Supported CMC claims..........................................................................................................6

Degrees of equivalence......................................................................................................... 7

Results................................................................................................................................... 7

Conclusions..........................................................................................................................12

References............................................................................................................................13

Annex 1 Measurement report from MKEH.........................................................................14

Annex 2 Measurement report from SMU............................................................................17

Introduction

This bilateral supplementary preparative comparison involves standard gas mixtures of automotive gas. The type of automotive gas samples in the comparison is the same as in the comparison EURAMET.QM-S4 which was organized in 2009/2010 by VSL. SMU participated in EURAMET.QM-S4 comparison with good results for all analysed components. In actual comparison SMU is the coordinating laboratory, responsible for collecting and reporting measurement results. Mixture prepared gravimetrically by participant is sent to SMU for the analysis and the analysed value of the amount–of–substance fraction in coordinating laboratory is adopted as reference value for this comparison.

Participants

This comparison has two participants SMU and MKEH which are listed in Table 1.

Table 1

| Acronym | Country | Institute |
| --- | --- | --- |
| SMU | SK | Slovak Institute of Metrology, Bratislava, Slovak Republic |
| MKEH | HU | Hungarian Trade Licensing Office, Budapest, Hungary |

Measurement standards

Participants were requested to prepare an automotive mixture with the following nominal composition:

Carbon monoxide			0.02   mol/mol

Carbon dioxide			0.12   mol/mol

Propane				0.001 mol/mol

Nitrogen			           	balance

The calculation of the gas composition and associated uncertainty evaluation should be done in accordance with ISO 6142 [1].

The pressure in the cylinders was approximately 100 bars; SMU used the cylinder of 5 dm3 nominal and MKEH used the cylinder of 4.67 dm3 nominal. Impurities of parent gases used for preparation were described in the participant measurement reports (see Annex 1 and Annex 2).

MKEH mixture was prepared in the three steps:

- premixture: 0.005 mol/mol propane/N2,

- premixture: 0.1 mol/mol CO/N2,

- 1. premixture, 2. premixture, pure CO2.

SMU mixture was prepared in the two steps:

- premixture: 0.025 mol/mol propane/N2,

- 1. premixture, pure CO, pure CO2.

Schedule

Measurements on gas standard received by the SMU were carried out at April 2014.

Measurement protocol

The submitted reference gas mixtures were measured three times on three different days against SMU 5 PSM´s (primary standard gas mixtures): 0726E_4, 0053F_3, 0088F_3, 0015F_10 and 0059F_4_ in accordance to ISO 6143[2] requirements.

All calculations of reference values were performed with B_LEAST, a computer programme which implements the methodology of ISO 6143:2001, and takes into consideration uncertainties in both axes for regression analysis.

Measurement system Varian 3600 GC used for this comparison was used also in previous comparison EURAMET.QM-S4 for validation SMU PSM´s. Measurement system was well characterized and validated in previous use.

The Varian 3600 GC is equipped with FID and TCD detectors and two sample loops (L1 and L2). As a carrier gas is used He. Settings of the measurements regarding temperature program and flows of the carrier gas (front carrier gas, middle carrier gas and auxiliary carrier gas) are listed in Table 2. Measured chromatograms are showed on Figure 1 and Figure 2.

Table 2 Settings of Varian 3600 GC

| GC | L1; cm3 | L2; cm3 | Oven temperature; °C | Total time; min | F_EFC |  | M_EFC |  | Flow ACG; cm3/min |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Pressure of FCG; kPa | Total flow; ml/min | Pressure of MCG; kPa | Total flow; ml/min |  |
| Varian GC 3800 | 0.25 | 0.25 | 95 | 9.00 | 172.4 | 200 | 479.2 | 100 | 30 |

All measurements were done in automatic way using selector gas valve. Before entering sample loops all gas mixtures went through a mass flow controller and pressure controller for regulation. Measurement method with 6 automated runs was used. All runs in first, third, fifth measurement sequence had rising molar fraction, second, fourth, sixth processed in reverse order. From each run was made one calibration curve with sample signals. Data were subjected to the b_least program (weighted least square regression). As a result from measurement sequence the average value of amount–of–substance fractions was taken. From repetitions of the signal measurements yi, the standard uncertainties u(yi) were calculated.

Figure 1 Chromatogram TCD

CO

CO2

N2

N2

Figure 2 Chromatogram FID

Propane

Measurement equation

The results from the gas chromatographic measurements measured by the pilot laboratory were taken as a reference values for this comparison x0 [3].

For quadratic type of the function is:

(1)

Where:

a0, a1, a2	are the regression coefficients

y 		is GC response, the area of the peak

x		amount–of–substance fraction.

At convergence:

(2)

The uncertainty associated with x0 is computed from [6]:

(3)

Where

(4)

Va denotes the covariance matrix associated with the parameter vector a. For a quadratic polynomial, m = 2.

Supported CMC claims

Components and ranges supported by this supplementary comparison are presented in Table 3.

Table 3

| Component | Supported amount-of-substance fraction range Δx (mmol/mol) |
| --- | --- |
| Carbon monoxide | 5 – 35 |
| Carbon dioxide | 60 – 140 |
| Propane | 0.1 – 2 |

Degrees of equivalence

The analysis function and recorded analyser responses and uncertainties were used to

determine amount–of–substance fractions (x0) and uncertainties u(x0). A degree of equivalence (d) is defined as the difference of a measurement result of the pilot laboratory and NMIs assigned values amount–of–substance fractions (xi) from gravimetric preparation.

(5)

where is :

d	degree of equivalence

x0	pilot laboratory measurement value

xi	value reported by participant from gravimetric preparation

and its associated uncertainty is defined as

(6)

where is:

u(d)			standard uncertainty of the degree of equivalence

u(x0)			standard uncertainty of the pilot laboratory measurement result

u(xi)			standard uncertainty participant´s reported value

Then expanded uncertainty (k=2) is

(7)

In previous preparative comparisons the KCRV was calculated from the consensus straight line obtained from calibrating the gas chromatograph with the submitted mixtures. This approach is in this comparison not possible, because it is bilateral comparison of the very limited data. As an alternative approach the gas chromatograph is calibrated with the PSMs (Primary Standard gas Mixtures) from the coordinating laboratory. In this case the analysed value of the amount–of–substance fraction is adopted as the reference value for each component in each mixture [3].

Results

Results for Carbon monoxide

The gravimetric data of SMU PSMs used for calibrating the GC Varian for Carbon monoxide and measured signals with associated standard uncertainties are shown in Table 4. These data are fitted with a quadratic polynomial.

Table 4 Calibration mixtures for Carbon monoxide

| Mixture | x (mol/mol) | ux (mol/mol) | y (count) | uy (count) |
| --- | --- | --- | --- | --- |
| 0726E_4 | 0.0050938 | 0.0000017 | 1505 | 0.9 |
| 0053F_3 | 0.0081706 | 0.0000035 | 2454 | 1.58 |
| 0088F_3 | 0.0100499 | 0.0000051 | 3049 | 2.51 |
| 0015F_10 | 0.0183856 | 0.0000038 | 5640 | 5.31 |
| 0059F_4 | 0.035085 | 0.000010 | 10912 | 6.62 |

Table 5 Regression coefficients for Carbon monoxide

| Coefficient | Value | u |
| --- | --- | --- |
| a0(count) | -69.4 | 7.95 |
| a1(count) | 3.08x106 | 1.371x103 |
| a2(count) | 1.30 x106 | 3.96x105 |

Using the calibration curve, the reference values have been obtained by analysing the mixtures.

The results are given in Table 6.

Table 6 Reference values for Carbon monoxide

| Mixture | y count | uy count | xi mmol/mol | x0 mmol/mol | u(x0) mmol/mol | u(x0)/x0 % rel. | ∆x mmol/mol | ∆x/x0 % rel. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 6113 | 3.17 | 19.974 | 19.991 | 0.023 | 0.11 | 0.017 | 0.09 |
| SMU0035F2 | 6328 | 4.54 | 20.528 | 20.539 | 0.025 | 0.12 | 0.011 | 0.05 |

The degrees–of–equivalence have been computed as described in equation 5 and are given in Table 7. The degrees–of–equivalence are shown in figure 3. All laboratories have satisfactory results for Carbon monoxide.

Table 7 Degrees–of–equivalence for Carbon monoxide (mmol/mol)

| Mixture | x0 | u(x0) | xi | u(xi) | d | U(d) |
| --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 19.991 | 0.023 | 19.974 | 0.050 | 0.017 | 0.110 |
| SMU0035F2 | 20.539 | 0.025 | 20.528 | 0.025 | 0.011 | 0.071 |

Figure 3 Degrees –of-equivalence for Carbon monoxide

Results for Carbon dioxide

The gravimetric data of SMU PSMs used for calibrating the GC Varian for Carbon dioxide and measured signals with associated standard uncertainties are shown in Table 8. These data are fitted with a quadratic polynomial.

Table 8 Calibration mixtures for Carbon dioxide

| Mixture | x (mol/mol) | ux (mol/mol) | y(count) | uy(count) |
| --- | --- | --- | --- | --- |
| 0726E_4 | 0.060612 | 0.000016 | 24744 | 88.37 |
| 0053F_3 | 0.08023 | 0.000025 | 33025 | 34.24 |
| 0088F_3 | 0.100364 | 0.000016 | 41321 | 23.06 |
| 0015F_10 | 0.111667 | 0.000021 | 46084 | 23.35 |
| 0059F_4 | 0.14132 | 0.000012 | 58442 | 40.3 |

Table 9 Regression coefficients for Carbon dioxide

| Coefficient | Value | u |
| --- | --- | --- |
| a0(count) | -4.12E+02 | 4.48E+02 |
| a1(count) | 4.16E+05 | 8.85E+03 |
| a2(count) | 3.75E+03 | 4.25E+04 |

Using the calibration curve, the reference values have been obtained by analysing the mixtures.

The results are given in Table 10.

Table 10 Reference values for Carbon dioxide

| Mixture | y count | uy count | xi mmol/mol | x0 mmol/mol | u(x0) mmol/mol | u(x0)/x0 % rel. | ∆x mmol/mol | ∆x/x0 % rel. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 49567 | 32.00 | 120.08 | 120.005 | 0.113 | 0.09 | -0.075 | -0.062 |
| SMU0035F2 | 49479 | 34.99 | 119.93 | 119.840 | 0.118 | 0.10 | -0.090 | -0.075 |

The degrees–of–equivalence have been computed as described in equation 5 and are given in Table 11. The degrees–of–equivalence are shown in figure 4. All laboratories have satisfactory results for Carbon dioxide.

Table 11 Degrees–of–equivalence for Carbon dioxide (mmol/mol)

| Mixture | x0 | u(x0) | xi | u(xi) | d | U(d) |
| --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 120.005 | 0.113 | 120.08 | 0.35 | -0.075 | 0.74 |
| SMU0035F2 | 119.840 | 0.118 | 119.93 | 0.12 | -0.090 | 0.34 |

Figure 4 Degrees –of-equivalence for Carbon dioxide

Results for Propane

The gravimetric data of SMU PSMs used for calibrating the GC Varian for Propane and measured signals with associated standard uncertainties are shown in Table 12. These data are fitted with a quadratic polynomial.

Table 12 Calibration mixtures for Propane

| Mixture | x (mol/mol) | ux (mol/mol) | y(count) | uy(count) |
| --- | --- | --- | --- | --- |
| 0726E_4 | 0.00020027 | 0.000000098 | 54381 | 53.65 |
| 0053F_3 | 0.00041345 | 0.00000021 | 114360 | 127.76 |
| 0088F_3 | 0.00060535 | 0.00000025 | 169500 | 148.55 |
| 0015F_10 | 0.00095572 | 0.00000050 | 269780 | 144.39 |
| 0059F_4 | 0.0020377 | 0.00000069 | 591370 | 833.7 |

Table 13 Regression coefficients for Propane

| Coefficient | Value | u |
| --- | --- | --- |
| a0(count) | -1.587E+03 | 2.95E+02 |
| a1(count) | 2.79E+08 | 1.18E+06 |
| a2(count) | 5.97E+09 | 7.03E+08 |

Using the calibration curve, the reference values have been obtained by analysing the mixtures.

The results are given in Table 14

.

Table 14 Reference values for Propane

| Mixture | y count | uy count | xi mmol/mol | x0 mmol/mol | u(x0) mmol/mol | u(x0)/x0 % rel. | ∆x mmol/mol | ∆x/x0 % rel. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 283950 | 303.5 | 0.9992 | 1.0032 | 0.0018 | 0.18 | 0.0041 | 0.40 |
| SMU0035F2 | 290450 | 348.23 | 1.0253 | 1.0267 | 0.0021 | 0.20 | 0.0014 | 0.14 |

The degrees–of–equivalence have been computed as described in equation 5 and are given in Table 15. The degrees–of–equivalence are shown in figure 5. All laboratories have satisfactory results for Propane.

Table 15 Degrees–of–equivalence for Propane (mmol/mol)

| Mixture | x0 | u(x0) | xi | u(xi) | d | U(d) |
| --- | --- | --- | --- | --- | --- | --- |
| OMH230 | 1.0032 | 0.0018 | 0.9992 | 0.0042 | 0.0041 | 0.0091 |
| SMU0035F2 | 1.0267 | 0.0021 | 1.0253 | 0.0017 | 0.0014 | 0.0054 |

Figure 5 Degrees –of-equivalence for Propane

Conclusions

The reported results for Carbon monoxide (Figure 3) agree with the reference value within 0.10 % relative. For Carbon dioxide (Figure 4) results agree with the reference value within 0.10 % relative and for Propane (Figure 5) results agree with the reference value within 0.40 % relative.

The agreement of the results in this supplementary comparison is good. All the results with their reported uncertainties are in agreement with the reference values for the participants. SMU participated in the previous preparative comparison organised within EURAMET in this field. Both laboratories have established facilities for automotive emission gas analysis, and have existing claims for their Calibration and Measurement Capabilities (CMCs) for automotive gas mixtures [5].

References

[1] ISO 6142: Gas Analysis – Preparation of calibration gas mixtures – Gravimetric method, 2nd ed., Switzerland, 2001;

[2] ISO 6143: Gas Analysis – Determination of the composition of calibration gas mixtures, Comparison Methods, 2nd ed., Switzerland, 2001;

[3] Veen van der, A.M.H. - van Wijk, J.I.T. – van Otterloo, R. P. et al: EURAMET.QM-S4, Final report, Metrologia 48 (2011), Tech. Suppl., 08016.

[4] Veen van der, A.M.H. - Oudwater, R. - Ziel, P.R. - van Wijk, J.I.T. et al: CCQM key comparison CCQM-K3 of measurements of CO, CO2, and C3H8 in N2, Final report, 2002 Metrologia 39 (2002).

[5] http://kcdb.bipm.org/AppendixC/

[6] Milton M.J.T., Harris P.M., Smith I.M., Brown A.S., Goody B.A., “Implementation of a generalised leastsquares method for determining calibration curves from data with general uncertainty structures”,Metrologia 43 (2006), pp. S291-S298.

Annex 1

EURAMET 1274 _: Comparison on automotive emission gases

1. Laboratory:  MKEH   Hungarian Trade Licensing Office

Metrology Authority

Division of Radiation Physics and Chemistry

2. Cylinder Details:

| Date of mixture preparation | 16-01-2014 |
| --- | --- |
| Volume (L) | 4.67 |
| Total Pressure (bar) | 100 |
| Connection type (e.g. DIN1, BS14 etc.) | DIN10 |

3. Source of co, co2, propane, nitrogen:

CO 4.7 from Messer Hungary

CO2 4.8 from SIAD Italy

Propane 3.5 from Messer Germany

N2 5.0 from Messer Hungary

4. Purity table  of pure co: (MP1292)

| Component | Method | Mole fraction (mol/mol) | Uncertainty (mol/mol) |
| --- | --- | --- | --- |
| CO | specifications | 0.9999895 | 0.0000061 |
| O2 | specifications | 0.0000015 | 0.0000009 |
| N2 | specifications | 0.0000025 | 0.0000015 |
| H2O | specifications | 0.0000015 | 0.0000009 |
| H2 | specifications | 0.0000005 | 0.0000003 |
| CnHm | specifications | 0.0000010 | 0.0000006 |
| Ar | specifications | 0.0000035 | 0.0000020 |

5. Purity table  of pure co2: (38458)

| Component | Method | Mole fraction (mol/mol) | Uncertainty (mol/mol) |
| --- | --- | --- | --- |
| CO2 | specifications | 0.9999879 | 0.0000070 |
| O2 | specifications | 0.0000015 | 0.0000009 |
| N2 | specifications | 0.0000050 | 0.0000029 |
| H2O | specifications | 0.0000025 | 0.0000015 |
| H2 | specifications | 0.00000005 | 0.00000003 |
| CnHm | specifications | 0.0000030 | 0.0000018 |
| CO | specifications | 0.00000005 | 0.00000003 |

6. Purity table  of pure propane: (2476B)

| Component | Method | Mole fraction (mol/mol) | Uncertainty (mol/mol) |
| --- | --- | --- | --- |
| propane | specifications | 0.9997675 | 0.0001343 |
| O2 | specifications | 0.0000050 | 0.0000029 |
| N2 | specifications | 0.0000200 | 0.0000116 |
| H2O | specifications | 0.0000025 | 0.0000015 |
| CnHm | specifications | 0.0002000 | 0.0001155 |
| CO2 | specifications | 0.0000050 | 0.0000029 |

7. Purity table  of pure nitrogen:

| Component | Method | Mole fraction (mol/mol) | Uncertainty (mol/mol) |
| --- | --- | --- | --- |
| N2 | specifications | 0.9999975 | 0.0000015 |
| O2 | specifications | 0.0000010 | 0.0000006 |
| H2O | specifications | 0.0000015 | 0.0000009 |
| CnHm | specifications | 0.00000005 | 0.00000003 |

8. Preparation of final mixture:

| Parent gases | x(grav+pur) (mol/mol) | u(x) (mol/mol) |
| --- | --- | --- |
| (Propane) OMH222 | 0.0058594 | 0.0000017 |
| (CO) OMH274 | 0.099988 | 0.00045 |
| (CO2) | 0.9999879 | 0.0000140 |
| (N2) | 0.9999975 | 0.0000029 |

9. Purity table  of final mixture: (OMH230)

| Component | x(grav+pur) (mol/mol) | u(x) (mol/mol) |
| --- | --- | --- |
| Ar | 0.00000007 | 0.00000004 |
| CO | 0.01997360 | 0.00000012 |
| CO2 | 0.12008155 | 0.00000084 |
| C3H8 | 0.00099915 | 0.00000013 |
| CnHm | 0.00000062 | 0.00000025 |
| H2 | 0.00000002 | 0.00000001 |
| H2O | 0.00000160 | 0.00000079 |
| N2 | 0.85894234 | 0.00000134 |
| O2 | 0.00000107 | 0.00000053 |

10. Results for co, co2, propane (including gravimetry and verification):

| Component | xprep | uprep | uver | ust | ucert | xi | U(k=2) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CO | 0.0199736 | 0.0000020 | 0.000050 | - | 0.000050 | 0.019974 | 0.00010 |
| CO2 | 0.1200816 | 0.0000080 | 0.00035 | - | 0.00035 | 0.12008 | 0.00070 |
| propane | 0.00099915 | 0.00000034 | 0.0000042 | - | 0.0000042 | 0.0009992 | 0.0000084 |

Annex 2

EURAMET 1274 _: Comparison on automotive emission gases

Participating Institute:

| Institute | Slovak Institute of Metrology |
| --- | --- |
| Address | Karloveska 63 SK-842 55 Bratislava Slovak Republic |
| Contact Person | Ing. M. Valkova, PhD., RNDr. Z. Durisova |
| Telephone | +421 2 60294 211 |
| Fax | +421 2 60294 561 |
| email | valkova@smu.gov.sk, durisova@smu.gov.sk |

1. Cylinder Details

| Date of mixture preparation | 2.10.2013 |
| --- | --- |
| Volume (L) | 5 |
| Total Pressure (bar) | 60 |
| Connection type (e.g. DIN1, BS14 etc.) | DIN1 |

2. Source of CO

The source of nominally pure CO gas was CO 4.7 Messer.

3. Purity Table for nominally pure CO

Complete for all components considered:

| Component | Method | Mole Fraction (mol/mol) | uncertainty (mol/mol) |
| --- | --- | --- | --- |
| CO | rest | 0.99994402 | 0.0000132 |
| Ar | specifications | 0.0000035 | 0.0000020 |
| CO2 | FTIR | 0.00000052 | 0.00000015 |
| H2 | specifications | 0.00000050 | 0.00000029 |
| O2 | specifications | 0.0000025 | 0.0000014 |
| N2 | GC-TCD | 0.000045 | 0.000013 |
| H2O | Dew-point meter | 0.00000387 | 0.00000045 |
| CH4 | FTIR-DL | 0.00000010 | 0.00000006 |

4. Purity Table for nominally pure N2

Complete for all components considered:

N2 BIP 6.0 (AIR Products)

| Component | Method | Mole Fraction (mol/mol) | uncertainty (mol/mol) |
| --- | --- | --- | --- |
| N2 | rest | 0.999998613 | 0.000000129 |
| CO | GC-FID (methaniser) | 0.0000000333 | 0.0000000032 |
| CO2 | GC-FID (methaniser DL) | 0.000000012 | 0.000000007 |
| H2 | specifications | 0.000000025 | 0.000000014 |
| O2 | specifications | 0.000000005 | 0.000000003 |
| H2O | Dew-point meter | 0.000000969 | 0.000000012 |
| CH4 | GC-FID (methaniser DL) | 0.000000023 | 0.000000013 |
| C3H8 | GC-FID (methaniser DL) | 0.000000020 | 0.000000012 |

5. Source of CO2

The source of nominally pure CO2 gas was CO2 5.5 Messer.

6. Purity Table for nominally pure CO2

Complete for all components considered:

| Component | Method | Mole Fraction (mol/mol) | Uncertainty (mol/mol) |
| --- | --- | --- | --- |
| CO2 | rest | 0.99999380 | 0.00000046 |
| CO | GC-FID (methaniser) | 0.00000025 | 0.00000014 |
| O2 | specifications | 0.00000025 | 0.00000014 |
| N2 | GC-TCD | 0.0000035 | 0.0000003 |
| H2O | Dew-point meter | 0.00000204 | 0.00000012 |
| CH4 | GC-FID (methaniser) | 0.00000016 | 0.00000003 |

7. Source of propane

The source of nominally pure propane gas was Propane 3.5 Linde.

8. Purity Table for nominally pure propane

Complete for all components considered:

| Component | Method | Mole Fraction (mol/mol) | uncertainty (mol/mol) |
| --- | --- | --- | --- |
| propane | rest | 0.99977259 | 0.00001268 |
| CO2 | GC-FID (methaniser) | 0.0000025 | 0.0000007 |
| O2 | specifications | 0.000005 | 0.0000014 |
| N2 | GC-TCD | 0.000204 | 0.000012 |
| H2O | specifications | 0.0000050 | 0.0000014 |
| CH4 | GC-FID (methaniser) | 0.0000055 | 0.0000005 |
| C2H6 | GC-FID | 0.000002678 | 0.000000007 |
| propene | GC-FID | 0.0000021 | 0.0000002 |
| Hydrocarbons | GC-FID (as C6+) | 0.000000632 | 0.00000002 |

9. Preparation of final mixture

| Parent gases | x(grav+pur) (mol/mol) | u(x) (mol/mol) |
| --- | --- | --- |
| 0103F_2 (Propane premixture) | 0.0268246 | 0.0000093 |
| CO 4.7 | 0.9999440 | 0.000013 |
| CO2 5.5 | 0.99999380 | 0.00000046 |
| N2 BIP | 0.99999861 | 0.00000013 |

10. Purity Table for final 0035F_2 mixture

Complete for all components considered:

| Component | x(grav+pur) (mol/mol) | u(x) (mol/mol) |
| --- | --- | --- |
| Ar | 0.000000072 | 0.000000041 |
| CO | 0.0205275 | 0.0000147 |
| CO2 | 0.1199261 | 0.0000174 |
| H2 | 0.000000032 | 0.000000013 |
| O2 | 0.000000181 | 0.000000075 |
| N2 | 0.858519 | 0.000021 |
| H2O | 0.000001223 | 0.000000106 |
| CH4 | 0.000000077 | 0.000000004 |
| C2H6 | 0.00000000275 | 0.000000000007 |
| C3H8 | 0.001025262 | 0.000000512 |
| Propene | 0.00000000215 | 0.000000000205 |
| Hydrocarbons | 0.00000000065 | 0.000000000205 |

11. Verification

The mixture was verified on GC Varian using Porapack and molsieve packed columns, 2x sample loops, TCD (for CO, CO2) and FID (for propane) detectors, oven temperature 95 °C, method time 9 min, carrier gas Helium. All measurements were done in automatic way using selector gas valve. Before entering sample loops all gas mixtures went through a mass flow controller and pressure controller for regulation. Calibration standards used for verification were made gravimetrically according to ISO 6142 and ISO 6143 in SMU. Measurement method with 6 automated runs was used. All runs in first, third, fifth measurement sequence had rising molar fraction, second, fourth, sixth processed in reverse order. From each run was made one calibration curve with sample signals. Data were subjected to the b_least program (weighted least square regression).

Results from 0035F_2 CO verification (quadratic curve):

| Cylinder | x(an) | u(an)(k=1) | %u(an) | x(grav) | u(gr)k=1 | ¦x(an)-x(gr)¦ | Criterion 2*odm(u(an)^2+u(gr)^2) | is OK ? |  |  | U(gr+st) k=2 do certifikatu |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0726E_4 | 0,0050995 | 0,0000066 | 0,13% | 0,005094 | 0,0000017 | 5,7E-06 | 1,36E-05 | OK |  |  | 1,81E-05 |
| 0053F_3 | 0,0081544 | 0,0000089 | 0,11% | 0,008171 | 0,0000035 | 1,62E-05 | 1,91E-05 | OK |  |  | 2,94E-05 |
| 0088F_3 | 0,010053 | 0,000010 | 0,10% | 0,01005 | 0,0000051 | 3,1E-06 | 2,25E-05 | OK |  |  | 2,71E-05 |
| 0035F_2 | 0,020553 | 0,000020 | 0,10% | 0,020528 | 0,000015 | 2,5E-05 | 5,0E-05 | OK |  |  | 5,94E-05 |
| 0059F_4 | 0,035076 | 0,000023 | 0,07% | 0,035085 | 0,00001 | 9E-06 | 5,02E-05 | OK |  |  | 9E-05 |

Results from 0035F_2 CO2 verification (linear curve):

| Cylinder | x(an) | u(an)(k=1) | %u(an) | x(grav) | u(gr)k=1 | ¦x(an)-x(gr)¦ | Criterion 2*odm(u(an)^2+u(gr)^2) | is OK ? |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0726E_4 | 0,060625 | 0,000062 | 0,10% | 0,060612 | 0,000016 | 1,3E-05 | 0,000128 | OK |
| 0053F_3 | 0,080212 | 0,000051 | 0,06% | 0,08023 | 0,000025 | 1,8E-05 | 0,000114 | OK |
| 0088F_3 | 0,100365 | 0,000091 | 0,09% | 0,100364 | 0,000016 | 1E-06 | 0,000185 | OK |
| 0035F_2 | 0,11991 | 0,00012 | 0,10% | 0,119926 | 0,000017 | 1,6E-05 | 0,000242 | OK |
| 0059F_4 | 0,14132 | 0,00011 | 0,08% | 0,14132 | 0,000012 | 0 | 0,000221 | OK |

Results from 0035F_2 propane verification (quadratic curve):

| Cylinder | x(an) | u(an)(k=1) | %u(an) | x(grav) | u(gr)k=1 | ¦x(an)-x(gr)¦ | Criterion 2*odm(u(an)^2+u(gr)^2) | is OK ? |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0726E_4 | 0,00020047 | 0,00000025 | 0,12% | 0,0002 | 0,000000098 | 2E-07 | 5,37E-07 | OK |
| 0053F_3 | 0,00041272 | 0,00000034 | 0,08% | 0,000413 | 0,00000021 | 7,3E-07 | 7,99E-07 | OK |
| 0088F_3 | 0,00060477 | 0,00000062 | 0,10% | 0,000605 | 0,00000025 | 5,8E-07 | 1,34E-06 | OK |
| 0035F_2 | 0,0010280 | 0,0000016 | 0,16% | 0,001025 | 0,00000051 | 2,7E-06 | 3,36E-06 | OK |
| 0059F_4 | 0,0020359 | 0,0000020 | 0,10% | 0,002038 | 0,00000069 | 1,8E-06 | 4,23E-06 | OK |

11. RESULTS

The results are presented in following table with data:

xprep	amount of substance fraction , from preparation (mol.mol-1)

uprep	uncertainty of xprep from gravimetrical preparation and purity (mol.mol-1)

uver	uncertainty from verification (mol.mol-1)

ucert	final uncertainty of x (mol.mol-1)

U(k=2)		stated expanded uncertainty of x, at 95% level of confidence (mol.mol-1)

Standard uncertainty of the mixture was calculated with following formula:

| Component | xprep | uprep | uver | ucert | xi | U(k=2) |
| --- | --- | --- | --- | --- | --- | --- |
| CO | 0.0205275 | 0.0000147 | 0.000020 | 0.0000248 | 0.020528 | 0.000050 |
| CO2 | 0.119926 | 0.000017 | 0.00012 | 0.00012 | 0.11993 | 0.00024 |
| propane | 0.00102526 | 0.00000051 | 0.0000016 | 0.0000017 | 0.0010253 | 0.0000034 |