## 35,5 °C
Analysis of comparison measurements at 35,5 °C | [11 empty] | Analysis of comparison measurements at 35,5 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift |  | Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005 |  | b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434 |  | b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434
b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty |  | b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty
Local IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Local IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (without probe cover) (°C) | Transfer BB - Local IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (without probe cover) (°C) | Local BB - Local IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (with probe cover) (°C) | Transfer BB - Local IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (with probe cover) (°C) | Local BB - Local IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 113.7846 | 35.2323 | 35.21 | 0.022332 |  | 35.2323 | 35.21 | 0.022332 | 0 | SI 1st set | 2007-11-14 | 0:06 | 113.7846 | 35.2323 | 35.12 | 0.11233 |  | 35.2323 | 35.12 | 0.11233 | 0 | 1
CH 1st set | 2007-12-10 | 0:07 | 113.7823 | 35.2264 | 35.17 | 0.056403 |  | 35.522 | 35.44 | 0.08197 | -0.025567 | CH 1st set | 2007-12-10 | 0:08 | 113.7823 | 35.2265 | 35.19 | 0.036454 |  | 35.5241 | 35.42 | 0.10413 | -0.067681 | 1
CH 2nd set | 2008-01-07 | 0:06 | 113.7828 | 35.2278 | 35.16 | 0.067794 |  | 35.5252 | 35.43 | 0.09524 | -0.027446 | CH 2nd set | 2008-01-07 | 0:07 | 113.7828 | 35.2277 | 35.2 | 0.027678 |  | 35.5239 | 35.43 | 0.09394 | -0.066262 | 1
TR 1st set | [11 empty] | TR 1st set | 2008-01-30 | 0:19 | 113.7839 | 35.2305 | 35.4 | -0.16946 |  | 35.4804 | 35.19 | 0.29041 | -0.45987 | 1 | ear mode
TR 2nd set | [11 empty] | TR 2nd set | 2008-02-19 | 0:12 | 113.7826 | 35.2273 | 35.4 | -0.17271 |  | 35.7751 | 35.43 | 0.34515 | -0.51786 | 1 | ear mode
NO 1st set | [11 empty] | NO 1st set
NO 2nd set | 2008-03-10 | 0:08 | 113.7853 | 35.2342 | 35.18 | 0.054192 |  | 35.5036 | 35.44 | 0.063565 | -0.009373 | NO 2nd set | 2008-03-10 | 0:07 | 113.7853 | 35.2341 | 35.19 | 0.044097 |  | 35.5037 | 35.44 | 0.063705 | -0.019608 | JV
DE 1st set | [11 empty] | DE 1st set | 2008-06-30 | 0:09 | 113.7805 | 35.2219 | 34.99 | 0.23192 |  | 35.5034 | 35.24 | 0.2634 | -0.031479 |  | tympanic cal
DE 2nd set | [11 empty] | DE 2nd set | 2008-07-07 | 0:09 | 113.7807 | 35.2224 | 34.99 | 0.23244 |  | 35.5064 | 35.24 | 0.26635 | -0.033913 |  | tympanic cal
UK 1st set | 2008-07-18 | 936:33 | 113.7865 | 35.2373 | 36.16 | -0.92273 |  | 35.2154 | 36.13 | -0.91459 | -0.0081416 | UK 1st set | 2008-07-18 | 0:29 | 113.7864 | 35.237 | 35.1075 | 0.12953 |  | 35.2159 | 35.08 | 0.13593 | -0.0063935 | B
UK 2nd set | 2008-07-24 | 0:25 | 113.7864 | 35.2369 | 36.12 | -0.88308 |  | 35.2168 | 36.1 | -0.88317 | 8.7939e-05 | UK 2nd set | 2008-07-24 | 0:24 | 113.7863 | 35.2368 | 35.09 | 0.14679 |  | 35.2171 | 35.06 | 0.15709 | -0.0103 | B
FR 1st set | 2008-10-06 | 0:20 | 113.7842 | 35.2313 | 35.72 | -0.48872 |  | 35.4885 | 35.98 | -0.49155 | 0.0028335 | FR 1st set | 2008-10-06 | 0:20 | 113.7841 | 35.2312 | 34.6 | 0.63119 |  | 35.4893 | 34.8 | 0.68935 | -0.058157 | 1
FR 2nd set | 2008-10-20 | 0:25 | 113.7842 | 35.2313 | 35.66 | -0.4287 |  | 35.4913 | 35.91 | -0.41875 | -0.0099536 | FR 2nd set | 2008-10-20 | 0:20 | 113.7842 | 35.2314 | 34.67 | 0.56136 |  | 35.4925 | 34.9 | 0.5925 | -0.031139 | 1
SI 2nd set | 2008-12-01 | 0:00 | 113.7839 | 35.2371 | 35.19 | 0.0471 |  | 35.2371 | 35.19 | 0.0471 | 0 | SI 2nd set | 2008-12-01 | 0:00 | 113.7839 | 35.2371 | 35.12 | 0.1171 |  | 35.2371 | 35.12 | 0.1171 | 0
Uncertainty budget in calibration of local IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0068633 | 0.00021467 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0099337 | 0.013974 | SI 1st set | 0.026178 | 0.00036686 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.027146 | 0.038364
CH 1st set | 0.01 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01235 | 0.020078 | CH 1st set | 0.02 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.021273 | 0.033153
CH 2nd set | 0.012 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.014019 | 0.022872 | CH 2nd set | 0.017 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01848 | 0.025848
TR 1st set | [11 empty] | TR 1st set | 0.0394 | 0.0015 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.028868 | 0.0023 | 0.049339 | 0.19504
TR 2nd set | [11 empty] | TR 2nd set | 0.02236 | 0.0085 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.028868 | 0.0023 | 0.038105 | 0.18758
NO 1st set | [11 empty] | NO 1st set
NO 2nd set | 0.023396 | 0.00016902 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.024473 | 0.036023 | NO 2nd set | 0.022821 | 0.00033085 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.023925 | 0.039088
DE 1st set | [11 empty] | DE 1st set | 0.0076 | 0.0005 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.010466 | 0.022874
DE 2nd set | [11 empty] | DE 2nd set | 0.0064 | 0.0006 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0096357 | 0.022103
UK 1st set | 0.08 | 0.0003 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.081213 | 0.11528 | UK 1st set | 0.034 | 0.0004 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.036765 | 0.05103
UK 2nd set | 0.07 | 0.0003 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.071384 | 0.10162 | UK 2nd set | 0.026 | 0.0002 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.029522 | 0.042919
FR 1st set | 0.042 | 0.00036 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.042611 | 0.064328 | FR 1st set | 0.0399 | 0.00027 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.040541 | 0.05846
FR 2nd set | 0.043 | 0.00031 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.043596 | 0.062443 | FR 2nd set | 0.0461 | 0.00025 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.046656 | 0.067428
SI 2nd set | 0.0072548 | 0.00030779 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01021 | 0.014367 | SI 2nd set | 0.01916 | 0.00075394 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.020474 | 0.028919
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of local IRET without probe cover compared to local BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0068633 | 0.00021467 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0098283 |  | SI 1st set | 0.026178 | 0.00036686 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.027108
CH 1st set | 0.013 | 0.006 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.01583 |  | CH 1st set | 0.024 | 0.005 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.025428
CH 2nd set | 0.016 | 0.005 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.018072 |  | CH 2nd set | 0.016 | 0.005 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.018072
TR 1st set | [11 empty] | TR 1st set | 0.05871 | 0.06013 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.1887
TR 2nd set | [11 empty] | TR 2nd set | 0.06387 | 0.03333 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.18367
NO 1st set | [11 empty] | NO 1st set
NO 2nd set | 0.014654 | 0.0044737 | 0.0025 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.026434 |  | NO 2nd set | 0.021831 | 0.0029717 | 0.0035 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.03091
DE 1st set | [11 empty] | DE 1st set | 0.0076 | 0.0005 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.020339
DE 2nd set | [11 empty] | DE 2nd set | 0.0063 | 0.0006 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.019892
UK 1st set | 0.078 | 0.0021 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.081821 |  | UK 1st set | 0.025 | 0.0046 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.03539
UK 2nd set | 0.068 | 0.0009 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.072326 |  | UK 2nd set | 0.019 | 0.0018 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.031153
FR 1st set | 0.0446 | 0.0008 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.048192 |  | FR 1st set | 0.03796 | 0.0006 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.042119
FR 2nd set | 0.0408 | 0.0011 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.044705 |  | FR 2nd set | 0.04513 | 0.00061 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.04868
SI 2nd set | 0.0072548 | 0.00030779 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.010108 |  | SI 2nd set | 0.01916 | 0.00075394 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.020423
Note on V38: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W38: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 and 0,012 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,006 added to the total uncertainty
Note on J45: Igor: 0,0003 and 0,012 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,006 added to the total uncertainty
Note on V72: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W72: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W73: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J78: Igor: 0,010 added to the total uncertainty
Note on W78: Igor: 0,006 added to the total uncertainty
Note on J79: Igor: 0,010 added to the total uncertainty
Note on W79: Igor: 0,006 added to the total uncertainty

## 38,0 °C
Analysis of comparison measurements at 38,0 °C | [11 empty] | Analysis of comparison measurements at 38,0 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift |  | Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005 |  | b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434 |  | b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434
b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty |  | b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty
Local IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Local IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (without probe cover) (°C) | Transfer BB - Local IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (without probe cover) (°C) | Local BB - Local IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (with probe cover) (°C) | Transfer BB - Local IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (with probe cover) (°C) | Local BB - Local IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 114.7538 | 37.7296 | 37.72 | 0.0095848 |  | 37.73 | 37.72 | 0.01 | -0.00041522 | SI 1st set | 2007-11-14 | 0:06 | 114.7538 | 37.7296 | 37.65 | 0.079559 |  | 37.73 | 37.65 | 0.08 | -0.000441 | 1
CH 1st set | 2007-12-10 | 0:07 | 114.7509 | 37.7222 | 37.63 | 0.092156 |  | 38.0162 | 37.9 | 0.11616 | -0.023999 | CH 1st set | 2007-12-10 | 0:07 | 114.751 | 37.7223 | 37.67 | 0.052285 |  | 38.0181 | 37.88 | 0.13811 | -0.08583 | 1
CH 2nd set | 2008-01-07 | 0:07 | 114.751 | 37.7224 | 37.62 | 0.10239 |  | 38.0175 | 37.89 | 0.12754 | -0.025152 | CH 2nd set | 2008-01-07 | 0:08 | 114.7511 | 37.7226 | 37.68 | 0.042581 |  | 38.0155 | 37.88 | 0.13545 | -0.092874 | 1
TR 1st set | [11 empty] | TR 1st set | 2008-01-31 | 0:14 | 114.754 | 37.7302 | 37.9 | -0.16984 |  | 38.1527 | 37.78 | 0.37275 | -0.54259 | 1 | ear mode
TR 2nd set | [11 empty] | TR 2nd set | 2008-01-31 | 0:14 | 114.7496 | 37.7188 | 37.9 | -0.18116 |  | 38.3174 | 38 | 0.31735 | -0.49851 | 1 | ear mode
NO 1st set | [11 empty] | NO 1st set | 2008-03-07 | 0:11 | 114.7543 | 37.7308 | 37.66 | 0.070766 |  | 38.001 | 37.87 | 0.13099 | -0.060224 | JV
NO 2nd set | 2008-03-10 | 0:07 | 114.7543 | 37.7309 | 37.67 | 0.060856 |  | 38.004 | 37.9 | 0.10399 | -0.043129 | NO 2nd set | 2008-03-10 | 0:07 | 114.7543 | 37.7309 | 37.66 | 0.070928 |  | 38.0038 | 37.87 | 0.13378 | -0.062847 | JV
DE 1st set | [11 empty] | DE 1st set | 2008-06-30 | 0:10 | 114.7493 | 37.718 | 37.34 | 0.37797 |  | 38.0047 | 37.6 | 0.40465 | -0.026683 |  | tympanic cal
DE 2nd set | [11 empty] | DE 2nd set | 2008-07-07 | 0:11 | 114.7493 | 37.7181 | 37.29 | 0.4281 |  | 38.0032 | 37.55 | 0.45315 | -0.025054 |  | tympanic cal
UK 1st set | 2008-07-18 | 936:27 | 114.7553 | 37.7335 | 38.76 | -1.0265 |  | 37.7125 | 38.75 | -1.0375 | 0.011075 | UK 1st set | 2008-07-18 | 0:24 | 114.7552 | 37.7332 | 37.66 | 0.073222 |  | 37.7114 | 37.63 | 0.081435 | -0.0082131 | B
UK 2nd set | 2008-07-24 | 0:21 | 114.7552 | 37.7331 | 38.8 | -1.0669 |  | 37.8124 | 38.89 | -1.0776 | 0.010736 | UK 2nd set | 2008-07-24 | 0:17 | 114.7552 | 37.7331 | 37.6 | 0.13307 |  | 37.8115 | 37.66 | 0.15147 | -0.018398 | B
FR 1st set | 2008-10-07 | 0:30 | 114.7534 | 37.7284 | 38.23 | -0.50155 |  | 37.9995 | 38.49 | -0.49055 | -0.011005 | FR 1st set | 2008-10-07 | 0:20 | 114.7534 | 37.7285 | 37.24 | 0.48847 |  | 38.0004 | 37.45 | 0.55045 | -0.061979 | 1
FR 2nd set | 2008-10-21 | 0:20 | 114.7532 | 37.7281 | 38.28 | -0.55186 |  | 38.0006 | 38.55 | -0.5494 | -0.0024638 | FR 2nd set | 2008-10-21 | 0:20 | 114.7533 | 37.7282 | 37.04 | 0.68817 |  | 38.0015 | 37.27 | 0.73155 | -0.043375 | 1
SI 2nd set | 2008-12-01 | 0:00 | 114.7529 | 37.7335 | 37.64 | 0.09355 |  | 37.734 | 37.64 | 0.094 | -0.00045 | SI 2nd set | 2008-12-01 | 0:00 | 114.7529 | 37.7337 | 37.58 | 0.15365 |  | 37.734 | 37.58 | 0.154 | -0.00035
Uncertainty budget in calibration of local IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0065695 | 0.0011418 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0098827 | 0.013902 | SI 1st set | 0.014681 | 0.0010515 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.016427 | 0.023186
CH 1st set | 0.012 | 0.00125 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.014099 | 0.019833 | CH 1st set | 0.026 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.027022 | 0.04011
CH 2nd set | 0.011 | 0.00075 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01322 | 0.022635 | CH 2nd set | 0.025 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026062 | 0.033133
TR 1st set | [11 empty] | TR 1st set | 0.02236 | 0.00148 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.028868 | 0.0023 | 0.037224 | 0.21277
TR 2nd set | [11 empty] | TR 2nd set | 0.03078 | 0.0009 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.028868 | 0.0023 | 0.042799 | 0.1999
NO 1st set | [11 empty] | NO 1st set | 0.027237 | 0.00021668 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.028197 | 0.0408
NO 2nd set | 0.026808 | 0.00050904 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.027788 | 0.043615 | NO 2nd set | 0.020622 | 0.00043464 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.021878 | 0.039336
DE 1st set | [11 empty] | DE 1st set | 0.0089 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011511 | 0.024013
DE 2nd set | [11 empty] | DE 2nd set | 0.0134 | 0.0004 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015262 | 0.02767
UK 1st set | 0.085 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.086153 | 0.11558 | UK 1st set | 0.036 | 0.0004 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.038644 | 0.055367
UK 2nd set | 0.065 | 0.0002 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.0665 | 0.093554 | UK 2nd set | 0.02 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.02444 | 0.039616
FR 1st set | 0.0392 | 0.00036 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.039875 | 0.059631 | FR 1st set | 0.0383 | 0.00027 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.038989 | 0.052328
FR 2nd set | 0.0265 | 0.00041 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.027489 | 0.039612 | FR 2nd set | 0.02294 | 0.00031 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.024074 | 0.038019
SI 2nd set | 0.0068056 | 0.00060481 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0099944 | 0.01406 | SI 2nd set | 0.01927 | 0.00048936 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.02061 | 0.029111
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of local IRET without probe cover compared to local BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0065695 | 0.0011418 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0097768 |  | SI 1st set | 0.014681 | 0.0010515 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.016363
CH 1st set | 0.01 | 0.007 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.013949 |  | CH 1st set | 0.028 | 0.007 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.029641
CH 2nd set | 0.016 | 0.006 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.018373 |  | CH 2nd set | 0.018 | 0.007 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.020459
TR 1st set | [11 empty] | TR 1st set | 0.10563 | 0.06467 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.20948
TR 2nd set | [11 empty] | TR 2nd set | 0.06387 | 0.07421 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.19527
NO 1st set | [11 empty] | NO 1st set | 0.019601 | 0.0019625 | 0.0025 | 0.021457 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.029488
NO 2nd set | 0.025397 | 0.0018972 | 0.0025 | 0.021457 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.033617 |  | NO 2nd set | 0.024165 | 0.0017923 | 0.0025 | 0.021457 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.03269
DE 1st set | [11 empty] | DE 1st set | 0.0094 | 0.0003 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021074
DE 2nd set | [11 empty] | DE 2nd set | 0.0133 | 0.0004 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.02308
UK 1st set | 0.073 | 0.0013 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.077052 |  | UK 1st set | 0.031 | 0.0022 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.03965
UK 2nd set | 0.061 | 0.0017 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.065804 |  | UK 2nd set | 0.019 | 0.0022 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.031179
FR 1st set | 0.0404 | 0.00105 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.044339 |  | FR 1st set | 0.02975 | 0.0006 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.034901
FR 2nd set | 0.0219 | 0.0011 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.028522 |  | FR 2nd set | 0.02308 | 0.00076 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.029427
SI 2nd set | 0.0068056 | 0.00060481 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0098896 |  | SI 2nd set | 0.01927 | 0.00048936 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.020559
Note on V38: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W38: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W39: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 and 0,012 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,016 added to the total uncertainty
Note on J45: Igor: 0,0003 and 0,012 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,016 added to the total uncertainty
Note on V72: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W72: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W73: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J78: Igor: 0,012 added to the total uncertainty
Note on W78: Igor: 0,020 added to the total uncertainty
Note on J79: Igor: 0,012 added to the total uncertainty
Note on W79: Igor: 0,020 added to the total uncertainty

## 41,0 °C
Analysis of comparison measurements at 41,0 °C | [11 empty] | Analysis of comparison measurements at 41,0 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift |  | Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005 |  | b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  |  | 0.005
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434 |  | b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  |  | 0.0014434
b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty |  | b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  |  | uncertainty
Local IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Local IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (without probe cover) (°C) | Transfer BB - Local IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (without probe cover) (°C) | Local BB - Local IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Local IRET (with probe cover) (°C) | Transfer BB - Local IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Local IRET (with probe cover) (°C) | Local BB - Local IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 115.9156 | 40.7258 | 40.74 | -0.014231 |  | 40.726 | 40.74 | -0.014 | -0.00023112 | SI 1st set | 2007-11-14 | 0:06 | 115.9155 | 40.7256 | 40.7 | 0.025588 |  | 40.726 | 40.7 | 0.026 | -0.00041173 | 1
CH 1st set | 2007-12-10 | 0:06 | 115.9129 | 40.7188 | 40.58 | 0.13875 |  | 41.0083 | 40.84 | 0.16825 | -0.029504 | CH 1st set | 2007-12-10 | 0:07 | 115.9129 | 40.7187 | 40.63 | 0.088699 |  | 41.0099 | 40.81 | 0.19992 | -0.11123 | 1
CH 2nd set | 2008-01-07 | 0:06 | 115.9129 | 40.7187 | 40.59 | 0.12873 |  | 41.0079 | 40.85 | 0.1579 | -0.029175 | CH 2nd set | 2008-01-07 | 0:07 | 115.9129 | 40.7186 | 40.64 | 0.078648 |  | 41.0104 | 40.82 | 0.19043 | -0.11178 | 1
TR 1st set | [11 empty] | TR 1st set | 2008-01-01 | 0:14 | 115.9148 | 40.7236 | 40.97 | -0.2464 |  | 41.5666 | 41.17 | 0.3966 | -0.643 | 1 | ear mode
TR 2nd set | [11 empty] | TR 2nd set | 2008-02-20 | 0:15 | 115.9115 | 40.7152 | 40.9 | -0.18484 |  | 41.343 | 40.9 | 0.443 | -0.62784 | 1 | ear mode
NO 1st set | 2008-03-07 | 0:09 | 115.9163 | 40.7275 | 40.68 | 0.047467 |  | 41.0124 | 40.94 | 0.072405 | -0.024938 | NO 1st set | 2008-03-07 | 0:09 | 115.9163 | 40.7274 | 40.59 | 0.13743 |  | 41.0119 | 40.83 | 0.18193 | -0.044502 | JV
NO 2nd set | 2008-03-10 | 0:11 | 115.9162 | 40.7273 | 40.66 | 0.067285 |  | 41.0153 | 40.9 | 0.11526 | -0.047975 | NO 2nd set | 2008-03-10 | 0:08 | 115.9162 | 40.7274 | 40.63 | 0.097382 |  | 41.0142 | 40.86 | 0.15424 | -0.056858 | JV
DE 1st set | [11 empty] | DE 1st set | 2008-07-01 | 0:10 | 115.9112 | 40.7144 | 40.2 | 0.51436 |  | 41.0051 | 40.46 | 0.5451 | -0.030735 |  | tympanic cal
DE 2nd set | [11 empty] | DE 2nd set | 2008-07-08 | 0:09 | 115.912 | 40.7164 | 40.12 | 0.59643 |  | 41.0037 | 40.39 | 0.6137 | -0.017271 |  | tympanic cal
UK 1st set | 2008-07-21 | 936:23 | 115.9171 | 40.7297 | 42.04 | -1.3103 |  | 40.8054 | 42.11 | -1.3046 | -0.005725 | UK 1st set | 2008-07-21 | 936:20 | 115.9171 | 40.7295 | 40.6 | 0.12949 |  | 40.8045 | 40.64 | 0.16449 | -0.035 | B
UK 2nd set | 2008-07-25 | 0:24 | 115.9174 | 40.7303 | 42.11 | -1.3797 |  | 40.8042 | 42.19 | -1.3858 | 0.0060238 | UK 2nd set | 2008-07-25 | 0:20 | 115.9173 | 40.73 | 40.68 | 0.050001 |  | 40.8032 | 40.71 | 0.0932 | -0.0432 | B
FR 1st set | 2008-10-08 | 0:25 | 115.9153 | 40.725 | 41.32 | -0.59503 |  | 41.0254 | 41.62 | -0.5946 | -0.00043096 | FR 1st set | 2008-10-08 | 0:25 | 115.9153 | 40.7249 | 39.95 | 0.7749 |  | 41.0256 | 40.18 | 0.8456 | -0.070695 | 1
FR 2nd set | 2008-10-22 | 0:25 | 115.9153 | 40.7248 | 41.34 | -0.61519 |  | 41.0203 | 41.62 | -0.59965 | -0.015536 | FR 2nd set | 2008-10-22 | 0:20 | 115.9153 | 40.7248 | 39.83 | 0.89483 |  | 41.0204 | 40.06 | 0.96045 | -0.065623 | 1
SI 2nd set | 2008-12-01 | 0:00 | 115.9149 | 40.731 | 40.58 | 0.151 |  | 40.731 | 40.58 | 0.151 | 0 | SI 2nd set | 2008-12-01 | 0:00 | 115.9149 | 40.731 | 40.52 | 0.211 |  | 40.731 | 40.52 | 0.211 | 0
Uncertainty budget in calibration of local IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0091047 | 0.00032851 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01178 | 0.016596 | SI 1st set | 0.027028 | 0.00032313 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.028043 | 0.039632
CH 1st set | 0.013 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015025 | 0.022189 | CH 1st set | 0.032 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.032875 | 0.041465
CH 2nd set | 0.017 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.018595 | 0.027043 | CH 2nd set | 0.03 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.030931 | 0.03959
TR 1st set | [11 empty] | TR 1st set | 0.05871 | 0.00103 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.028868 | 0.0023 | 0.065817 | 0.20649
TR 2nd set | [11 empty] | TR 2nd set | 0.06959 | 0.00094 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.028868 | 0.0023 | 0.075681 | 0.21342
NO 1st set | 0.016944 | 0.00019671 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.018518 | 0.032637 | NO 1st set | 0.043456 | 0.00034158 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.044071 | 0.057277
NO 2nd set | 0.028004 | 0.00024302 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.028983 | 0.041382 | NO 2nd set | 0.015009 | 0.00043185 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.016769 | 0.032447
DE 1st set | [11 empty] | DE 1st set | 0.013 | 0.0007 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015008 | 0.027585
DE 2nd set | [11 empty] | DE 2nd set | 0.0132 | 0.0004 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015171 | 0.026714
UK 1st set | 0.083 | 0.0003 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.084195 | 0.11823 | UK 1st set | 0.019 | 0.0006 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.023688 | 0.039314
UK 2nd set | 0.078 | 0.0003 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.079271 | 0.11075 | UK 2nd set | 0.024 | 0.0002 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.027853 | 0.045252
FR 1st set | 0.0469 | 0.00035 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.047492 | 0.072396 | FR 1st set | 0.0603 | 0.00035 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.060762 | 0.089127
FR 2nd set | 0.0344 | 0.00045 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.035204 | 0.052692 | FR 2nd set | 0.03703 | 0.00021 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.037776 | 0.05634
SI 2nd set | 0.01118 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013445 | 0.018959 | SI 2nd set | 0.024942 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026036 | 0.036792
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of local IRET without probe cover compared to local BB | [11 empty] | Uncertainty budget in calibration of local IRET with probe cover compared to local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0091047 | 0.00032851 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.011691 |  | SI 1st set | 0.027028 | 0.00032313 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.028005
CH 1st set | 0.014 | 0.005 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.016327 |  | CH 1st set | 0.023 | 0.008 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.02527
CH 2nd set | 0.018 | 0.004 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.019636 |  | CH 2nd set | 0.023 | 0.006 | 0.0022361 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.02471
TR 1st set | [11 empty] | TR 1st set | 0.08751 | 0.04589 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.19572
TR 2nd set | [11 empty] | TR 2nd set | 0.08335 | 0.06579 | 0.151 | 0.07 | 0.002 |  |  | 0.028868 | 0.0023 | 0.19955
NO 1st set | 0.01469 | 0.0031728 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.026876 |  | NO 1st set | 0.028887 | 0.0027382 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.036584
NO 2nd set | 0.019084 | 0.0034361 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.029537 |  | NO 2nd set | 0.016376 | 0.0026454 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.027778
DE 1st set | [11 empty] | DE 1st set | 0.0134 | 0.0007 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.023145
DE 2nd set | [11 empty] | DE 2nd set | 0.0113 | 0.0004 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021989
UK 1st set | 0.079 | 0.0016 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.082999 |  | UK 1st set | 0.018 | 0.0039 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.031376
UK 2nd set | 0.073 | 0.0026 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.077337 |  | UK 2nd set | 0.025 | 0.0013 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.035664
FR 1st set | 0.0515 | 0.0009 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.054642 |  | FR 1st set | 0.0626 | 0.0005 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.065205
FR 2nd set | 0.0347 | 0.00059 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.039206 |  | FR 2nd set | 0.0376 | 0.0009 | 0.0065 | 0.009 | 0.002 | 0.001 |  | 0.0029 | 0.014 | 0.0418
SI 2nd set | 0.01118 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.013367 |  | SI 2nd set | 0.024942 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.025996
Note on V38: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W38: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W39: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 and 0,018 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,023 added to the total uncertainty
Note on J45: Igor: 0,0003 and 0,018 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,023 added to the total uncertainty
Note on V72: ahmetd: ume local IRET had 1 digit reading. Therefore 0.05/sqrt(3)
Note on W72: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W73: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J78: Igor: 0,021 added to the total uncertainty
Note on W78: Igor: 0,020 added to the total uncertainty
Note on J79: Igor: 0,021 added to the total uncertainty
Note on W79: Igor: 0,020 added to the total uncertainty