## 35,5 °C
Analysis of comparison measurements at 35,5 °C | [11 empty] | Analysis of comparison measurements at 35,5 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison |  |  |  |  | max drift |  |  | Transfer PRT (LMK 1204) coefficients |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  | 0.005 | °C |  | b0 | -243.7645 |  | b0 | -243.7648 |  |  |  |  | 0.005 | °C
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  | 0.0014434 | °C |  | b1 | 2.3284 |  | b1 | 2.3248 |  |  |  |  | 0.0014434 | °C
b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  | uncertainty |  |  | b2 | 0.0010858 |  | b2 | 0.0011031 |  |  |  |  | uncertainty
Transfer IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Transfer IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Transfer BB - Transfer IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Local BB - Transfer IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Transfer BB - Transfer IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Local BB - Transfer IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 113.7846 | 35.2323 | 35.21 | 0.022332 |  | 35.2323 | 35.21 | 0.022332 | 0 | SI 1st set | 2007-11-14 | 0:06 | 113.7846 | 35.2323 | 35.12 | 0.11233 |  | 35.2323 | 35.12 | 0.11233 | 0 | 1
CH 1st set | 2007-12-10 | 0:06 | 113.7824 | 35.2267 | 35.19 | 0.036686 |  | 35.5196 | 35.47 | 0.049625 | -0.012939 | CH 1st set | 2007-12-10 | 0:06 | 113.7823 | 35.2265 | 35.27 | -0.043455 |  | 35.5228 | 35.51 | 0.01282 | -0.056275 | 1
CH 2nd set | 2008-01-07 | 0:06 | 113.7829 | 35.2281 | 35.21 | 0.01809 |  | 35.521 | 35.48 | 0.041 | -0.02291 | CH 2nd set | 2008-01-07 | 0:06 | 113.7829 | 35.2279 | 35.28 | -0.052052 |  | 35.525 | 35.52 | 0.005 | -0.057052 | 1
TR 1st set | 2008-01-30 | 0:13 | 113.7833 | 35.229 | 35.22 | 0.0089526 |  | 35.4637 | 35.07 | 0.39365 | -0.3847 | TR 1st set | 2008-01-30 | 0:13 | 113.7833 | 35.2289 | 35.18 | 0.048888 |  | 35.4802 | 35.09 | 0.39019 | -0.3413 | 1
TR 2nd set | 2008-02-18 | 0:13 | 113.7843 | 35.2315 | 35.18 | 0.051464 |  | 35.7857 | 35.16 | 0.6257 | -0.57424 | TR 2nd set | 2008-02-18 | 0:13 | 113.7849 | 35.2331 | 35.02 | 0.21306 |  | 35.7837 | 35.09 | 0.69365 | -0.48059 | 1
NO 1st set | 2008-03-07 | 0:12 | 113.7854 | 35.2345 | 35.16 | 0.07446 |  | 35.502 | 35.38 | 0.12201 | -0.047555 | NO 1st set | 2008-03-07 | 0:10 | 113.7853 | 35.2342 | 35.05 | 0.1842 |  | 35.5011 | 35.26 | 0.24106 | -0.056858 | 1
NO 2nd set | 2008-03-10 | 0:08 | 113.7854 | 35.2343 | 35.17 | 0.064318 |  | 35.5045 | 35.4 | 0.10455 | -0.040232 | NO 2nd set | 2008-03-10 | 0:07 | 113.7854 | 35.2344 | 35.02 | 0.21439 |  | 35.505 | 35.24 | 0.26503 | -0.050644 | 1
DK 1st set | 2008-05-21 | 0:21 | 113.786 | 35.2359 | 35.14 | 0.095919 |  | 35.4987 | 35.36 | 0.13865 | -0.042731 | DK 1st set | 2008-05-21 | 0:17 | 113.7859 | 35.2358 | 34.86 | 0.3758 |  | 35.4991 | 35.07 | 0.42905 | -0.053246 | 1
DK 2nd set | 2008-06-06 | 0:17 | 113.7864 | 35.237 | 35.1 | 0.13704 |  | 35.5007 | 35.33 | 0.17075 | -0.03371 | DK 2nd set | 2008-06-06 | 0:16 | 113.7864 | 35.2369 | 34.82 | 0.41687 |  | 35.5005 | 35.03 | 0.47045 | -0.053578 | 1
DE 1st set | 2008-06-30 | 0:07 | 113.7805 | 35.2219 | 35.1 | 0.12192 |  | 35.5024 | 35.3455 | 0.1569 | -0.034979 | DE 1st set | 2008-06-30 | 0:08 | 113.7803 | 35.2213 | 34.96 | 0.26128 |  | 35.5035 | 35.1 | 0.4035 | -0.14222 | 1
DE 2nd set | 2008-07-07 | 0:07 | 113.781 | 35.2231 | 35.11 | 0.11308 |  | 35.5062 | 35.35 | 0.15615 | -0.04307 | DE 2nd set | 2008-07-07 | 0:06 | 113.781 | 35.2231 | 34.97 | 0.25308 |  | 35.5071 | 35.11 | 0.3971 | -0.14402 | 1
UK 1st set | 2008-07-17 | 936:39 | 113.7868 | 35.2381 | 35.11 | 0.12811 |  | 35.2202 | 35.09 | 0.13019 | -0.0020813 | UK 1st set | 2008-07-17 | 0:31 | 113.7866 | 35.2376 | 34.86 | 0.37758 |  | 35.2193 | 34.8 | 0.41928 | -0.041709 | 1
UK 2nd set | 2008-07-24 | 0:29 | 113.7867 | 35.2378 | 35.11 | 0.12781 |  | 35.217 | 35.05 | 0.16696 | -0.039154 | UK 2nd set | 2008-07-24 | 0:25 | 113.7865 | 35.2372 | 34.85 | 0.3872 |  | 35.217 | 34.72 | 0.49696 | -0.10976 | 1
NL 1st set | 2008-09-17 | 1:00 | 113.7888 | 35.2432 | 35.09 | 0.15319 |  | 35.5486 | 35.32 | 0.2286 | -0.075414 | NL 1st set | 2008-09-17 | 0:40 | 113.7876 | 35.2401 | 34.83 | 0.41011 |  | 35.5139 | 35 | 0.5139 | -0.10379 | 1
NL 2nd set | 2008-09-22 | 0:35 | 113.7862 | 35.2364 | 35.09 | 0.1464 |  | 35.5089 | 35.3 | 0.20885 | -0.062446 | NL 2nd set | 2008-09-22 | 0:48 | 113.7862 | 35.2365 | 34.78 | 0.45646 |  | 35.5367 | 35.03 | 0.50675 | -0.050287 | 1
FR 1st set | 2008-10-06 | 0:20 | 113.7841 | 35.2312 | 35.06 | 0.17117 |  | 35.4888 | 35.28 | 0.20885 | -0.037682 | FR 1st set | 2008-10-06 | 0:20 | 113.7841 | 35.2312 | 34.98 | 0.25118 |  | 35.489 | 35.14 | 0.349 | -0.09782 | 2
FR 2nd set | 2008-10-20 | 0:25 | 113.7842 | 35.2313 | 34.81 | 0.42132 |  | 35.49 | 35.02 | 0.47 | -0.048678 | FR 2nd set | 2008-10-20 | 0:20 | 113.7842 | 35.2314 | 34.78 | 0.45136 |  | 35.4903 | 34.95 | 0.5403 | -0.088939 | 2
SI 2nd set | 2008-12-01 | 0:00 | 113.7839 | 35.2371 | 35.19 | 0.0471 |  | 35.2371 | 35.19 | 0.0471 | 0 | SI 2nd set | 2008-12-01 | 0:00 | 113.7839 | 35.2371 | 35.12 | 0.1171 |  | 35.2371 | 35.12 | 0.1171 | 0 | 2
Uncertainty budget in calibration of transfer IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0068633 | 0.00021467 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0099337 | 0.013974 | SI 1st set | 0.026178 | 0.00036686 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.027146 | 0.038364
CH 1st set | 0.018 | 0.00125 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.019419 | 0.028652 | CH 1st set | 0.025 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026029 | 0.038449
CH 2nd set | 0.007 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.010076 | 0.016534 | CH 2nd set | 0.029 | 0.001 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.029892 | 0.03782
TR 1st set | 0.0059 | 0.00081 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.0096064 | 0.18005 | TR 1st set | 0.0225 | 0.00098 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.023749 | 0.18288
TR 2nd set | 0.0075 | 0.00118 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.010699 | 0.18441 | TR 2nd set | 0.0192 | 0.00109 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.020655 | 0.18452
NO 1st set | 0.0082078 | 0.00037505 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01091 | 0.02642 | NO 1st set | 0.024301 | 0.0004881 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.025343 | 0.038487
NO 2nd set | 0.0055012 | 0.00039585 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0090524 | 0.024687 | NO 2nd set | 0.014244 | 0.00019805 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015952 | 0.031903
DK 1st set | 0.0067082 | 0.00033917 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0098307 | 0.023121 | DK 1st set | 0.013328 | 0.00031266 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015141 | 0.032732
DK 2nd set | 0.0067082 | 0.00046224 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0098357 | 0.022844 | DK 2nd set | 0.01214 | 0.00019192 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.014104 | 0.030987
DE 1st set | 0.0067082 | 0.0013146 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0099124 | 0.021993 | DE 1st set | 0.01619 | 0.0012109 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.017751 | 0.028086
DE 2nd set | 0.0082558 | 0 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01094 | 0.023592 | DE 2nd set | 0.013338 | 0 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015147 | 0.026593
UK 1st set | 0.009 | 0.0003 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.016632 | 0.030341 | UK 1st set | 0.043 | 0.0005 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.045219 | 0.06403
UK 2nd set | 0.006 | 0.0003 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.015219 | 0.029629 | UK 2nd set | 0.024 | 0.0004 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.027779 | 0.039031
NL 1st set | 0.033 | 0.014 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.07026 | 0.07032 | NL 1st set | 0.036 | 0.013 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.07153 | 0.071588
NL 2nd set | 0.029 | 0.023 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.070863 | 0.070922 | NL 2nd set | 0.033 | 0.016 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.070686 | 0.070745
FR 1st set | 0.014609 | 0.00024218 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.016279 | 0.029822 | FR 1st set | 0.025438 | 0.00022846 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026433 | 0.045711
FR 2nd set | 0.024252 | 0.00027714 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.025294 | 0.042956 | FR 2nd set | 0.013416 | 0.00047472 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.015223 | 0.032889
SI 2nd set | 0.0072548 | 0.00030779 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01021 | 0.014367 | SI 2nd set | 0.01916 | 0.00075394 | 0.0035 | 0.0019 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.020474 | 0.028919
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of transfer IRET without probe cover compared to Local BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0068633 | 0.00021467 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0098283 |  | SI 1st set | 0.026178 | 0.00036686 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.027108
CH 1st set | 0.019 | 0.006 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.021067 |  | CH 1st set | 0.027 | 0.005 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.028299
CH 2nd set | 0.01 | 0.005 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.013109 |  | CH 2nd set | 0.021 | 0.007 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.02317
TR 1st set | 0.0362 | 0.05741 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.17979 |  | TR 1st set | 0.04727 | 0.05412 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18133
TR 2nd set | 0.0565 | 0.05461 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.1841 |  | TR 2nd set | 0.04861 | 0.05949 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18336
NO 1st set | 0.0091191 | 0.0056395 | 0.0025 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.024062 |  | NO 1st set | 0.019324 | 0.0012258 | 0.0025 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.028964
NO 2nd set | 0.0076777 | 0.0021294 | 0.0025 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.022967 |  | NO 2nd set | 0.01638 | 0.00557 | 0.0025 | 0.021046 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.027629
DK 1st set | 0.010699 | 0.00067082 | 0.005 | 0.012 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.020927 |  | DK 1st set | 0.022775 | 0.00067082 | 0.005 | 0.012 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.02902
DK 2nd set | 0.010052 | 0.0010195 | 0.005 | 0.012 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.020618 |  | DK 2nd set | 0.020926 | 0.00060481 | 0.005 | 0.012 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.027591
DE 1st set | 0.0051042 | 0.0019304 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.019632 |  | DE 1st set | 0.010563 | 0.0025443 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021765
DE 2nd set | 0.0087509 | 0.0021588 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.020902 |  | DE 2nd set | 0.010712 | 0.0027125 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021857
UK 1st set | 0.006 | 0.0013 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.025376 |  | UK 1st set | 0.038 | 0.0022 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.045333
UK 2nd set | 0.006 | 0.002 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.025422 |  | UK 2nd set | 0.012 | 0.0012 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.027417
NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868
NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868
FR 1st set | 0.013803 | 0.00048936 | 0.012 | 0.009 | 0.002 | 0.0005 |  | 0.0028868 | 0.014 | 0.024987 |  | FR 1st set | 0.030933 | 0.00064888 | 0.012 | 0.009 | 0.002 | 0.0005 |  | 0.0028868 | 0.014 | 0.037294
FR 2nd set | 0.027773 | 0.00072548 | 0.012 | 0.009 | 0.002 | 0.0005 |  | 0.0028868 | 0.014 | 0.034719 |  | FR 2nd set | 0.020391 | 0.00073269 | 0.012 | 0.009 | 0.002 | 0.0005 |  | 0.0028868 | 0.014 | 0.029153
SI 2nd set | 0.0072548 | 0.00030779 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.010108 |  | SI 2nd set | 0.01916 | 0.00075394 | 0.0035 | 0.0019 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.020423
Note on J36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,03 added to the total uncertainty
Note on J45: Igor: 0,0003 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,03 added to the total uncertainty
Note on J66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J74: Igor: 0,01 added to the total uncertainty
Note on W74: Igor: 0,021 added to the total uncertainty
Note on J75: Igor: 0,01 added to the total uncertainty
Note on W75: Igor: 0,021 added to the total uncertainty

## 38,0 °C
Analysis of comparison measurements at 38,0 °C | [11 empty] | Analysis of comparison measurements at 38,0 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison |  |  |  |  | max drift |  |  | Transfer PRT (LMK 1204) coefficients |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  | 0.005 | °C |  | b0 | -243.7645 |  | b0 | -243.7648 |  |  |  |  | 0.005 | °C
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  | 0.0014434 | °C |  | b1 | 2.3284 |  | b1 | 2.3248 |  |  |  |  | 0.0014434 | °C
b2 | 0.0010858 |  |  | b2 | 0.0011031 |  |  |  | uncertainty |  |  | b2 | 0.0010858 |  | b2 | 0.0011031 |  |  |  |  | uncertainty
Transfer IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Transfer IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Transfer BB - Transfer IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Local BB - Transfer IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Transfer BB - Transfer IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Local BB - Transfer IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 114.7538 | 37.7296 | 37.72 | 0.0095848 |  | 37.77 | 37.72 | 0.05 | -0.040415 | SI 1st set | 2007-11-14 | 0:06 | 114.7538 | 37.7296 | 37.65 | 0.079559 |  | 37.73 | 37.65 | 0.08 | -0.000441 | 1
CH 1st set | 2007-12-10 | 0:07 | 114.7509 | 37.7221 | 37.72 | 0.0021432 |  | 38.014 | 37.99 | 0.024 | -0.021857 | CH 1st set | 2007-12-10 | 0:07 | 114.7509 | 37.7221 | 37.79 | -0.067921 |  | 38.016 | 38 | 0.016 | -0.083921 | 1
CH 2nd set | 2008-01-07 | 0:07 | 114.7511 | 37.7225 | 37.73 | -0.0075217 |  | 38.016 | 37.99 | 0.026 | -0.033522 | CH 2nd set | 2008-01-07 | 0:07 | 114.751 | 37.7224 | 37.81 | -0.087573 |  | 38.015 | 38.02 | -0.005 | -0.082573 | 1
TR 1st set | 2008-01-31 | 0:14 | 114.7552 | 37.7331 | 37.72 | 0.013059 |  | 38.1292 | 37.59 | 0.5392 | -0.52614 | TR 1st set | 2008-01-31 | 0:12 | 114.7552 | 37.7331 | 37.63 | 0.10312 |  | 38.1329 | 37.5 | 0.63285 | -0.52973 | 1
TR 2nd set | 2008-02-19 | 0:12 | 114.7495 | 37.7184 | 37.7 | 0.01838 |  | 38.3155 | 37.68 | 0.6355 | -0.61712 | TR 2nd set | 2008-02-19 | 0:11 | 114.7497 | 37.7189 | 37.5 | 0.21892 |  | 38.3197 | 37.54 | 0.77965 | -0.56073 | 1
NO 1st set | 2008-03-07 | 0:09 | 114.7542 | 37.7307 | 37.66 | 0.070726 |  | 38.0011 | 37.88 | 0.12108 | -0.050354 | NO 1st set | 2008-03-07 | 0:07 | 114.7543 | 37.7309 | 37.5 | 0.23091 |  | 38.0009 | 37.68 | 0.32089 | -0.089976 | 1
NO 2nd set | 2008-03-10 | 0:06 | 114.7543 | 37.7309 | 37.66 | 0.070915 |  | 38.0047 | 37.88 | 0.12473 | -0.05382 | NO 2nd set | 2008-03-10 | 0:07 | 114.7543 | 37.731 | 37.51 | 0.22098 |  | 38.0038 | 37.68 | 0.32385 | -0.10287 | 1
DK 1st set | 2008-05-21 | 0:15 | 114.7551 | 37.7329 | 37.63 | 0.10289 |  | 37.9995 | 37.85 | 0.14945 | -0.046558 | DK 1st set | 2008-05-21 | 0:17 | 114.755 | 37.7328 | 37.3 | 0.43275 |  | 37.9995 | 37.51 | 0.4895 | -0.05675 | 1
DK 2nd set | 2008-06-06 | 0:16 | 114.7556 | 37.7343 | 37.59 | 0.14427 |  | 38.0014 | 37.83 | 0.17135 | -0.027079 | DK 2nd set | 2008-06-06 | 0:17 | 114.7556 | 37.7342 | 37.25 | 0.48422 |  | 38.0006 | 37.45 | 0.55065 | -0.066431 | 1
DE 1st set | 2008-06-30 | 0:07 | 114.7491 | 37.7175 | 37.61 | 0.10745 |  | 38.0073 | 37.86 | 0.1473 | -0.039848 | DE 1st set | 2008-06-30 | 0:08 | 114.749 | 37.7173 | 37.38 | 0.33732 |  | 38.0079 | 37.48 | 0.5279 | -0.19058 | 1
DE 2nd set | 2008-07-07 | 0:07 | 114.7493 | 37.7181 | 37.62 | 0.098096 |  | 38.0056 | 37.86 | 0.14565 | -0.047554 | DE 2nd set | 2008-07-07 | 0:08 | 114.7493 | 37.7181 | 37.46 | 0.2581 |  | 38.004 | 37.56 | 0.444 | -0.1859 | 1
UK 1st set | 2008-07-18 | 936:29 | 114.7555 | 37.7339 | 37.6 | 0.13391 |  | 37.7117 | 37.55 | 0.16169 | -0.027785 | UK 1st set | 2008-07-18 | 0:22 | 114.7554 | 37.7338 | 37.32 | 0.41375 |  | 37.7122 | 37.14 | 0.57221 | -0.15846 | 1
UK 2nd set | 2008-07-24 | 0:24 | 114.7553 | 37.7334 | 37.6 | 0.13342 |  | 37.8135 | 37.64 | 0.17353 | -0.040118 | UK 2nd set | 2008-07-24 | 0:28 | 114.7553 | 37.7333 | 37.3 | 0.43331 |  | 37.8131 | 37.23 | 0.58315 | -0.14983 | 1
NL 1st set | 2008-09-17 | 0:30 | 114.7564 | 37.7364 | 37.56 | 0.17636 |  | 38.0007 | 37.74 | 0.26065 | -0.084289 | NL 1st set | 2008-09-17 | ? | 114.7564 | 37.7363 | 37.24 | 0.49635 |  | 38.0114 | 37.39 | 0.62145 | -0.1251 | 1
NL 2nd set | 2008-09-22 | 0:37 | 114.7546 | 37.7317 | 37.55 | 0.18167 |  | 38.0139 | 37.76 | 0.2539 | -0.072226 | NL 2nd set | 2008-09-22 | 0:29 | 114.7548 | 37.7321 | 37.25 | 0.48213 |  | 38.0388 | 37.42 | 0.6188 | -0.13667 | 1
FR 1st set | 2008-10-07 | 0:30 | 114.7535 | 37.7287 | 37.52 | 0.20869 |  | 38.0015 | 37.76 | 0.2415 | -0.03281 | FR 1st set | 2008-10-07 | 0:30 | 114.7533 | 37.7284 | 37.35 | 0.37837 |  | 38.0002 | 37.52 | 0.48015 | -0.10178 | 2
FR 2nd set | 2008-10-21 | 0:25 | 114.7533 | 37.7284 | 37.43 | 0.29837 |  | 37.9979 | 37.64 | 0.3579 | -0.059532 | FR 2nd set | 2008-10-21 | 0:20 | 114.7533 | 37.7283 | 37.31 | 0.4183 |  | 37.9991 | 37.49 | 0.50914 | -0.090836 | 2
SI 2nd set | 2008-12-01 | 0:00 | 114.7529 | 37.7335 | 37.64 | 0.09355 |  | 37.734 | 37.64 | 0.094 | -0.00045 | SI 2nd set | 2008-12-01 | 0:00 | 114.7529 | 37.7337 | 37.58 | 0.15365 |  | 37.734 | 37.58 | 0.154 | -0.00035 | 2
Uncertainty budget in calibration of transfer IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0065695 | 0.0011418 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0098827 | 0.013902 | SI 1st set | 0.014681 | 0.0010515 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.016427 | 0.023186
CH 1st set | 0.005 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0088998 | 0.016494 | CH 1st set | 0.041 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.041656 | 0.053917
CH 2nd set | 0.01 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.012418 | 0.017834 | CH 2nd set | 0.028 | 0.001 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.028952 | 0.039485
TR 1st set | 0.0061 | 0.00079 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.0098148 | 0.18873 | TR 1st set | 0.0248 | 0.00118 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.025979 | 0.18646
TR 2nd set | 0.0061 | 0.00126 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.0098638 | 0.19552 | TR 2nd set | 0.0164 | 0.00111 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.01813 | 0.19241
NO 1st set | 0.0095 | 0.00028 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011981 | 0.026928 | NO 1st set | 0.03011 | 0.00035 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.030983 | 0.043621
NO 2nd set | 0.0072 | 0.00019 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.010251 | 0.025163 | NO 2nd set | 0.01714 | 0.00028 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01863 | 0.033458
DK 1st set | 0.01099 | 0.00037772 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013196 | 0.026494 | DK 1st set | 0.020641 | 0.00038574 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.021895 | 0.036508
DK 2nd set | 0.008507 | 0.00054117 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011219 | 0.025163 | DK 2nd set | 0.01572 | 0.00046918 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.017336 | 0.042838
DE 1st set | 0.010699 | 0.00079337 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.012973 | 0.023666 | DE 1st set | 0.019494 | 0.00057637 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.020822 | 0.030195
DE 2nd set | 0.0060481 | 0.0012614 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0095591 | 0.022644 | DE 2nd set | 0.021981 | 0.0012614 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.023194 | 0.035067
UK 1st set | 0.007 | 0.0005 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.015699 | 0.030616 | UK 1st set | 0.018 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.022832 | 0.038639
UK 2nd set | 0.005 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.01491 | 0.029265 | UK 2nd set | 0.029 | 0.0003 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.032223 | 0.04156
NL 1st set | 0.02 | 0.013 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.064979 | 0.065043 | NL 1st set | 0.04 | 0.019 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.074928 | 0.074984
NL 2nd set | 0.03 | 0.028 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.073056 | 0.073113 | NL 2nd set | 0.025 | 0.008 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.065895 | 0.065959
FR 1st set | 0.020743 | 0.00028237 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.02199 | 0.036931 | FR 1st set | 0.021026 | 0.00024057 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.022257 | 0.038289
FR 2nd set | 0.026532 | 0.00041373 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.02752 | 0.040954 | FR 2nd set | 0.022942 | 0.00030613 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.024075 | 0.039388
SI 2nd set | 0.0068056 | 0.00060481 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.0099944 | 0.01406 | SI 2nd set | 0.01927 | 0.00048936 | 0.0035 | 0.0023 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.02061 | 0.029111
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of transfer IRET without probe cover compared to Local BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0065695 | 0.0011418 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0097768 |  | SI 1st set | 0.014681 | 0.0010515 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.016363
CH 1st set | 0.011 | 0.005 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.013886 |  | CH 1st set | 0.033 | 0.006 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.034232
CH 2nd set | 0.009 | 0.006 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.0128 |  | CH 2nd set | 0.025 | 0.007 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.026848
TR 1st set | 0.0554 | 0.06881 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18847 |  | TR 1st set | 0.04303 | 0.06724 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18464
TR 2nd set | 0.0672 | 0.0768 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.19528 |  | TR 2nd set | 0.06565 | 0.06831 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.19156
NO 1st set | 0.01 | 0.00175 | 0.0025 | 0.021388 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.024116 |  | NO 1st set | 0.021546 | 0.00035 | 0.0025 | 0.021388 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.030706
NO 2nd set | 0.0067 | 0.00216 | 0.0025 | 0.021388 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.02298 |  | NO 2nd set | 0.01714 | 0.00028 | 0.0025 | 0.021388 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.027792
DK 1st set | 0.012312 | 0.001099 | 0.005 | 0.014 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.022974 |  | DK 1st set | 0.021861 | 0.00068825 | 0.005 | 0.014 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.029213
DK 2nd set | 0.011459 | 0.00098809 | 0.005 | 0.014 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.022524 |  | DK 2nd set | 0.034043 | 0.00074516 | 0.005 | 0.014 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.039173
DE 1st set | 0.0057124 | 0.0018666 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.019793 |  | DE 1st set | 0.010712 | 0.0027891 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021867
DE 2nd set | 0.0076089 | 0.0027961 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.020527 |  | DE 2nd set | 0.018238 | 0.0018638 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.026301
UK 1st set | 0.009 | 0.0019 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.026284 |  | UK 1st set | 0.019 | 0.0021 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.031172
UK 2nd set | 0.005 | 0.0017 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.025182 |  | UK 2nd set | 0.009 | 0.0013 | 0 | 0.019 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.026248
NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868
NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868
FR 1st set | 0.021001 | 0.00082717 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.02967 |  | FR 1st set | 0.02305 | 0.00087509 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.031155
FR 2nd set | 0.021909 | 0.0011425 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.03033 |  | FR 2nd set | 0.023079 | 0.00075915 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.031174
SI 2nd set | 0.0068056 | 0.00060481 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.0098896 |  | SI 2nd set | 0.01927 | 0.00048936 | 0.0035 | 0.0023 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.020559
Note on J36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 and 0,002 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,005 added to the total uncertainty
Note on J45: Igor: 0,0003 and 0,002 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,005 added to the total uncertainty
Note on J66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J74: Igor: 0,003 added to the total uncertainty
Note on W74: Igor: 0,003 added to the total uncertainty
Note on J75: Igor: 0,003 added to the total uncertainty
Note on W75: Igor: 0,003 added to the total uncertainty

## 41,0 °C
Analysis of comparison measurements at 41,0 °C | [11 empty] | Analysis of comparison measurements at 41,0 °C
Transfer PRT (LMK 1204) coefficients |  |  |  | Transfer PRT (LMK 1204) coefficients after comparison |  |  |  |  | max drift |  |  | Transfer PRT (LMK 1204) coefficients |  |  | Transfer PRT (LMK 1204) coefficients after comparison | [5 empty] | max drift
b0 | -243.7645 |  |  | b0 | -243.7648 |  |  |  | 0.005 | °C |  | b0 | -243.7645 |  | b0 | -243.7648 |  |  |  |  | 0.005 | °C
b1 | 2.3284 |  |  | b1 | 2.3248 |  |  |  | 0.0014434 | °C |  | b1 | 2.3284 |  | b1 | 2.3248 |  |  |  |  | 0.0014434 | °C
b2 | 0.0010858 |  |  | b2 | 0.0005 |  |  |  | uncertainty |  |  | b2 | 0.0010858 |  | b2 | 0.0011031 |  |  |  |  | uncertainty
0.0005
Transfer IRET without probe cover compared to Transfer BB and Local BB | [11 empty] | Transfer IRET with probe cover compared to Transfer BB and Local BB
country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Transfer BB - Transfer IRET (without probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (without probe cover) (°C) | Local BB - Transfer IRET (without probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (without probe cover) | country/set | Date | Duration (minutes) | Transfer PRT resistance (ohm) | Transfer BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Transfer BB - Transfer IRET (with probe cover) (°C) |  | Local BB temperature (°C) | Transfer IRET (with probe cover) (°C) | Local BB - Transfer IRET (with probe cover) (°C) | Transfer BB-Local BB via Transfer IRET (with probe cover) | probe cover number
SI 1st set | 2007-11-14 | 0:06 | 115.9156 | 40.7258 | 40.74 | -0.014231 |  | 40.726 | 40.74 | -0.014 | -0.00023112 | SI 1st set | 2007-11-14 | 0:06 | 115.9155 | 40.7256 | 40.7 | 0.025588 |  | 40.726 | 40.7 | 0.026 | -0.00041173 | 1
CH 1st set | 2007-12-10 | 0:07 | 115.9129 | 40.7187 | 40.73 | -0.011314 |  | 41.014 | 40.99 | 0.024 | -0.035314 | CH 1st set | 2007-12-10 | 0:07 | 115.9129 | 40.7188 | 40.82 | -0.10118 |  | 41.01 | 41.02 | -0.01 | -0.091185 | 1
CH 2nd set | 2008-01-07 | 0:07 | 115.9129 | 40.7188 | 40.75 | -0.031249 |  | 41.009 | 41.01 | -0.001 | -0.030249 | CH 2nd set | 2008-01-07 | 0:06 | 115.9129 | 40.7187 | 40.84 | -0.12131 |  | 41.012 | 41.03 | -0.018 | -0.10331 | 1
TR 1st set | 2008-01-01 | 0:12 | 115.9154 | 40.7251 | 40.73 | -0.0048891 |  | 41.5637 | 40.83 | 0.73365 | -0.73854 | TR 1st set | 2008-01-01 | 0:11 | 115.9152 | 40.7246 | 40.61 | 0.11457 |  | 41.5694 | 40.79 | 0.77935 | -0.66478 | 1
TR 2nd set | 2008-02-20 | 0:11 | 115.9121 | 40.7168 | 40.67 | 0.046777 |  | 41.3376 | 40.48 | 0.8576 | -0.81082 | TR 2nd set | 2008-02-20 | 0:12 | 115.9115 | 40.7151 | 40.41 | 0.30513 |  | 41.3655 | 40.37 | 0.99555 | -0.69042 | 1
NO 1st set | 2008-03-07 | 0:07 | 115.9163 | 40.7275 | 40.68 | 0.047529 |  | 41.0123 | 40.91 | 0.10228 | -0.054751 | NO 1st set | 2008-03-07 | 0:07 | 115.9163 | 40.7276 | 40.5 | 0.2276 |  | 41.0126 | 40.68 | 0.33263 | -0.10503 | 1
NO 2nd set | 2008-03-10 | 0:07 | 115.9162 | 40.7273 | 40.68 | 0.047303 |  | 41.0154 | 40.91 | 0.1054 | -0.058092 | NO 2nd set | 2008-03-10 | 0:07 | 115.9162 | 40.7272 | 40.5 | 0.22723 |  | 41.0161 | 40.68 | 0.33615 | -0.10892 | 1
DK 1st set | 2008-05-21 | 0:19 | 115.9172 | 40.7299 | 40.61 | 0.11992 |  | 41.0003 | 40.83 | 0.17035 | -0.050427 | DK 1st set | 2008-05-21 | ? | 115.9172 | 40.7297 | 40.25 | 0.47973 |  | 41.0004 | 40.42 | 0.58045 | -0.10072 | 1
DK 2nd set | 2008-06-06 | 0:16 | 115.5304 | 39.7321 | 39.59 | 0.14211 |  | 41.0001 | 40.8 | 0.20015 | -0.058042 | DK 2nd set | 2008-06-06 | 0:20 | 115.5303 | 39.7318 | 39.21 | 0.5218 |  | 41.0004 | 40.37 | 0.6304 | -0.1086 | 1
DE 1st set | 2008-07-01 | 0:07 | 115.9111 | 40.714 | 40.61 | 0.10398 |  | 41.0054 | 40.84 | 0.1654 | -0.061422 | DE 1st set | 2008-07-01 | 0:07 | 115.911 | 40.7138 | 40.37 | 0.34385 |  | 41.0045 | 40.44 | 0.56445 | -0.2206 | 1
DE 2nd set | 2008-07-08 | 0:06 | 115.912 | 40.7163 | 40.6 | 0.1163 |  | 41.0049 | 40.82 | 0.1849 | -0.0686 | DE 2nd set | 2008-07-08 | 0:07 | 115.9117 | 40.7157 | 40.41 | 0.30565 |  | 41.005 | 40.47 | 0.53495 | -0.2293 | 1
UK 1st set | 2008-07-21 | 936:28 | 115.9173 | 40.7302 | 40.61 | 0.12022 |  | 40.8058 | 40.64 | 0.16579 | -0.045567 | UK 1st set | 2008-07-21 | 936:25 | 115.9173 | 40.73 | 40.26 | 0.47004 |  | 40.807 | 40.19 | 0.61695 | -0.14691 | 1
UK 2nd set | 2008-07-25 | 0:32 | 115.9176 | 40.7308 | 40.58 | 0.15082 |  | 40.8044 | 40.6 | 0.20436 | -0.053547 | UK 2nd set | 2008-07-25 | 0:26 | 115.9174 | 40.7304 | 40.23 | 0.50043 |  | 40.8035 | 40.1 | 0.70346 | -0.20303 | 1
NL 1st set | 2008-09-17 | 0:39 | 115.9205 | 40.7384 | 40.55 | 0.18836 |  | 41.038 | 40.76 | 0.27795 | -0.089586 | NL 1st set | 15.&17.9.2008 | 0:37 | 115.9204 | 40.7382 | 40.17 | 0.56817 |  | 41.0489 | 40.33 | 0.71895 | -0.15078 | 1
NL 2nd set | 2008-09-22 | 0:30 | 115.9202 | 40.7377 | 40.54 | 0.19771 |  | 41.0637 | 40.78 | 0.28365 | -0.085939 | NL 2nd set | 2008-09-22 | 0:27 | 115.9203 | 40.7378 | 40.14 | 0.5978 |  | 41.0682 | 40.36 | 0.7082 | -0.1104 | 1
FR 1st set | 2008-10-08 | 0:20 | 115.9153 | 40.7249 | 40.58 | 0.14489 |  | 41.0216 | 40.82 | 0.2016 | -0.056708 | FR 1st set | 2008-10-08 | 0:25 | 115.9152 | 40.7248 | 40.4 | 0.32479 |  | 41.0223 | 40.54 | 0.4823 | -0.15751 | 2
FR 2nd set | 2008-10-22 | 0:25 | 115.9153 | 40.7248 | 40.34 | 0.38483 |  | 41.0193 | 40.57 | 0.4493 | -0.064473 | FR 2nd set | 2008-10-22 | 0:20 | 115.9153 | 40.7249 | 40.48 | 0.24489 |  | 41.0209 | 40.63 | 0.3909 | -0.14601 | 2
SI 2nd set | 2008-12-01 | 0:00 | 115.9149 | 40.731 | 40.58 | 0.151 |  | 40.731 | 40.58 | 0.151 | 0 | SI 2nd set | 2008-12-01 | 0:00 | 115.9149 | 40.731 | 40.52 | 0.211 |  | 40.731 | 40.52 | 0.211 | 0 | 2
Uncertainty budget in calibration of transfer IRET without probe cover compared to Transfer BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Transfer BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s) |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) | utotal (1s)
SI 1st set | 0.0091047 | 0.00032851 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.01178 | 0.016596 | SI 1st set | 0.027028 | 0.00032313 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.028043 | 0.039632
CH 1st set | 0.009 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011737 | 0.017366 | CH 1st set | 0.034 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.034825 | 0.049645
CH 2nd set | 0.009 | 0.00125 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011761 | 0.01828 | CH 2nd set | 0.04 | 0.001 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.040703 | 0.053708
TR 1st set | 0.0126 | 0.00095 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.014856 | 0.18463 | TR 1st set | 0.0164 | 0.00128 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.018211 | 0.18093
TR 2nd set | 0.0089 | 0.00099 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.011884 | 0.18611 | TR 2nd set | 0.028 | 0.00086 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.0023 | 0.029082 | 0.18944
NO 1st set | 0.010052 | 0.00025374 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.012525 | 0.026775 | NO 1st set | 0.025021 | 0.00034256 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026114 | 0.037606
NO 2nd set | 0.012258 | 0.00045836 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.014361 | 0.027735 | NO 2nd set | 0.02473 | 0.00030112 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.025835 | 0.039065
DK 1st set | 0.011673 | 0.0007317 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013876 | 0.02885 | DK 1st set | 0.010699 | 0.00039509 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013053 | 0.03274
DK 2nd set | 0.0076472 | 0.0011416 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.010749 | 0.030624 | DK 2nd set | 0.018489 | 0.00044669 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.019945 | 0.047301
DE 1st set | 0.0092338 | 0.00057694 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011889 | 0.024985 | DE 1st set | 0.018064 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.019547 | 0.029291
DE 2nd set | 0.0088704 | 0.00057694 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.011609 | 0.024583 | DE 2nd set | 0.01713 | 0.0012131 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.018726 | 0.029951
UK 1st set | 0.015 | 0.0003 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.020612 | 0.033136 | UK 1st set | 0.031 | 0.0002 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.03407 | 0.04423
UK 2nd set | 0.006 | 0.0005 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.015363 | 0.031156 | UK 2nd set | 0.027 | 0.0002 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.012 | 0.030476 | 0.042802
NL 1st set | 0.025 | 0.021 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.068715 | 0.068776 | NL 1st set | 0.052 | 0.024 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.083281 | 0.083331
NL 2nd set | 0.027 | 0.023 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.070098 | 0.070158 | NL 2nd set | 0.065 | 0.02 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 | 0.06 | 0.090999 | 0.091044
FR 1st set | 0.011459 | 0.00035119 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013682 | 0.027905 | FR 1st set | 0.02593 | 0.0004539 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026988 | 0.045309
FR 2nd set | 0.16325 | 0.00021301 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.16342 | 0.23439 | FR 2nd set | 0.025021 | 0.00030871 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026113 | 0.044958
SI 2nd set | 0.01118 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.013445 | 0.018959 | SI 2nd set | 0.024942 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 | 0.0014434 | 0.0028868 |  | 0.026036 | 0.036792
[11 empty] | for BB ILC | [11 empty] | for BB ILC
Uncertainty budget in calibration of transfer IRET without probe cover compared to Local BB | [11 empty] | Uncertainty budget in calibration of transfer IRET with probe cover compared to Local BB
| uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s) |  |  | uIRET, std | uBB, temp. stab. | uBB, temp. hom. | uBB, emis. | uref. term. | uresistance | udrift | uIRET, resol. | uadd. | utotal (1s)
SI 1st set | 0.0091047 | 0.00032851 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.011691 |  | SI 1st set | 0.027028 | 0.00032313 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.028005
CH 1st set | 0.009 | 0.006 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.0128 |  | CH 1st set | 0.034 | 0.007 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.035381
CH 2nd set | 0.01 | 0.007 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.013994 |  | CH 2nd set | 0.034 | 0.005 | 0.0025 | 0.004 | 0.004 | 0.0005 |  | 0.0028868 |  | 0.03504
TR 1st set | 0.0566 | 0.05426 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18403 |  | TR 1st set | 0.04324 | 0.05306 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18001
TR 2nd set | 0.0528 | 0.06315 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.18573 |  | TR 2nd set | 0.05857 | 0.0624 | 0.151 | 0.07 | 0.002 |  |  | 0.0028868 | 0.0023 | 0.1872
NO 1st set | 0.007327 | 0.003143 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.023664 |  | NO 1st set | 0.015044 | 0.003088 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.027061
NO 2nd set | 0.0076777 | 0.0027598 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.023728 |  | NO 2nd set | 0.018806 | 0.0029235 | 0.0025 | 0.021803 | 0.0025 | 0.0005 |  | 0.0028868 |  | 0.029303
DK 1st set | 0.013089 | 0.00067082 | 0.005 | 0.017 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.025294 |  | DK 1st set | 0.0208 | 0.00094451 | 0.005 | 0.017 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.030026
DK 2nd set | 0.018806 | 0.00074516 | 0.005 | 0.017 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.028675 |  | DK 2nd set | 0.037031 | 0.00050262 | 0.005 | 0.017 | 0.012 | 0.0013 |  | 0.0028868 |  | 0.04289
DE 1st set | 0.010894 | 0.0029272 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021975 |  | DE 1st set | 0.010699 | 0.0023946 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021814
DE 2nd set | 0.010311 | 0.0027511 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.021669 |  | DE 2nd set | 0.013563 | 0.0026052 | 0.005 | 0.017 | 0.0056 | 0.0014 |  | 0.0028868 |  | 0.023375
UK 1st set | 0.005 | 0.0017 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.025945 |  | UK 1st set | 0.012 | 0.0025 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.028205
UK 2nd set | 0.009 | 0.0029 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.027105 |  | UK 2nd set | 0.016 | 0.0014 | 0 | 0.02 | 0.0077 | 0.0058 |  | 0.0028868 | 0.012 | 0.030054
NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 1st set | [7 empty] | 0.0028868 |  | 0.0028868
NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868 |  | NL 2nd set | [7 empty] | 0.0028868 |  | 0.0028868
FR 1st set | 0.012354 | 0.00050263 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.02432 |  | FR 1st set | 0.029753 | 0.0008645 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.036395
FR 2nd set | 0.1667 | 0.0012607 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.16802 |  | FR 2nd set | 0.029982 | 0.0013338 | 0.012 | 0.009 | 0.003 | 0.0005 |  | 0.0028868 | 0.014 | 0.036597
SI 2nd set | 0.01118 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.013367 |  | SI 2nd set | 0.024942 | 0 | 0.0035 | 0.0028 | 0.005 | 0.0005 |  | 0.0028868 |  | 0.025996
Note on J36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W36: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W37: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J44: Igor: 0,0003 and 0,01 added to the total uncertainty
Note on W44: Igor: 0,0003 and 0,01 added to the total uncertainty
Note on J45: Igor: 0,0003 and 0,01 added to the total uncertainty
Note on W45: Igor: 0,0003 and 0,01 added to the total uncertainty
Note on J66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W66: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on W67: ahmetd: uncertainty from the measuring instrument (multimeter)
Note on J74: Igor: 0,013 added to the total uncertainty
Note on W74: Igor: 0,026 added to the total uncertainty
Note on J75: Igor: 0,013 added to the total uncertainty
Note on W75: Igor: 0,026 added to the total uncertainty