## contents
code for sheet labels:
S-S x | Schiltknecht MINI (Small probe, 22 mm diameter); air speed x m/s; data and comparison plots
S-L x | Schiltknecht MACRO (Large probe, 85 mm diameter); air speed x m/s; data and comparison plots
T-S x | Testo 0635 9540 (Small probe, 16 mm diameter); air speed x m/s; data and comparison plots
T-L x | Testo 0635 9340 (Large probe, 107 mm diameter); air speed x m/s; data and comparison plots
v dependencies | plots of E(d) curves for each lab separately with all velocities in one plot
stability | results of repeated calibrations at CMI, evaluation of long term stability of the anemometers
DWG … | tests at DWG - vane and cup anemometers

## S-S 0.5
| 0.5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 2.4821
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 0.49207 | 0.505 | 0.0052705 | 2.6283 | 1.829 | 45 | -0.55161 | 1.8304
| 52.5 |  | 0.49133 | 0.473 | 0.0048305 | -3.7315 | 1.8317 | 60 | -1.7928 | 2.0354 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 0.49128 | 0.492 | 0.010328 | 0.14592 | 2.239 | 75 | 2.8396 | 2.1382 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 0.49084 | 0.518 | 0.0063246 | 5.5333 | 2.0373 | 90 | 5.294 | 1.9348 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 0.49117 | 0.516 | 0.005164 | 5.0546 | 1.8323 | 105 | 4.0145 | 1.832 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 0.49138 | 0.506 | 0.005164 | 2.9744 | 1.8316 | 120 | 4.065 | 1.8311 |  | 1.1474 | 1.6524 |  | 3.7828 |  | 0.77674 |  | 3.2358 |  |  | 2.3804 | [7 empty] | 1.0297 | 0.097459 | 0.56604 |  | 2.817
| 127.5 | 0.49176 | 0.49165 | 0.517 | 0.0048305 | 5.1557 | 1.8306 | 135 | 6.2837 | 1.8342 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 0.4897 | 0.526 | 0.005164 | 7.4118 | 1.8378 | 150 | 6.3966 | 1.9381 | 45 |  | 0.86289 |  | 2.5336 | [6 empty] | 1.799
| 157.5 |  | 0.4906 | 0.517 | 0.0067495 | 5.3814 | 2.0383 | 165 | 7.2491 | 1.937 | 60 | 1.4215 | 1.8817 |  | 2.1623 |  | 0.77791 |  | 1.3201 |  |  | 0.67958
| 172.5 |  | 0.4903 | 0.535 | 0.0052705 | 9.1168 | 1.8356 | 180 | 9.0517 | 1.8328 | 75 | 2.4425 | 1.3188 |  | 3.0803 |  | 0.89962 |  | 1.5244 |  |  | 2.0068
| 187.5 |  | 0.4918 | 0.536 | 0.005164 | 8.9865 | 1.83 | 195 | 8.4211 | 1.829 | 90 | 1.2178 | 3.6069 |  | 3.785 |  | 2.9337 |  | 3.2007 |  |  | 0.85208
| 202.5 |  | 0.49232 | 0.531 | 0.0031623 | 7.8558 | 1.8281 | 210 | 8.1465 | 1.8261 | 105 | 1.4393 | 1.1035 |  | 3.5405 |  | 0.12328 |  | 2.7629 |  |  | 2.5991
| 217.5 |  | 0.49337 | 0.535 | 0.0052705 | 8.4372 | 1.8242 | 225 |  |  | 120 | 0.7244 | 2.6224 |  | 3.3081 |  | 2.2618 |  | 3.0277 |  |  | 1.113 | [7 empty] | 2.3687 | 1.9754 | 0.2652 |  | 1.3402
[8 empty] | 240 |  |  | 135 | 1.3074 | 1.6687 |  | 3.7025 |  | 0.63664 |  | 3.0057 |  |  | 2.2401 | [7 empty] | 2.9672 | 2.0876 | 1.2319 |  | 1.2119
[8 empty] | 255 |  |  | 150 | 2.3872 | 2.7026 |  |  |  | 0.7639 | [12 empty] | 2.877 | 1.0105 | 0.23377
[8 empty] | 270 |  |  | 165 | 1.8748 | 2.2806 |  |  |  | 0.79318 | [12 empty] | 3.1081 | 1.7274 | 0.78985
[8 empty] | 285 |  |  | 180 | 2.6352 | 3.0529 |  |  |  | 0.93622 | [12 empty] | 1.9708 | 0.24808 | 1.0052
[8 empty] | 300 |  |  | 195 | 1.6864 | 2.1128 |  |  |  | 0.78703 | [12 empty] | 1.7787 | 0.42918 | 0.29572
[8 empty] | 315 |  |  | 210 | 0.93128 | 2.5784 |  |  |  | 2.0217 | [12 empty] | 1.6608 | 0.99783 | 0.84768
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.82721 | [13 empty] | 0.035609 | 0.66749
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.2588 | [13 empty] | 1.3819 | 0.73293
[8 empty] | 360 |  |  | 255 | [5 empty] | 2.7605 | [13 empty] | 0.53279 | 1.8798
[8 empty] | 375 |  |  | 270 | [20 empty] | 0.28049
[8 empty] | 390 |  |  | 285 | [20 empty] | 0.85558
[8 empty] | 405 |  |  | 300 | [20 empty] | 0.39555
[8 empty] | 420 |  |  | 315 | [20 empty] | 1.1144
[8 empty] | 435 |  |  | 330 | [20 empty] | 1.5167
[8 empty] | 450 |  |  | 345 | [20 empty] | 0.38981
[8 empty] | 465 |  |  | 360 | [20 empty] | 2.1029
[8 empty] | 480 |  |  | 375 | [20 empty] | 1.5604
[8 empty] | 495 |  |  | 390 | [20 empty] | 0.5208
[8 empty] | 510 |  |  | 405 | [20 empty] | 1.2909
[8 empty] | 525 |  |  | 420
435
450
45
PTB | 55 | 0.49542 | 0.49821 | 0.47308 | 0.0061769 | -5.0438 | 0.65 | 60 | -6.4493 | 0.65333
| 70 | 0.49542 | 0.49804 | 0.45192 | 0.0069393 | -9.2603 | 0.66 | 75 | -5.3199 | 0.65333
| 85 | 0.49542 | 0.49802 | 0.51077 | 0.0062757 | 2.5609 | 0.64 | 90 | 1.3771 | 0.66333
| 100 | 0.49542 | 0.49801 | 0.49308 | 0.0097033 | -0.99045 | 0.71 | 105 | -0.53746 | 0.69667
| 115 | 0.49542 | 0.49586 | 0.49769 | 0.0081524 | 0.36852 | 0.67 | 120 | 1.7806 | 0.65667
| 130 | 0.49542 | 0.49601 | 0.51885 | 0.0058835 | 4.6047 | 0.63 | 135 | 2.151 | 0.68333
| 145 | 0.49542 | 0.49618 | 0.4825 | 0.012247 | -2.7565 | 0.79 | 150 | -1.3223 | 0.73333
| 160 | 0.49542 | 0.49542 | 0.50308 | 0.0047068 | 1.546 | 0.62 | 165 | 1.2217 | 0.65
| 175 | 0.49542 | 0.49715 | 0.5 | 0.009798 | 0.57309 | 0.71 | 180 | 0.72769 | 0.67667
| 190 | 0.49542 | 0.49601 | 0.50115 | 0.0032581 | 1.0369 | 0.61 | 195 | 3.1196 | 0.61333
| 205 | 0.49542 | 0.49766 | 0.53391 | 0.0049901 | 7.2849 | 0.62 | 210 | 5.2186 | 0.62333
| 220 | 0.49542 | 0.497 | 0.5024 | 0.0052281 | 1.086 | 0.63 | 225 | 1.7163 | 0.63333
| 235 | 0.49542 | 0.49779 | 0.51261 | 0.0061919 | 2.9767 | 0.64 | 240 | 4.4363 | 0.63333
| 250 | 0.49542 | 0.4989 | 0.5356 | 0.0050662 | 7.3555 | 0.62 | 255 | 5.6827 | 0.62667
| 265 | 0.49542 | 0.49913 | 0.5108 | 0.0064031 | 2.3373 | 0.64 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 0.54463 | 0.524 | 0.0050262 | -3.7879 | 2.1343 | 45 | -3.7879 | 2.1343
| 60 |  | 0.54753 | 0.498 | 0.0083351 | -9.0461 | 2.1343 | 60 | -9.0461 | 2.1343
| 75 |  | 0.54513 | 0.5325 | 0.0044426 | -2.3169 | 2.1343 | 75 | -2.3169 | 2.1343
| 90 |  | 0.54598 | 0.5 | 0.0079472 | -8.4216 | 2.1343 | 90 | -8.4216 | 2.1343
| 105 |  | 0.54468 | 0.544 | 0.0050262 | -0.12484 | 2.1343 | 105 | -0.12484 | 2.1343
| 120 |  | 0.54548 | 0.514 | 0.0050262 | -5.7711 | 2.1343 | 120 | -5.7711 | 2.1343
| 135 |  | 0.54538 | 0.5455 | 0.0051042 | 0.022003 | 2.1343 | 135 | 0.022003 | 2.1343
| 150 |  | 0.54518 | 0.524 | 0.0050262 | -3.885 | 2.1343 | 150 | -3.885 | 2.1343
| 165 |  | 0.54578 | 0.538 | 0.0041039 | -1.4255 | 2.1343 | 165 | -1.4255 | 2.1343
| 180 |  | 0.54458 | 0.5315 | 0.0048936 | -2.4019 | 2.1343 | 180 | -2.4019 | 2.1343
| 195 |  | 0.54578 | 0.5485 | 0.0048936 | 0.49837 | 2.1343 | 195 | 0.49837 | 2.1343
| 210 |  | 0.54528 | 0.537 | 0.0047016 | -1.5185 | 2.1343 | 210 | -1.5185 | 2.1343
| 225 | 0.54518 | 0.54518 | 0.5395 | 0.0039403 | -1.0419 | 2.1343 | 225 | -1.0419 | 2.1343
| 240 |  | 0.54538 | 0.5285 | 0.0036635 | -3.0951 | 2.1343 | 240 | -3.0951 | 2.1343
| 255 |  | 0.54518 | 0.526 | 0.0050262 | -3.5181 | 2.1343 | 255 | -3.5181 | 2.1343
| 270 |  | 0.54448 | 0.545 | 0.0051299 | 0.095504 | 2.1343 | 270 | 0.095504 | 2.1343
| 285 |  | 0.54448 | 0.5365 | 0.0048936 | -1.4656 | 2.1343 | 285 | -1.4656 | 2.1343
| 300 |  | 0.54458 | 0.5355 | 0.0051042 | -1.6673 | 2.1343 | 300 | -1.6673 | 2.1343
| 315 |  | 0.54428 | 0.5205 | 0.0039403 | -4.3691 | 2.1343 | 315 | -4.3691 | 2.1343
| 330 |  | 0.54558 | 0.5235 | 0.0067082 | -4.0471 | 2.1343 | 330 | -4.0471 | 2.1343
| 345 |  | 0.54438 | 0.534 | 0.0050262 | -1.9068 | 2.1343 | 345 | -1.9068 | 2.1343
| 360 |  | 0.54578 | 0.509 | 0.0055251 | -6.739 | 2.1343 | 360 | -6.739 | 2.1343
| 375 |  | 0.54503 | 0.5195 | 0.0039403 | -4.6841 | 2.1343 | 375 | -4.6841 | 2.1343
| 390 |  | 0.54683 | 0.531 | 0.0055251 | -2.8949 | 2.1343 | 390 | -2.8949 | 2.1343
| 405 |  | 0.55353 | 0.532 | 0.0052315 | -3.8896 | 2.1343 | 405 | -3.8896 | 2.1343
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 0.5039 | 0.48 | 0.01 | -4.7429 | 2.0768 | 120 | -4.7429 | 2.0768
| 135 |  | 0.50396 | 0.48 | 0.01 | -4.7535 | 2.0763 | 135 | -4.7535 | 2.0763
| 150 |  | 0.50431 | 0.48 | 0.02 | -4.82 | 2.2984 | 150 | -4.82 | 2.2984
| 165 |  | 0.50258 | 0.48 | 0.01 | -4.492 | 2.0873 | 165 | -4.492 | 2.0873
| 180 |  | 0.50212 | 0.51 | 0.01 | 1.5694 | 2.212 | 180 | 1.5694 | 2.212
| 195 |  | 0.50164 | 0.51 | 0.01 | 1.6674 | 2.2162 | 195 | 1.6674 | 2.2162
| 210 |  | 0.50081 | 0.51 | 0.01 | 1.8359 | 2.2233 | 210 | 1.8359 | 2.2233
| 225 |  | 0.50199 | 0.51 | 0.01 | 1.5957 | 2.2131 | 225 | 1.5957 | 2.2131
| 240 |  | 0.50106 | 0.5 | 0.01 | -0.21229 | 2.1805 | 240 | -0.21229 | 2.1805
| 255 |  | 0.50054 | 0.52 | 0 | 3.8872 | 2.1916 | 255 | 3.8872 | 2.1916
| 270 | 0.501 | 0.5006 | 0.50667 | 0.013333 | 1.2127 | 2.2692 | 270 | 1.2127 | 2.2692
| 285 |  | 0.50038 | 0.51 | 0.01 | 1.9218 | 2.2269 | 285 | 1.9218 | 2.2269
| 300 |  | 0.50055 | 0.5 | 0.01 | -0.11062 | 2.1847 | 300 | -0.11062 | 2.1847
| 315 |  | 0.4999 | 0.5 | 0.01 | 0.020263 | 2.1903 | 315 | 0.020263 | 2.1903
| 330 |  | 0.5002 | 0.51 | 0.01 | 1.9591 | 2.2285 | 330 | 1.9591 | 2.2285
| 345 |  | 0.50164 | 0.5 | 0.02 | -0.32607 | 2.3931 | 345 | -0.32607 | 2.3931
| 360 |  | 0.50212 | 0.51 | 0.01 | 1.5694 | 2.212 | 360 | 1.5694 | 2.212
| 375 |  | 0.50258 | 0.51 | 0.01 | 1.4772 | 2.2082 | 375 | 1.4772 | 2.2082
| 390 |  | 0.50431 | 0.5 | 0.01 | -0.85418 | 2.1535 | 390 | -0.85418 | 2.1535
| 405 |  | 0.50396 | 0.51 | 0.01 | 1.1994 | 2.1965 | 405 | 1.1994 | 2.1965
| 420 |  | 0.5039 | 0.51 | 0.01 | 1.2106 | 2.197 | 420 | 1.2106 | 2.197
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
UL-LMPS | 42.5 | 0.55646 | 0.55499 | 0.48567 | 0.010726 | -12.4905 | 3.5275 | 45 | -12.4958 | 3.5656
| 57.5 | 0.55326 | 0.55024 | 0.48133 | 0.010417 | -12.5224 | 3.7565 | 60 | -12.4119 | 3.7168
| 72.5 | 0.55412 | 0.55744 | 0.49133 | 0.01008 | -11.8597 | 3.5181 | 75 | -11.9472 | 3.5088
| 87.5 | 0.55588 | 0.55926 | 0.49 | 0.010171 | -12.3843 | 3.4623 | 90 | -12.5003 | 3.4925
| 102.5 | 0.54722 | 0.55224 | 0.48 | 0.010171 | -13.0805 | 3.6431 | 105 | -12.7736 | 3.6011
| 117.5 | 0.55764 | 0.56857 | 0.50467 | 0.01383 | -11.2394 | 3.3908 | 120 | -10.9768 | 3.3409
| 132.5 | 0.56506 | 0.57046 | 0.51533 | 0.0062881 | -9.6636 | 3.0911 | 135 | -10.3118 | 3.2503
| 147.5 | 0.56112 | 0.56836 | 0.49133 | 0.012794 | -13.553 | 4.0466 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-S 2
| 2 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.54705
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 1.9648 | 1.858 | 0.0042164 | -5.4361 | 0.81433 | 45 | -4.7878 | 0.81465
| 52.5 |  | 1.9633 | 1.882 | 0.0042164 | -4.1395 | 0.81497 | 60 | -4.0177 | 0.83947 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 1.9677 | 1.891 | 0.0073786 | -3.8958 | 0.86397 | 75 | -3.848 | 0.86372 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 1.9688 | 1.894 | 0.0084327 | -3.8002 | 0.86346 | 90 | -3.1505 | 0.88876 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 1.9692 | 1.92 | 0.011547 | -2.5008 | 0.91406 | 105 | -2.0305 | 0.91371 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 1.9707 | 1.94 | 0.011547 | -1.5602 | 0.91336 | 120 | -1.4661 | 0.88815 |  | 0.2166 | 1.2609 | 0.26794 | 4.3593 | 0.84819 | 1.7649 | 0.37784 | 5.1277 | 1.0925 | 0.3712 | 5.8382 | 0.096512 | 2.4628 | 0.27957 | 4.1703 | 0.30479 | 0.45426 | 0.35339 | 1.595 | 2.0988 | 0.42345 | 0.58393 | 5.8771 | 0.41429 | 1.1213 | 1.2587 | 0.52098 | 0.6732 | 3.213 | 0.53508 | 0.0088857 | 0.029065 | 0.2653 | 0.11041 | 0.86424 | 0.23573
| 127.5 | 1.9707 | 1.97 | 1.943 | 0.0067495 | -1.3721 | 0.86293 | 135 | -1.3011 | 0.8373 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 1.9712 | 1.947 | 0.0048305 | -1.23 | 0.81167 | 150 | -1.2017 | 0.8375 | 45 |  | 0.10141 |  | 3.3587 | [6 empty] | 3.4916
| 157.5 |  | 1.9691 | 1.946 | 0.0069921 | -1.1733 | 0.86334 | 165 | -1.0593 | 0.86322 | 60 | 0.088909 | 0.32992 | 1.15 | 3.0894 |  | 0.30607 | 1.2367 | 3.7586 |  | 1.3213 | 3.4774 |  | 0.43968 | [11 empty] | 0.24977 | 0.29314 | 0.39931 | 0.62021 | 1.1779
| 172.5 |  | 1.9696 | 1.951 | 0.0073786 | -0.94526 | 0.86311 | 180 | -0.74622 | 0.86308 | 75 | 0.50654 | 0.6339 | 0.36889 | 3.0049 | 0.57405 | 0.24298 | 0.61512 | 4.2453 | 0.9954 | 0.69522 | 3.7805 | 1.0679 | 1.1381 | 0.017585 | 1.6816 |  | 0.12012 | [7 empty] | 0.58805 | 0.40098 | 0.2964 | 0.72962 | 1.9772 | 0.87806 | 0.12911 | 0.041049 | 0.0003199 | 0.26235 | 0.73027 | 0.28087
| 187.5 |  | 1.9698 | 1.959 | 0.0056765 | -0.54717 | 0.86304 | 195 | -0.39055 | 0.8379 | 90 | 0.41793 | 0.54791 | 0.066033 | 2.8074 | 0.33243 | 0.23184 | 0.26218 | 3.956 | 0.67287 | 0.34996 | 3.53 | 0.76409 | 1.3672 | 0.14718 | 1.8091 |  | 0.14483 | [7 empty] | 0.61516 | 0.46353 | 0.36064 | 0.53297 | 1.9432 | 0.7676 | 0.11366 | 0.039831 | 0.00096472 | 0.13276 | 0.68395 | 0.20169
| 202.5 |  | 1.9686 | 1.964 | 0.005164 | -0.23393 | 0.81276 | 210 | -0.2141 | 0.81251 | 105 | 0.031368 | 0.057875 | 0.17782 | 2.925 | 0.67477 | 0.10649 | 0.17232 | 3.6608 | 0.73478 | 0.21082 | 3.1731 | 0.7428 | 1.3327 | 0.26093 | 1.5692 |  | 0.15851 | [7 empty] | 0.41346 | 0.44528 | 0.39084 | 0.45642 | 1.8203 | 0.78486 | 0.010741 | 0.016528 | 0.0013145 | 0.078191 | 0.61243 | 0.19385
| 217.5 |  | 1.9698 | 1.966 | 0.005164 | -0.19428 | 0.81225 | 225 |  |  | 120 | 0.37239 | 0.30505 | 0.27136 | 2.7737 | 0.70888 | 0.019795 | 0.1125 | 3.026 | 0.52104 | 0.11685 | 2.586 | 0.49193 | 1.164 | 0.19374 | 1.4184 | 0.89546 | 0.13217 | 0.24739 | 1.2855 | 1.183 | 1.0365 | 0.41479 | 1.4337 | 0.3011 | 0.2592 | 0.4283 | 0.40785 | 0.40191 | 1.5936 | 0.65751 | 0.028797 | 0.037498 | 0.033993 | 0.076118 | 0.54484 | 0.16269
[8 empty] | 240 |  |  | 135 | 0.36911 | 0.10172 | 0.15446 | 2.8025 | 0.51837 | 0.50205 | 0.0034892 | 2.9189 | 0.31195 | 0.20641 | 2.9441 | 0.60062 | 1.3575 | 0.17811 | 1.6931 | 0.84858 | 0.18642 | 0.16956 | 0.94618 | 0.76294 | 1.0599 | 0.34732 | 1.8057 | 0.2272 | 0.38672 | 0.55256 | 0.34079 | 0.41415 | 1.7884 | 0.65871 | 0.03291 | 0.030508 | 0.053352 | 0.027394 | 0.58204 | 0.10529
[8 empty] | 255 |  |  | 150 | 0.30863 | 0.19642 | 0.14972 |  | 0.34288 | 0.080924 | 0.29299 |  | 0.16135 | 0.24908 |  | 0.19904 |  | 0.34815 |  | 0.69163 | 0.14109 | 0.14869 | 0.67814 | 0.50407 | 0.5007 | 0.50216 |  | 0.18459 | 0.36386 | 0.50407 | 0.45743 | 0.17476 |  | 0.53581 | 0.0056729 | 0.058916 | 0.045104 | 0.05145 |  | 0.096679
[8 empty] | 270 |  |  | 165 | 0.39405 | 0.24564 | 0.25123 |  | 0.097291 | 0.11295 | 0.4406 |  | 0.17227 | 0.37825 |  | 0.088034 |  | 0.29119 |  | 0.78433 | 0.13673 | 0.19371 | 0.67514 | 0.42518 | 0.45882 | 0.60945 |  | 0.43348 | 0.44981 | 0.63334 | 0.57006 | 0.16735 |  | 0.47442 | 0.047256 | 0.11675 | 0.097279 | 0.050992 |  | 0.072604
[8 empty] | 285 |  |  | 180 | 0.38474 | 0.081584 | 0.28852 |  | 0.15221 | 0.30332 | 0.47534 |  | 0.44176 | 0.33242 |  | 0.21719 |  | 0.17082 |  | 0.96273 | 0.22753 | 0.16987 | 0.63747 | 0.38618 | 0.57713 | 0.62376 |  | 0.64641 | 0.64781 | 0.83678 | 0.69175 | 0.29226 |  | 0.51061 | 0.033568 | 0.10132 | 0.050205 | 0.078123 |  | 0.0076542
[8 empty] | 300 |  |  | 195 | 0.35744 | 0.26125 | 0.25559 |  | 0.19278 | 0.68192 | 0.42506 |  | 0.46133 | 0.12569 |  | 0.0015026 |  | 0.11546 |  | 0.79933 | 0.228 | 0.09927 | 0.50687 | 0.25609 | 0.76771 | 0.51846 |  | 0.57903 | 0.55324 | 0.72282 | 0.43188 | 0.24198 |  | 0.39816 | 0.0081586 | 0.053502 | 0.060615 | 0.1047 |  | 0.059256
[8 empty] | 315 |  |  | 210 | 0.014367 | 0.43993 | 0.32891 |  | 0.28945 | 0.53546 | 0.34739 |  | 0.32508 | 0.11152 |  | 0.035927 |  | 0.12351 |  | 0.77218 | 0.32729 | 0.018662 | 0.3412 | 0.38186 | 0.76091 | 0.50123 |  | 0.54096 | 0.61196 | 0.6381 | 0.4066 | 0.23228 |  | 0.39477 | 0.09043 | 0.088501 | 0.17764 | 0.20988 |  | 0.16538
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.089376 | 0.23228 |  | 0.19076 | 0.18696 |  | 0.11983 |  | 0.097447 |  | 0.58639 | 0.33166 | 0.099716 |  | 0.31572 | 0.34631 | 0.36614 |  | 0.37991 |  | 0.47166 | 0.42184 | 0.18893 |  | 0.3182 |  | 0.16188 | 0.17595 | 0.23675 |  | 0.20463
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.039327 | 0.19545 |  | 0.030912 | 0.17227 |  | 0.05301 |  | 0.18927 |  | 0.38502 | 0.34901 | 0.2062 |  | 0.04753 | 0.072168 | 0.2064 |  | 0.0045019 |  | 0.38335 | 0.35591 | 0.14872 |  | 0.35807 |  | 0.21695 | 0.22232 | 0.27501 |  | 0.20486
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.020017 | 0.24908 |  | 0.10552 | 0.2321 |  | 0.10962 |  | 0.27825 |  | 0.27532 | 0.40603 | 0.3146 |  | 0.44629 | 0.37171 | 0.035814 |  | 0.3877 |  | 0.48857 | 0.46542 | 0.18888 |  | 0.49165 |  | 0.23519 | 0.23724 | 0.3117 |  | 0.20467
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.17344 |  | 0.35634 |  | 0.37486 |  | 0.16981 | 0.32158 | 0.26972 |  |  | 0.15945 | 0.088079 |  | 0.46421 |  |  | 0.2518 | 0.067609 |  | 0.43859 |  |  | 0.23693 | 0.28927 |  | 0.13955
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.15195 |  | 0.52474 |  | 0.45781 |  | 0.022783 | 0.319 | 0.34938 |  |  | 0.48758 | 0.10325 |  | 0.87511 |  |  | 0.21997 | 0.058862 |  | 0.50705 |  |  | 0.24761 | 0.29123 |  | 0.10581
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.4448 |  | 0.092447 |  | 0.46165 |  | 0.091816 | 0.3611 | 0.42384 |  |  | 1.3123 | 0.24414 |  | 1.0864 |  |  | 0.5617 | 0.10931 |  | 0.57029 |  |  | 0.14795 | 0.30763 |  | 0.12131
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.49835 |  | 0.047245 |  | 0.42539 |  | 0.18675 | 0.35231 | 0.45542 |  |  | 1.5707 | 0.32589 |  | 1.1486 |  |  | 0.59293 | 0.095149 |  | 0.51813 |  |  | 0.12502 | 0.30616 |  | 0.13544
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.4714 |  | 0.045626 |  | 0.40183 |  | 0.0039986 | 0.38752 | 0.4141 |  |  | 1.4584 | 0.30459 |  | 1.0725 |  |  | 0.73466 | 0.22398 |  | 0.64939 |  |  | 0.10272 | 0.27495 |  | 0.11307
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.49587 |  | 0.29294 |  | 0.63025 |  | 0.17083 | 0.34446 | 0.44125 |  |  | 1.3881 | 0.23136 |  | 1.3358 |  |  | 0.51062 | 0.036261 |  | 0.63898 |  |  | 0.15026 | 0.32907 |  | 0.070909
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.54917 | [5 empty] | 0.32789 | 0.30847 | 0.47115 |  |  | 1.496 | 0.23477 |  |  |  |  | 0.40269 | 0.08531 |  |  |  |  | 0.15765 | 0.35613
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.59264 | [5 empty] | 0.41678 |  |  |  |  | 1.9106 | 0.44673 |  |  |  |  | 0.55393 | 0.0010618
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.37615 | [5 empty] | 0.40308 |  |  |  |  | 1.4812 | 0.42632 |  |  |  |  | 0.34677 | 0.0031557
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.013805 | [5 empty] | 0.33411 |  |  |  |  | 0.61884 | 0.31038 |  |  |  |  | 0.032693 | 0.035391
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.35554 | [5 empty] | 0.53032 | [5 empty] | 0.11203
[11 empty] | 435 | [27 empty] | 0.093665
[11 empty] | 450 | [27 empty] | 0.14752
45
PTB | 55 | 2.0129 | 2.0195 | 1.9356 | 0.0076811 | -4.1544 | 0.44 | 60 | -3.9204 | 0.44
| 70 | 2.0129 | 2.0185 | 1.9488 | 0.0072572 | -3.4523 | 0.44 | 75 | -3.2842 | 0.44
| 85 | 2.0129 | 2.0174 | 1.9579 | 0.0065801 | -2.9479 | 0.44 | 90 | -2.6771 | 0.44
| 100 | 2.0129 | 2.0146 | 1.9716 | 0.006245 | -2.1355 | 0.44 | 105 | -2.0666 | 0.44
| 115 | 2.0129 | 2.0151 | 1.9762 | 0.0049454 | -1.9288 | 0.44 | 120 | -1.8877 | 0.44
| 130 | 2.0129 | 2.0139 | 1.9775 | 0.0067566 | -1.8056 | 0.44 | 135 | -1.7044 | 0.44
| 145 | 2.0129 | 2.0127 | 1.9825 | 0.0053161 | -1.5019 | 0.44 | 150 | -1.539 | 0.44
| 160 | 2.0129 | 2.0129 | 1.9804 | 0.007506 | -1.613 | 0.44 | 165 | -1.4977 | 0.44
| 175 | 2.0129 | 2.0139 | 1.9883 | 0.008165 | -1.2672 | 0.44 | 180 | -1.1743 | 0.44
| 190 | 2.0129 | 2.0128 | 1.9929 | 0.0069025 | -0.98841 | 0.44 | 195 | -0.78176 | 0.44333
| 205 | 2.0129 | 2.0134 | 2.006 | 0.010408 | -0.36844 | 0.45 | 210 | -0.22959 | 0.45
| 220 | 2.0129 | 2.0138 | 2.0148 | 0.011225 | 0.048106 | 0.45 | 225 | 0.16701 | 0.44667
| 235 | 2.0129 | 2.0138 | 2.022 | 0.006455 | 0.40482 | 0.44 | 240 | 0.41981 | 0.44
| 250 | 2.0129 | 2.0165 | 2.0256 | 0.0065064 | 0.4498 | 0.44 | 255 | 0.39013 | 0.44
| 265 | 2.0129 | 2.0193 | 2.0248 | 0.0077028 | 0.27077 | 0.44 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 1.9965 | 1.9035 | 0.0048936 | -4.6596 | 0.79758 | 45 | -4.6596 | 0.79758
| 60 |  | 2.0035 | 1.9315 | 0.0098809 | -3.5952 | 0.79758 | 60 | -3.5952 | 0.79758
| 75 |  | 2.0036 | 1.943 | 0.007327 | -3.026 | 0.79758 | 75 | -3.026 | 0.79758
| 90 |  | 2.0088 | 1.96 | 0.0056195 | -2.4308 | 0.79758 | 90 | -2.4308 | 0.79758
| 105 |  | 2.0082 | 1.969 | 0.0064072 | -1.9535 | 0.79758 | 105 | -1.9535 | 0.79758
| 120 |  | 2.0105 | 1.973 | 0.0057124 | -1.8667 | 0.79758 | 120 | -1.8667 | 0.79758
| 135 |  | 2.0095 | 1.986 | 0.0068056 | -1.1709 | 0.79758 | 135 | -1.1709 | 0.79758
| 150 |  | 2.0117 | 1.9825 | 0.0055012 | -1.453 | 0.79758 | 150 | -1.453 | 0.79758
| 165 |  | 2.0127 | 1.985 | 0.0082717 | -1.3777 | 0.79758 | 165 | -1.3777 | 0.79758
| 180 |  | 2.0106 | 1.9935 | 0.0087509 | -0.85197 | 0.79758 | 180 | -0.85197 | 0.79758
| 195 |  | 2.0091 | 2.008 | 0.010052 | -0.056243 | 0.79758 | 195 | -0.056243 | 0.79758
| 210 |  | 2.0111 | 2.018 | 0.0069585 | 0.3416 | 0.79758 | 210 | 0.3416 | 0.79758
| 225 | 2.0097 | 2.0097 | 2.015 | 0.0094591 | 0.26222 | 0.79758 | 225 | 0.26222 | 0.79758
| 240 |  | 2.0082 | 2.0175 | 0.0063867 | 0.4616 | 0.79758 | 240 | 0.4616 | 0.79758
| 255 |  | 2.0102 | 2.0185 | 0.0067082 | 0.4114 | 0.79758 | 255 | 0.4114 | 0.79758
| 270 |  | 2.0076 | 2.02 | 0.0064889 | 0.61615 | 0.79758 | 270 | 0.61615 | 0.79758
| 285 |  | 2.0034 | 2.0165 | 0.0104 | 0.65489 | 0.79758 | 285 | 0.65489 | 0.79758
| 300 |  | 2.0056 | 2.0005 | 0.0075915 | -0.25329 | 0.79758 | 300 | -0.25329 | 0.79758
| 315 |  | 2.0059 | 1.994 | 0.0068056 | -0.59226 | 0.79758 | 315 | -0.59226 | 0.79758
| 330 |  | 2.0064 | 1.989 | 0.0064072 | -0.86624 | 0.79758 | 330 | -0.86624 | 0.79758
| 345 |  | 2.0064 | 1.9875 | 0.0055012 | -0.941 | 0.79758 | 345 | -0.941 | 0.79758
| 360 |  | 2.0041 | 1.9815 | 0.0058714 | -1.1267 | 0.79758 | 360 | -1.1267 | 0.79758
| 375 |  | 2.0107 | 1.977 | 0.0057124 | -1.6775 | 0.79758 | 375 | -1.6775 | 0.79758
| 390 |  | 2.0019 | 1.978 | 0.0061559 | -1.1953 | 0.79758 | 390 | -1.1953 | 0.79758
| 405 |  | 1.9827 | 1.974 | 0.0068056 | -0.4403 | 0.79758 | 405 | -0.4403 | 0.79758
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 2.0048 | 1.94 | 0.01 | -3.2299 | 0.89133 | 120 | -3.2299 | 0.89133
| 135 |  | 2.0019 | 1.95 | 0.02 | -2.5948 | 0.93236 | 135 | -2.5948 | 0.93236
| 150 |  | 2.0024 | 1.96 | 0.01 | -2.1153 | 0.90241 | 150 | -2.1153 | 0.90241
| 165 |  | 1.9999 | 1.96 | 0.02 | -1.9961 | 0.93849 | 165 | -1.9961 | 0.93849
| 180 |  | 2.0024 | 1.97 | 0.01 | -1.6172 | 0.90687 | 180 | -1.6172 | 0.90687
| 195 |  | 2.0016 | 1.98 | 0.01 | -1.0768 | 0.91211 | 195 | -1.0768 | 0.91211
| 210 |  | 2.0035 | 1.99 | 0.01 | -0.67142 | 0.91488 | 210 | -0.67142 | 0.91488
| 225 |  | 2.0044 | 2 | 0.03 | -0.22074 | 1.0048 | 225 | -0.22074 | 1.0048
| 240 |  | 2.0027 | 2.01 | 0.02 | 0.36338 | 0.95763 | 240 | 0.36338 | 0.95763
| 255 |  | 2.0018 | 2.02 | 0.01 | 0.91011 | 0.92987 | 255 | 0.91011 | 0.92987
| 270 | 2.003 | 2.0034 | 2.02 | 0.01 | 0.82992 | 0.92841 | 270 | 0.82992 | 0.92841
| 285 |  | 2.0038 | 2.03 | 0.01 | 1.31 | 0.93254 | 285 | 1.31 | 0.93254
| 300 |  | 1.9997 | 2.03 | 0.01 | 1.5132 | 0.93624 | 300 | 1.5132 | 0.93624
| 315 |  | 1.9996 | 2.03 | 0.01 | 1.5223 | 0.93641 | 315 | 1.5223 | 0.93641
| 330 |  | 1.9975 | 2.02 | 0.02 | 1.1282 | 0.96681 | 330 | 1.1282 | 0.96681
| 345 |  | 2.0016 | 2.02 | 0.01 | 0.92161 | 0.93008 | 345 | 0.92161 | 0.93008
| 360 |  | 2.0024 | 2.02 | 0.01 | 0.87987 | 0.92932 | 360 | 0.87987 | 0.92932
| 375 |  | 1.9999 | 2.02 | 0.03 | 1.0041 | 1.0171 | 375 | 1.0041 | 1.0171
| 390 |  | 2.0024 | 2.02 | 0.03 | 0.88118 | 1.0149 | 390 | 0.88118 | 1.0149
| 405 |  | 2.0019 | 2.01 | 0.02 | 0.40228 | 0.95834 | 405 | 0.40228 | 0.95834
| 420 |  | 2.0048 | 2.02 | 0.03 | 0.76061 | 1.0127 | 420 | 0.76061 | 1.0127
435
450
465
480
495
510
525
45
DTI | 60 |  | 2.0001 | 1.9055 | 0.0060481 | -4.7314 | 2.6759 | 60 | -4.7314 | 2.6759
| 75 |  | 2.0016 | 1.9575 | 0.0055012 | -2.2024 | 2.6049 | 75 | -2.2024 | 2.6049
| 90 |  | 2.0007 | 1.972 | 0.0052315 | -1.4352 | 2.5857 | 90 | -1.4352 | 2.5857
| 105 |  | 2.0006 | 1.983 | 0.0047016 | -0.87974 | 2.5714 | 105 | -0.87974 | 2.5714
| 120 |  | 2.0015 | 1.9865 | 0.0048936 | -0.74795 | 2.5668 | 120 | -0.74795 | 2.5668
| 135 |  | 1.9993 | 1.9945 | 0.0051042 | -0.23942 | 2.5565 | 135 | -0.23942 | 2.5565
| 150 |  | 1.9996 | 1.9955 | 0.0051042 | -0.20321 | 2.5553 | 150 | -0.20321 | 2.5553
| 165 |  | 2 | 2.0035 | 0.0087509 | 0.17433 | 2.5451 | 165 | 0.17433 | 2.5451
| 180 |  | 2.0006 | 2.021 | 0.0085224 | 1.0172 | 2.523 | 180 | 1.0172 | 2.523
| 195 |  | 2.0013 | 2.0235 | 0.0048936 | 1.1094 | 2.5199 | 195 | 1.1094 | 2.5199
| 210 |  | 2.0003 | 2.029 | 0.0044721 | 1.4365 | 2.5131 | 210 | 1.4365 | 2.5131
| 225 |  | 2.0025 | 2.0305 | 0.0039403 | 1.3974 | 2.5112 | 225 | 1.3974 | 2.5112
| 240 |  | 2.0026 | 2.031 | 0.0044721 | 1.4192 | 2.5106 | 240 | 1.4192 | 2.5106
| 255 | 2.0011 | 2.0027 | 2.036 | 0.0050262 | 1.6609 | 2.5044 | 255 | 1.6609 | 2.5044
| 270 |  | 2.0026 | 2.0285 | 0.0036635 | 1.2943 | 2.5137 | 270 | 1.2943 | 2.5137
| 285 |  | 2.0025 | 2.0275 | 0.0055012 | 1.2476 | 2.5149 | 285 | 1.2476 | 2.5149
| 300 |  | 2.0003 | 2.0255 | 0.0060481 | 1.2615 | 2.5174 | 300 | 1.2615 | 2.5174
| 315 |  | 2.0013 | 2.0215 | 0.0048936 | 1.0095 | 2.5224 | 315 | 1.0095 | 2.5224
| 330 |  | 2.0006 | 2.023 | 0.0047016 | 1.1171 | 2.5205 | 330 | 1.1171 | 2.5205
| 345 |  | 2 | 2.009 | 0.010712 | 0.44933 | 2.5453 | 345 | 0.44933 | 2.5453
| 360 |  | 1.9996 | 1.999 | 0.0064072 | -0.028173 | 2.5508 | 360 | -0.028173 | 2.5508
| 375 |  | 1.9993 | 1.996 | 0.0082078 | -0.16439 | 2.5546 | 375 | -0.16439 | 2.5546
| 390 |  | 2.0015 | 1.9965 | 0.0048936 | -0.24832 | 2.554 | 390 | -0.24832 | 2.554
| 405 |  | 2.0006 | 1.99 | 0.0036418 | -0.52984 | 2.5623 | 405 | -0.52984 | 2.5623
| 420 |  | 2.0007 | 1.996 | 0.0082078 | -0.23558 | 2.5546 | 420 | -0.23558 | 2.5546
| 435 |  | 2.0016 | 1.9975 | 0.010699 | -0.20401 | 2.5598 | 435 | -0.20401 | 2.5598
| 450 |  | 2.0001 | 1.9965 | 0.0048936 | -0.18165 | 2.554 | 450 | -0.18165 | 2.554
465
480
495
510
525
45
DWG | 50 |  | 2.0993 | 1.9187 | 0.025768 | -8.6043 | 2.3905 | 60 | -6.9725 | 2.366
| 65 |  | 2.1283 | 1.9973 | 0.01615 | -6.1566 | 2.3537 | 75 | -4.794 | 2.3521
| 80 |  | 2.1283 | 2.0408 | 0.012388 | -4.1128 | 2.3512 | 90 | -3.3203 | 2.3506
| 95 |  | 2.1283 | 2.0661 | 0.0063235 | -2.924 | 2.3503 | 105 | -2.4894 | 2.3507
| 110 |  | 2.1273 | 2.079 | 0.0070096 | -2.272 | 2.351 | 120 | -2.1635 | 2.3487
| 125 |  | 2.1303 | 2.0854 | 0.0072041 | -2.1092 | 2.3476 | 135 | -1.6958 | 2.3518
| 140 |  | 2.1243 | 2.0927 | 0.0096558 | -1.4891 | 2.3539 | 150 | -0.81827 | 2.3575
| 155 |  | 2.1193 | 2.1091 | 0.0077584 | -0.48286 | 2.3593 | 165 | -0.41353 | 2.3585
| 170 |  | 2.1203 | 2.1123 | 0.0082706 | -0.37887 | 2.3581 | 180 | -0.0047162 | 2.3581
| 185 |  | 2.1203 | 2.1242 | 0.0061673 | 0.18236 | 2.3581 | 195 | 0.26401 | 2.3574
| 200 |  | 2.1213 | 2.1278 | 0.0046299 | 0.30484 | 2.357 | 210 | 0.62589 | 2.3585
| 215 |  | 2.1193 | 2.136 | 0.0056562 | 0.78641 | 2.3593 | 225 | 0.73855 | 2.3571
| 230 |  | 2.1223 | 2.1375 | 0.0048226 | 0.71462 | 2.356 | 240 | 0.90058 | 2.3575
| 245 |  | 2.1203 | 2.1414 | 0.0045036 | 0.99355 | 2.3582 | 255 | 1.003 | 2.3582
| 260 |  | 2.1203 | 2.1417 | 0.0042878 | 1.0077 | 2.3582 | 270 | 1.0584 | 2.359
| 275 |  | 2.1193 | 2.1423 | 0.0058281 | 1.0837 | 2.3594 | 285 | 1.0421 | 2.3579
| 290 |  | 2.1213 | 2.143 | 0.0045879 | 1.0214 | 2.3571 | 300 | 0.87982 | 2.3567
| 305 |  | 2.1218 | 2.139 | 0.0061166 | 0.80905 | 2.3565 | 315 | 0.67699 | 2.3561
| 320 |  | 2.1223 | 2.1353 | 0.0049979 | 0.61096 | 2.3559 | 330 | 0.3346 | 2.3567
| 335 |  | 2.1213 | 2.1255 | 0.0061017 | 0.19642 | 2.357 | 345 | 0.322 | 2.3563
| 350 |  | 2.1223 | 2.1305 | 0.0054907 | 0.3848 | 2.3559 | 360 | 0.27168 | 2.3555
| 365 |  | 2.1228 | 2.1274 | 0.0048024 | 0.21512 | 2.3553 | 375 | -0.16813 | 2.3561
| 380 |  | 2.1218 | 2.1142 | 0.0060475 | -0.35975 | 2.3565 | 390 | -0.23722 | 2.3565
| 395 |  | 2.1218 | 2.1181 | 0.0048506 | -0.17595 | 2.3565 | 405 | -0.40515 | 2.3557
| 410 |  | 2.1228 | 2.1118 | 0.0055411 | -0.51975 | 2.3554 | 420 | -0.62966 | 2.3554
| 425 |  | 2.1228 | 2.1083 | 0.0086679 | -0.68462 | 2.3554 | 435 | -0.53381 | 2.355
| 440 |  | 2.1233 | 2.1136 | 0.0052176 | -0.4584 | 2.3548 | 450 | -0.70054 | 2.356
| 455 |  | 2.1218 | 2.1044 | 0.0096093 | -0.82162 | 2.3565 | 465 | -0.50423 | 2.3561
| 470 |  | 2.1223 | 2.115 | 0.0050041 | -0.34553 | 2.3559 | 480 | -0.64395 | 2.356
| 485 |  | 2.1223 | 2.1055 | 0.0090573 | -0.79315 | 2.356 | 495 | -0.72014 | 2.3569
| 500 | 2.121 | 2.121 | 2.1065 | 0.0090602 | -0.68364 | 2.3574 | 510 | -0.82533 | 2.3554
| 515 |  | 2.1238 | 2.1048 | 0.008784 | -0.89618 | 2.3543 | 525 | -0.64828 | 2.3547
| 530 |  | 2.1233 | 2.1122 | 0.0050659 | -0.52433 | 2.3548
UL-LMPS | 42.5 | 2.0558 | 2.0463 | 1.8587 | 0.0077608 | -9.169 | 0.76071 | 45 | -9.0012 | 0.78151
| 57.5 | 2.0493 | 2.0304 | 1.8647 | 0.007303 | -8.1623 | 0.88551 | 60 | -8.1058 | 0.86438
| 72.5 | 2.042 | 2.0251 | 1.8667 | 0.0071116 | -7.8235 | 0.75869 | 75 | -7.6752 | 0.75959
| 87.5 | 2.0365 | 2.0183 | 1.8783 | 0.0087428 | -6.9337 | 0.76409 | 90 | -6.7812 | 0.76381
| 102.5 | 2.023 | 2.0139 | 1.8927 | 0.0082768 | -6.0184 | 0.76244 | 105 | -5.8627 | 0.76319
| 117.5 | 2.0165 | 2 | 1.8983 | 0.0094989 | -5.0846 | 0.76695 | 120 | -5.1233 | 0.80651
| 132.5 | 2.0097 | 2.0007 | 1.8943 | 0.0056832 | -5.3166 | 1.0043 | 135 | -5.2361 | 0.98545
| 147.5 | 2.0005 | 1.986 | 1.89 | 0.0069481 | -4.8335 | 0.89139 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | 64 |  | 2.0008 | 1.93 | 0.011698 | -3.5391 | 6.3 | 75 | -3.024 | 6.3
| 79 |  | 2.0074 | 1.9505 | 0.0068633 | -2.8367 | 6.3 | 90 | -2.4246 | 6.3
| 94 |  | 2.0051 | 1.9595 | 0.0068633 | -2.2748 | 6.3 | 105 | -1.9618 | 6.3
| 109 |  | 2.0015 | 1.9645 | 0.0075915 | -1.848 | 6.3 | 120 | -1.65 | 6.3
| 124 |  | 1.9985 | 1.967 | 0.0057124 | -1.578 | 6.3 | 135 | -1.511 | 6.3
| 139 |  | 1.9886 | 1.959 | 0.0064072 | -1.4866 | 6.3 | 150 | -1.1655 | 6.3
| 154 |  | 1.9757 | 1.955 | 0.0051299 | -1.0487 | 6.3 | 165 | -0.75769 | 6.3
| 169 |  | 1.9678 | 1.955 | 0.0051299 | -0.65186 | 6.3 | 180 | -0.53197 | 6.3
| 184 |  | 1.9651 | 1.9555 | 0.0051042 | -0.48838 | 6.3 | 195 | -0.44259 | 6.3
| 199 |  | 1.9634 | 1.955 | 0.0060698 | -0.42594 | 6.3 | 210 | -0.79066 | 6.3
| 214 |  | 1.9737 | 1.9555 | 0.0060481 | -0.92328 | 6.3 | 225 | -0.85922 | 6.3
| 229 |  | 1.9831 | 1.9665 | 0.0067082 | -0.83593 | 6.3 | 240 | -0.95541 | 6.3
| 244 |  | 1.9858 | 1.966 | 0.0050262 | -0.99886 | 6.3 | 255 | -1.1007 | 6.3
| 259 |  | 1.9841 | 1.9615 | 0.0067082 | -1.1378 | 6.3 | 270 | -0.89402 | 6.3
| 274 |  | 1.9875 | 1.9715 | 0.0067082 | -0.80538 | 6.3 | 285 | -0.92334 | 6.3
| 289 |  | 1.9872 | 1.968 | 0.0052315 | -0.96623 | 6.3 | 300 | -1.1963 | 6.3
| 304 |  | 1.9803 | 1.955 | 0.0051299 | -1.2799 | 6.3 | 315 | -1.3891 | 6.3
| 319 | 1.9752 | 1.9752 | 1.947 | 0.0065695 | -1.4288 | 6.3 | 330 | -1.5209 | 6.3
| 334 |  | 1.9676 | 1.937 | 0.0080131 | -1.5544 | 6.3 | 345 | -1.8987 | 6.3
| 349 |  | 1.9607 | 1.921 | 0.011653 | -2.0239 | 6.3 | 360 | -2.1315 | 6.3
| 364 |  | 1.9636 | 1.921 | 0.010712 | -2.1707 | 6.3 | 375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | 68 |  | 2.048 | 1.933 | 0.008 | -5.615 | 1.4 | 75 | -4.8431 | 1.4
| 83 |  | 2.045 | 1.964 | 0.008 | -3.961 | 1.4 | 90 | -3.7309 | 1.4
| 98 |  | 2.047 | 1.976 | 0.007 | -3.468 | 1.4 | 105 | -3.2174 | 1.4
| 113 |  | 2.047 | 1.987 | 0.007 | -2.931 | 1.4 | 120 | -2.7037 | 1.4
| 128 |  | 2.046 | 1.996 | 0.01 | -2.444 | 1.4 | 135 | -2.1929 | 1.4
| 143 |  | 2.046 | 2.007 | 0.011 | -1.906 | 1.4 | 150 | -1.7917 | 1.4
| 158 |  | 2.047 | 2.013 | 0.01 | -1.661 | 1.4 | 165 | -1.2279 | 1.4
| 173 |  | 2.046 | 2.031 | 0.007 | -0.733 | 1.4 | 180 | -0.4824 | 1.4
| 188 |  | 2.046 | 2.042 | 0.007 | -0.196 | 1.4 | 195 | -0.0588 | 1.4
| 203 |  | 2.046 | 2.048 | 0.007 | 0.098 | 1.4 | 210 | 0.28047 | 1.4
| 218 |  | 2.045 | 2.055 | 0.005 | 0.489 | 1.4 | 225 | 0.46613 | 1.4
| 233 |  | 2.046 | 2.055 | 0.007 | 0.44 | 1.4 | 240 | 0.3714 | 1.4
| 248 |  | 2.046 | 2.052 | 0.009 | 0.293 | 1.4 | 255 | 0.22487 | 1.4
| 263 |  | 2.045 | 2.048 | 0.007 | 0.147 | 1.4 | 270 | 0.0098 | 1.4
| 278 |  | 2.045 | 2.042 | 0.007 | -0.147 | 1.4 | 285 | -0.238 | 1.4
| 293 |  | 2.044 | 2.037 | 0.007 | -0.342 | 1.4 | 300 | -0.4106 | 1.4
| 308 |  | 2.045 | 2.035 | 0.005 | -0.489 | 1.4 | 315 | -0.51187 | 1.4
| 323 |  | 2.045 | 2.034 | 0.006 | -0.538 | 1.4 | 330 | -0.7886 | 1.4
| 338 |  | 2.047 | 2.025 | 0.006 | -1.075 | 1.4 | 345 | -1.4395 | 1.4
| 353 |  | 2.047 | 2.009 | 0.048 | -1.856 | 1.4 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-S 5
| 5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.22862
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 4.9778 | 4.707 | 0.0094868 | -5.44 | 0.68303 | 45 | -4.7401 | 0.68272
| 52.5 |  | 4.9823 | 4.781 | 0.0099443 | -4.0402 | 0.68242 | 60 | -3.6545 | 0.68252 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 4.9808 | 4.818 | 0.0091894 | -3.2688 | 0.68262 | 75 | -2.8627 | 0.69226 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 4.9865 | 4.864 | 0.011738 | -2.4566 | 0.7019 | 90 | -1.6276 | 0.69212 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 4.9828 | 4.943 | 0.0067495 | -0.79868 | 0.68235 | 105 | -0.80168 | 0.74188 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 4.9912 | 4.951 | 0.033149 | -0.80468 | 0.80142 | 120 | -0.47513 | 0.74178 |  | 0.29815 | 0.83608 | 1.454 | 4.2701 | 0.18444 | 1.4178 | 1.3926 | 4.753 | 0.47079 | 2.2187 | 5.6237 | 0.53228 | 2.0466 | 1.5107 | 4.0491 | 0.26953 |  |  | 0.70013 | 1.1327 | 0.045298 | 2.0228 | 4.9644 | 0.44703 | 0.82305 | 1.1313 | 0.25583 | 1.9478 | 4.2204 | 0.61844
| 127.5 | 4.9841 | 4.9843 | 4.977 | 0.0048305 | -0.14559 | 0.68215 | 135 | -0.10427 | 0.68181 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 4.9891 | 4.986 | 0.005164 | -0.06295 | 0.68148 | 150 | 0.23295 | 0.68226 | 45 |  | 1.1689 |  | 2.4981 | [6 empty] | 1.6708
| 157.5 |  | 4.9777 | 5.004 | 0.0069921 | 0.52885 | 0.68305 | 165 | 0.64478 | 0.68268 | 60 | 0.57423 | 1.2147 | 1.371 | 2.6087 |  | 0.84702 | 1.0911 | 2.5233 |  | 0.54182 | 1.7496 |  | 0.76941 | [11 empty] | 2.1948 | 2.0075 | 1.53 | 0.90578 | 0.34218
| 172.5 |  | 4.9831 | 5.021 | 0.0056765 | 0.76072 | 0.68231 | 180 | 0.82608 | 0.68166 | 75 | 0.5186 | 0.82784 | 0.41742 | 2.6899 | 2.0209 | 0.42461 | 0.06805 | 2.6828 | 1.8714 | 0.19487 | 2.2331 | 1.5074 | 1.7986 | 1.3146 | 0.44922 | [9 empty] | 0.72159 | 1.1568 | 1.3688 | 0.97938 | 2.7739 | 2.279
| 187.5 |  | 4.9925 | 5.037 | 0.0048305 | 0.89144 | 0.68102 | 195 | 0.88992 | 0.68115 | 90 | 0.79712 | 0.72208 | 0.30938 | 2.5724 | 1.9331 | 0.050278 | 0.26718 | 2.2628 | 1.5217 | 0.229 | 2.2041 | 1.5038 | 1.8082 | 1.3407 | 0.42995 | [9 empty] | 0.95978 | 1.6534 | 1.582 | 1.101 | 3.0056 | 2.4897
| 202.5 |  | 4.9907 | 5.035 | 0.0052705 | 0.88841 | 0.68127 | 210 | 0.57794 | 0.67978 | 105 | 0.75699 | 0.61206 | 0.28826 | 2.1748 | 1.6214 | 0.14884 | 0.27896 | 1.8928 | 1.2262 | 0.17969 | 1.9188 | 1.2832 | 1.554 | 1.1302 | 0.39389 | [9 empty] | 0.89035 | 1.6028 | 1.4701 | 1.0453 | 2.708 | 2.2318
| 217.5 |  | 5.0126 | 5.026 | 0.005164 | 0.26747 | 0.67829 | 225 |  |  | 120 | 0.84977 | 0.64877 | 0.080491 | 2.092 | 1.5282 | 0.21818 | 0.5911 | 1.6947 | 1.03 | 0.44033 | 1.7837 | 1.1426 | 1.6949 | 1.248 | 0.41289 | 1.7615 |  |  | 1.1248 | 0.49042 | 0.63985 | 0.86588 | 1.0415 | 0.53383 | 0.82176 | 1.5838 | 1.4135 | 0.80402 | 2.5479 | 2.0731
[8 empty] | 240 |  |  | 135 | 1.1173 | 0.51432 | 0.030873 | 2.1746 | 1.6039 | 0.69314 | 0.88127 | 1.5056 | 0.87069 | 0.43172 | 1.9454 | 1.3167 | 1.8286 | 1.3727 | 0.45812 | 1.4748 |  |  | 0.80133 | 0.15622 | 0.39425 | 0.68158 | 1.411 | 0.87391 | 0.84272 | 1.7545 | 1.2938 | 0.70579 | 2.5678 | 2.1029
[8 empty] | 255 |  |  | 150 | 1.238 | 0.92384 | 0.042293 |  | 1.575 | 0.31675 | 0.98507 |  | 0.73592 | 0.76146 |  | 0.92476 |  | 1.3594 |  | 1.6008 |  |  | 0.84302 | 0.24173 | 0.022625 | 0.72436 |  | 0.82575 | 0.92441 | 1.9485 | 1.7001 | 0.76486 |  | 2.1667
[8 empty] | 270 |  |  | 165 | 1.049 | 0.66844 | 0.073492 |  | 1.419 | 0.41693 | 0.71439 |  | 0.71484 | 0.43853 |  | 0.97753 |  | 1.1208 |  | 1.0506 |  |  | 0.57238 | 0.37372 | 0.023932 | 0.38985 |  | 0.9084 | 0.59441 | 1.4356 | 1.1465 | 0.57823 |  | 1.7247
[8 empty] | 285 |  |  | 180 | 0.70416 | 0.57059 | 0.035268 |  | 1.0926 | 0.12523 | 0.49464 |  | 0.64045 | 0.40362 |  | 0.70323 |  | 0.88235 |  | 0.77978 |  |  | 0.405 | 0.22461 | 0.11417 | 0.29258 |  | 0.73002 | 0.45318 | 1.0332 | 0.92751 | 0.42198 |  | 1.3329
[8 empty] | 300 |  |  | 195 | 0.32933 | 0.47323 | 0.035541 |  | 0.88845 | 0.20258 | 0.28909 |  | 0.73206 | 0.40455 |  | 0.56358 |  | 0.77706 |  | 0.64267 |  |  | 0.0056838 | 0.32074 | 0.46441 | 0.040059 |  | 0.88045 | 0.63975 | 0.95351 | 1.0483 | 0.52396 |  | 1.3308
[8 empty] | 315 |  |  | 210 | 0.34578 | 0.080789 | 0.32285 |  | 0.4749 | 0.51815 | 0.098045 |  | 0.84453 | 0.411 |  | 0.45231 |  | 0.69813 |  | 0.36125 |  |  | 0.70616 | 0.49643 | 0.87676 | 0.25114 |  | 1.1092 | 0.93557 | 0.78056 | 1.0715 | 0.52992 |  | 1.268
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.36584 | 0.069924 |  | 0.86481 | 0.29009 |  | 0.57681 |  | 0.68828 |  | 0.053475 |  |  |  | 0.65619 | 0.90864 | 0.38251 |  | 1.2476 |  | 0.54474 | 0.74594 | 0.37185 |  | 1.0781
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.14828 | 0.26177 |  | 0.84383 | 0.16349 |  | 0.91495 |  | 0.8357 |  | 0.34195 |  |  |  | 1.1661 | 0.99584 | 0.56967 |  | 1.6192 |  | 0.48472 | 0.38443 | 0.18383 |  | 1.0097
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.0090349 | 0.27536 |  | 1.1422 | 0.27285 |  | 1.0927 |  | 1.0651 |  | 0.074032 |  |  |  | 0.92659 | 0.88864 | 0.38645 |  | 1.6973 |  | 0.60503 | 0.59369 | 0.26433 |  | 1.3291
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.2403 |  | 1.2021 |  | 1.12 |  | 0.039368 |  |  |  |  | 0.72523 | 0.29634 |  | 1.6691 |  |  | 0.59097 | 0.2861 |  | 1.4111
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.43159 |  | 0.98538 |  | 1.1206 |  | 0.012044 |  |  |  |  | 1.0457 | 0.35637 |  | 1.7241 |  |  | 0.76162 | 0.28778 |  | 1.3836
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.59139 |  | 0.77249 |  | 1.0968 |  | 0.058881 |  |  |  |  | 1.3072 | 0.39671 |  | 1.7421 |  |  | 0.92166 | 0.28454 |  | 1.3731
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.56508 |  | 0.8579 |  | 1.139 |  | 0.053432 |  |  |  |  | 1.3361 | 0.44106 |  | 1.8423 |  |  | 0.94369 | 0.32551 |  | 1.454
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.49527 |  | 0.94244 |  | 1.1429 |  | 0.10992 |  |  |  |  | 1.1926 | 0.40152 |  | 1.8038 |  |  | 0.99234 | 0.43044 |  | 1.5518
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.28182 |  | 1.8015 |  | 1.611 |  | 0.18908 |  |  |  |  | 0.89572 | 0.39001 |  | 2.3313 |  |  | 0.87328 | 0.49449 |  | 2.0743
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.01672 | [5 empty] | 0.0030643 |  |  |  |  | 1.0614 | 0.77895 |  |  |  |  | 0.78442 | 0.65052
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.053693 | [5 empty] | 0.28068 |  |  |  |  | 1.4702 | 1.1079 |  |  |  |  | 0.79687 | 0.688
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.30509 | [5 empty] | 0.00034695 |  |  |  |  | 1.1505 | 1.1061 |  |  |  |  | 0.85416 | 0.92916
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.26152 | [5 empty] | 0.47521 |  |  |  |  | 1.6083 | 1.3944 |  |  |  |  | 0.66285 | 0.74127
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.27602 | [5 empty] | 1.3884 | [5 empty] | 0.92306
[11 empty] | 435 | [27 empty] | 1.0778
[11 empty] | 450 | [27 empty] | 1.1285
45
PTB | 55 | 4.9771 | 5.003 | 4.7832 | 0.01464 | -4.3939 | 0.42 | 60 | -4.1331 | 0.42
| 70 | 4.9771 | 4.9993 | 4.8188 | 0.0097125 | -3.6114 | 0.42 | 75 | -3.2999 | 0.42333
| 85 | 4.9771 | 4.9928 | 4.8592 | 0.028271 | -2.6768 | 0.43 | 90 | -2.3009 | 0.42667
| 100 | 4.9771 | 4.9889 | 4.9116 | 0.0089815 | -1.5489 | 0.42 | 105 | -1.4698 | 0.42
| 115 | 4.9771 | 4.9862 | 4.9208 | 0.0070238 | -1.3116 | 0.42 | 120 | -1.2251 | 0.42
| 130 | 4.9771 | 4.9804 | 4.928 | 0.006455 | -1.052 | 0.42 | 135 | -1.0329 | 0.41667
| 145 | 4.9771 | 4.9778 | 4.9283 | 0.0049103 | -0.99462 | 0.41 | 150 | -0.7944 | 0.41333
| 160 | 4.9771 | 4.9771 | 4.9575 | 0.012597 | -0.39397 | 0.42 | 165 | -0.22959 | 0.42
| 175 | 4.9771 | 4.9784 | 4.9833 | 0.0076139 | 0.099181 | 0.42 | 180 | 0.23975 | 0.42
| 190 | 4.9771 | 4.9809 | 5.0068 | 0.0069041 | 0.52089 | 0.42 | 195 | 0.61584 | 0.42
| 205 | 4.9771 | 4.9827 | 5.0228 | 0.0067823 | 0.80575 | 0.42 | 210 | 0.86475 | 0.41667
| 220 | 4.9771 | 4.9846 | 5.0336 | 0.0056862 | 0.98273 | 0.41 | 225 | 0.83855 | 0.41333
| 235 | 4.9771 | 4.9937 | 5.0212 | 0.0095192 | 0.5502 | 0.42 | 240 | 0.47537 | 0.42
| 250 | 4.9771 | 4.9964 | 5.0127 | 0.0087442 | 0.32571 | 0.42 | 255 | 0.20061 | 0.42
| 265 | 4.9771 | 5.0025 | 5 | 0.0084853 | -0.049603 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 5.0958 | 4.802 | 0.0076777 | -5.7648 | 0.5 | 45 | -5.7648 | 0.5
| 60 |  | 5.0934 | 4.853 | 0.013416 | -4.7191 | 0.5 | 60 | -4.7191 | 0.5
| 75 |  | 5.09 | 4.907 | 0.0092338 | -3.5945 | 0.5 | 75 | -3.5945 | 0.5
| 90 |  | 5.0868 | 4.9715 | 0.013089 | -2.2659 | 0.5 | 90 | -2.2659 | 0.5
| 105 |  | 5.0891 | 5.0195 | 0.0068633 | -1.3669 | 0.5 | 105 | -1.3669 | 0.5
| 120 |  | 5.0887 | 5.034 | 0.0068056 | -1.0742 | 0.5 | 120 | -1.0742 | 0.5
| 135 |  | 5.0852 | 5.057 | 0.011286 | -0.55475 | 0.5 | 135 | -0.55475 | 0.5
| 150 |  | 5.0838 | 5.0545 | 0.0088704 | -0.57654 | 0.5 | 150 | -0.57654 | 0.5
| 165 |  | 5.079 | 5.082 | 0.0069585 | 0.05887 | 0.5 | 165 | 0.05887 | 0.5
| 180 |  | 5.0829 | 5.0995 | 0.0068633 | 0.32639 | 0.5 | 180 | 0.32639 | 0.5
| 195 |  | 5.0853 | 5.1095 | 0.0068633 | 0.47568 | 0.5 | 195 | 0.47568 | 0.5
| 210 |  | 5.0837 | 5.1095 | 0.0060481 | 0.50731 | 0.5 | 210 | 0.50731 | 0.5
| 225 | 5.0842 | 5.0842 | 5.114 | 0.0075394 | 0.58692 | 0.5 | 225 | 0.58692 | 0.5
| 240 |  | 5.0851 | 5.1145 | 0.0082558 | 0.57796 | 0.5 | 240 | 0.57796 | 0.5
| 255 |  | 5.0886 | 5.0985 | 0.0081273 | 0.19436 | 0.5 | 255 | 0.19436 | 0.5
| 270 |  | 5.0894 | 5.085 | 0.010513 | -0.086651 | 0.5 | 270 | -0.086651 | 0.5
| 285 |  | 5.0909 | 5.0595 | 0.014681 | -0.61698 | 0.5 | 285 | -0.61698 | 0.5
| 300 |  | 5.0928 | 5.033 | 0.010809 | -1.1744 | 0.5 | 300 | -1.1744 | 0.5
| 315 |  | 5.0924 | 5.0225 | 0.010195 | -1.3728 | 0.5 | 315 | -1.3728 | 0.5
| 330 |  | 5.0897 | 5.0115 | 0.010894 | -1.5357 | 0.5 | 330 | -1.5357 | 0.5
| 345 |  | 5.0937 | 5.0165 | 0.0081273 | -1.5148 | 0.5 | 345 | -1.5148 | 0.5
| 360 |  | 5.0883 | 5.0215 | 0.015313 | -1.312 | 0.5 | 360 | -1.312 | 0.5
| 375 |  | 5.0929 | 5.0205 | 0.011459 | -1.4218 | 0.5 | 375 | -1.4218 | 0.5
| 390 |  | 5.0978 | 5.0365 | 0.014965 | -1.2027 | 0.5 | 390 | -1.2027 | 0.5
| 405 |  | 5.0791 | 5.0125 | 0.012927 | -1.3115 | 0.5 | 405 | -1.3115 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 5.0015 | 4.92 | 0.02 | -1.6291 | 0.67081 | 120 | -1.6291 | 0.67081
| 135 |  | 4.9951 | 4.95 | 0.03 | -0.90218 | 0.68872 | 135 | -0.90218 | 0.68872
| 150 |  | 4.9998 | 4.97 | 0.01 | -0.59615 | 0.67046 | 150 | -0.59615 | 0.67046
| 165 |  | 4.996 | 5 | 0.01 | 0.079713 | 0.67549 | 165 | 0.079713 | 0.67549
| 180 |  | 4.9987 | 5.02 | 0.01 | 0.426 | 0.67745 | 180 | 0.426 | 0.67745
| 195 |  | 5.0057 | 5.05 | 0.02 | 0.88428 | 0.68686 | 195 | 0.88428 | 0.68686
| 210 |  | 5.0059 | 5.07 | 0.02 | 1.2805 | 0.68946 | 210 | 1.2805 | 0.68946
| 225 |  | 5.0107 | 5.08 | 0.01 | 1.3831 | 0.68222 | 225 | 1.3831 | 0.68222
| 240 |  | 5.0066 | 5.08 | 0.03 | 1.4663 | 0.70253 | 240 | 1.4663 | 0.70253
| 255 |  | 5.0011 | 5.05 | 0.02 | 0.97705 | 0.68811 | 255 | 0.97705 | 0.68811
| 270 | 5.005 | 5.0026 | 5.03 | 0.013333 | 0.54732 | 0.67964 | 270 | 0.54732 | 0.67964
| 285 |  | 5.0045 | 5.02 | 0.03 | 0.30993 | 0.6953 | 285 | 0.30993 | 0.6953
| 300 |  | 5.0009 | 5 | 0.03 | -0.017347 | 0.69368 | 300 | -0.017347 | 0.69368
| 315 |  | 4.9996 | 4.99 | 0.03 | -0.19121 | 0.69273 | 315 | -0.19121 | 0.69273
| 330 |  | 4.9941 | 4.97 | 0.03 | -0.48198 | 0.69159 | 330 | -0.48198 | 0.69159
| 345 |  | 5.0057 | 4.97 | 0.04 | -0.71389 | 0.70518 | 345 | -0.71389 | 0.70518
| 360 |  | 4.9987 | 4.98 | 0.03 | -0.37421 | 0.69165 | 360 | -0.37421 | 0.69165
| 375 |  | 4.996 | 4.99 | 0.03 | -0.12045 | 0.69368 | 375 | -0.12045 | 0.69368
| 390 |  | 4.9998 | 4.99 | 0.02 | -0.19614 | 0.68052 | 390 | -0.19614 | 0.68052
| 405 |  | 4.9951 | 5 | 0.02 | 0.098807 | 0.68312 | 405 | 0.098807 | 0.68312
| 420 |  | 5.0015 | 5 | 0.03 | -0.029582 | 0.69351 | 420 | -0.029582 | 0.69351
435
450
465
480
495
510
525
45
DTI | 60 |  | 5.0737 | 4.73 | 0.02938 | -6.7749 | 1.2261 | 60 | -6.7749 | 1.2261
| 75 |  | 5.0617 | 4.965 | 0.022124 | -1.9097 | 1.1012 | 75 | -1.9097 | 1.1012
| 90 |  | 5.073 | 5.052 | 0.015079 | -0.41363 | 1.0337 | 90 | -0.41363 | 1.0337
| 105 |  | 5.0741 | 5.0915 | 0.014965 | 0.34318 | 1.0251 | 105 | 0.34318 | 1.0251
| 120 |  | 5.0639 | 5.094 | 0.01818 | 0.5942 | 1.0444 | 120 | 0.5942 | 1.0444
| 135 |  | 5.0663 | 5.1145 | 0.016051 | 0.95212 | 1.0268 | 135 | 0.95212 | 1.0268
| 150 |  | 5.0616 | 5.1315 | 0.01387 | 1.3802 | 1.0112 | 150 | 1.3802 | 1.0112
| 165 |  | 5.0672 | 5.1375 | 0.015517 | 1.3864 | 1.019 | 165 | 1.3864 | 1.019
| 180 |  | 5.0777 | 5.148 | 0.011965 | 1.3838 | 0.99867 | 180 | 1.3838 | 0.99867
| 195 |  | 5.0656 | 5.151 | 0.015183 | 1.6852 | 1.0145 | 195 | 1.6852 | 1.0145
| 210 |  | 5.0617 | 5.149 | 0.011192 | 1.7255 | 0.99509 | 210 | 1.7255 | 0.99509
| 225 |  | 5.0636 | 5.137 | 0.015252 | 1.4497 | 1.0176 | 225 | 1.4497 | 1.0176
| 240 |  | 5.0708 | 5.123 | 0.018382 | 1.0302 | 1.0399 | 240 | 1.0302 | 1.0399
| 255 | 5.0576 | 5.0657 | 5.1105 | 0.015381 | 0.88418 | 1.0236 | 255 | 0.88418 | 1.0236
| 270 |  | 5.0708 | 5.101 | 0.013727 | 0.59629 | 1.0165 | 270 | 0.59629 | 1.0165
| 285 |  | 5.0636 | 5.0785 | 0.020333 | 0.29439 | 1.0628 | 285 | 0.29439 | 1.0628
| 300 |  | 5.0617 | 5.057 | 0.015927 | -0.092065 | 1.0377 | 300 | -0.092065 | 1.0377
| 315 |  | 5.0656 | 5.0525 | 0.016819 | -0.25926 | 1.0441 | 315 | -0.25926 | 1.0441
| 330 |  | 5.0777 | 5.0605 | 0.021145 | -0.33946 | 1.0728 | 330 | -0.33946 | 1.0728
| 345 |  | 5.0672 | 5.0435 | 0.019541 | -0.46863 | 1.0644 | 345 | -0.46863 | 1.0644
| 360 |  | 5.0616 | 5.0425 | 0.018317 | -0.37814 | 1.056 | 360 | -0.37814 | 1.056
| 375 |  | 5.0663 | 5.042 | 0.017045 | -0.47892 | 1.0477 | 375 | -0.47892 | 1.0477
| 390 |  | 5.0639 | 5.054 | 0.01667 | -0.1957 | 1.0429 | 390 | -0.1957 | 1.0429
| 405 |  | 5.0741 | 5.048 | 0.020417 | -0.51412 | 1.0699 | 405 | -0.51412 | 1.0699
| 420 |  | 5.073 | 5.0535 | 0.018715 | -0.38406 | 1.0565 | 420 | -0.38406 | 1.0565
| 435 |  | 5.0617 | 5.055 | 0.018778 | -0.13164 | 1.0566 | 435 | -0.13164 | 1.0566
| 450 |  | 5.0737 | 5.0715 | 0.023458 | -0.044215 | 1.089 | 450 | -0.044215 | 1.089
465
480
495
510
525
45
DWG | 50 |  | 5.0157 | 4.6676 | 0.041658 | -6.9396 | 0.99927 | 60 | -5.3326 | 0.98998
| 65 |  | 5.0797 | 4.8496 | 0.015809 | -4.5292 | 0.98533 | 75 | -3.3745 | 0.98595
| 80 |  | 5.0717 | 4.9298 | 0.014076 | -2.7972 | 0.98625 | 90 | -2.0072 | 0.98686
| 95 |  | 5.0657 | 4.984 | 0.010607 | -1.6122 | 0.98717 | 105 | -1.1638 | 0.9875
| 110 |  | 5.0627 | 5.0151 | 0.008328 | -0.93956 | 0.98767 | 120 | -0.57627 | 0.98803
| 125 |  | 5.0597 | 5.0397 | 0.010641 | -0.39462 | 0.98822 | 135 | -0.066539 | 0.98821
| 140 |  | 5.0597 | 5.0646 | 0.007499 | 0.097503 | 0.98821 | 150 | 0.28465 | 0.98834
| 155 |  | 5.0587 | 5.0778 | 0.0073027 | 0.37823 | 0.98841 | 165 | 0.55491 | 0.98855
| 170 |  | 5.0577 | 5.0902 | 0.0061458 | 0.64325 | 0.98862 | 180 | 0.78297 | 0.98863
| 185 |  | 5.0577 | 5.1008 | 0.0043675 | 0.85283 | 0.98863 | 195 | 0.93337 | 0.98877
| 200 |  | 5.0567 | 5.1059 | 0.0080613 | 0.97363 | 0.98884 | 210 | 0.97231 | 0.98884
| 215 |  | 5.0567 | 5.1058 | 0.0060316 | 0.97165 | 0.98884 | 225 | 0.9152 | 0.9891
| 230 |  | 5.0547 | 5.0995 | 0.0075219 | 0.88697 | 0.98922 | 240 | 0.76299 | 0.98921
| 245 |  | 5.0547 | 5.0901 | 0.0070819 | 0.701 | 0.98921 | 255 | 0.50311 | 0.98907
| 260 |  | 5.0557 | 5.0761 | 0.0084807 | 0.40417 | 0.989 | 270 | 0.18528 | 0.98912
| 275 |  | 5.0547 | 5.0585 | 0.010911 | 0.075838 | 0.98919 | 285 | -0.12858 | 0.98912
| 290 |  | 5.0552 | 5.0435 | 0.012636 | -0.23079 | 0.98909 | 300 | -0.50514 | 0.98917
| 305 |  | 5.0547 | 5.0222 | 0.011354 | -0.64231 | 0.98921 | 315 | -0.73332 | 0.98921
| 320 |  | 5.0547 | 5.0153 | 0.0085986 | -0.77882 | 0.98921 | 330 | -0.97519 | 0.9891
| 335 |  | 5.0557 | 5.0014 | 0.0075992 | -1.0734 | 0.98905 | 345 | -1.1959 | 0.989
| 350 |  | 5.0562 | 4.9926 | 0.0062852 | -1.2572 | 0.98897 | 360 | -1.331 | 0.98891
| 365 |  | 5.0567 | 4.9875 | 0.007411 | -1.3678 | 0.98889 | 375 | -1.4825 | 0.9889
| 380 |  | 5.0567 | 4.9788 | 0.0090819 | -1.5399 | 0.98891 | 390 | -1.5479 | 0.98898
| 395 |  | 5.0562 | 4.9777 | 0.0093163 | -1.5519 | 0.98901 | 405 | -1.6074 | 0.98908
| 410 |  | 5.0557 | 4.973 | 0.010658 | -1.6351 | 0.98912 | 420 | -1.7364 | 0.98901
| 425 |  | 5.0567 | 4.9663 | 0.01079 | -1.7871 | 0.98895 | 435 | -1.7109 | 0.98913
| 440 |  | 5.0552 | 4.9706 | 0.0092335 | -1.6729 | 0.98923 | 450 | -1.7244 | 0.9893
| 455 |  | 5.0547 | 4.9662 | 0.010533 | -1.7502 | 0.98934 | 465 | -1.8054 | 0.98922
| 470 |  | 5.0557 | 4.963 | 0.012804 | -1.8329 | 0.98916 | 480 | -1.8424 | 0.98929
| 485 |  | 5.0547 | 4.9613 | 0.011233 | -1.8471 | 0.98935 | 495 | -1.8986 | 0.98936
| 500 | 5.055 | 5.0547 | 4.9574 | 0.013253 | -1.9243 | 0.98937 | 510 | -1.8769 | 0.98942
| 515 |  | 5.0542 | 4.9605 | 0.011398 | -1.8533 | 0.98945 | 525 | -1.819 | 0.98945
| 530 |  | 5.0542 | 4.9631 | 0.012403 | -1.8018 | 0.98944
UL-LMPS | 42.5 | 5.1095 | 5.1215 | 4.7405 | 0.012344 | -7.439 | 0.70797 | 45 | -7.2619 | 0.70757
| 57.5 | 5.0768 | 5.0677 | 4.7445 | 0.0094451 | -6.3768 | 0.70561 | 60 | -6.2846 | 0.70593
| 72.5 | 5.0467 | 5.0373 | 4.744 | 0.011877 | -5.8233 | 0.70754 | 75 | -5.5959 | 0.70781
| 87.5 | 5.0111 | 5.0062 | 4.783 | 0.013534 | -4.4591 | 0.70911 | 90 | -4.2423 | 0.70841
| 102.5 | 4.9825 | 4.9844 | 4.827 | 0.0083666 | -3.1585 | 0.70489 | 105 | -3.0828 | 0.7053
| 117.5 | 4.9521 | 4.9601 | 4.826 | 0.011626 | -2.7039 | 0.70733 | 120 | -2.6717 | 0.70707
| 132.5 | 4.9271 | 4.9363 | 4.8123 | 0.0097143 | -2.5105 | 0.7058 | 135 | -2.3947 | 0.76955
| 147.5 | 4.8857 | 4.8932 | 4.8043 | 0.013047 | -1.8154 | 1.0883 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | 68 |  | 5.238 | 4.946 | 0.008 | -5.575 | 0.83 | 75 | -5.0953 | 0.83
| 83 |  | 5.234 | 4.996 | 0.011 | -4.547 | 0.83 | 90 | -3.763 | 0.83
| 98 |  | 5.232 | 5.082 | 0.008 | -2.867 | 0.83 | 105 | -2.6444 | 0.83
| 113 |  | 5.23 | 5.105 | 0.006 | -2.39 | 0.83 | 120 | -2.2117 | 0.83
| 128 |  | 5.23 | 5.125 | 0.005 | -2.008 | 0.83 | 135 | -1.8657 | 0.83
| 143 |  | 5.226 | 5.137 | 0.008 | -1.703 | 0.83 | 150 | -1.4972 | 0.83
| 158 |  | 5.229 | 5.163 | 0.006 | -1.262 | 0.83 | 165 | -0.91433 | 0.83
| 173 |  | 5.225 | 5.198 | 0.008 | -0.517 | 0.83 | 180 | -0.37373 | 0.83
| 188 |  | 5.227 | 5.216 | 0.006 | -0.21 | 0.83 | 195 | -0.0854 | 0.83
| 203 |  | 5.224 | 5.227 | 0.005 | 0.057 | 0.83 | 210 | 0.057 | 0.83
| 218 |  | 5.223 | 5.226 | 0.007 | 0.057 | 0.83 | 225 | 0.012667 | 0.83
| 233 |  | 5.221 | 5.219 | 0.009 | -0.038 | 0.83 | 240 | -0.33293 | 0.83
| 248 |  | 5.22 | 5.185 | 0.014 | -0.67 | 0.83 | 255 | -0.89353 | 0.83
| 263 |  | 5.22 | 5.16 | 0.012 | -1.149 | 0.83 | 270 | -1.2834 | 0.83
| 278 |  | 5.218 | 5.143 | 0.007 | -1.437 | 0.83 | 285 | -1.598 | 0.83
| 293 |  | 5.218 | 5.125 | 0.007 | -1.782 | 0.83 | 300 | -1.9435 | 0.83
| 308 |  | 5.217 | 5.106 | 0.009 | -2.128 | 0.83 | 315 | -2.2269 | 0.83
| 323 |  | 5.214 | 5.092 | 0.011 | -2.34 | 0.83 | 330 | -2.4739 | 0.83
| 338 |  | 5.215 | 5.078 | 0.012 | -2.627 | 0.83 | 345 | -3.3083 | 0.83
| 353 |  | 5.212 | 4.999 | 0.017 | -4.087 | 0.83 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-S 8
| 8 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.35109
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 7.983 | 7.608 | 0.0063246 | -4.697 | 0.65139 | 45 | -4.0251 | 0.65123
| 52.5 |  | 7.9868 | 7.719 | 0.0056765 | -3.3531 | 0.65107 | 60 | -2.9514 | 0.65747 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 7.9836 | 7.78 | 0.011547 | -2.5497 | 0.66386 | 75 | -1.9512 | 0.66364 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 7.9891 | 7.881 | 0.016633 | -1.3527 | 0.66341 | 90 | -0.65721 | 0.66354 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 7.9859 | 7.989 | 0.0073786 | 0.038262 | 0.66367 | 105 | 0.12864 | 0.66331 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 7.9945 | 8.012 | 0.010328 | 0.21902 | 0.66296 | 120 | 0.43638 | 0.66344 |  | 0.25167 | 0.61096 | 3.2023 | 3.7534 | 0.15987 | 0.4463 | 3.4859 | 4.0684 | 0.070884 | 2.9496 | 3.5541 | 0.44472 | 0.73946 | 3.0916 | 3.6543 | 0.48558 | 1.2185 | 1.0436 | 1.1117 | 1.0388 | 0.61204 | 2.1445 | 2.7632 | 0.96612 | 0.60297 | 0.4474 | 0.062864 | 2.5596 | 3.1418 | 0.45364 | 1.4485 | 1.392 | 1.2583 | 0.27653 | 0.0027085 | 1.3923
| 127.5 | 7.9822 | 7.9828 | 8.035 | 0.0052705 | 0.65375 | 0.66393 | 135 | 0.71489 | 0.6635 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 7.993 | 8.055 | 0.0052705 | 0.77604 | 0.66308 | 150 | 1.0226 | 0.6638 | 45 |  | 1.5109 |  | 2.3953 | [6 empty] | 1.1757
| 157.5 |  | 7.9758 | 8.077 | 0.0067495 | 1.2691 | 0.66451 | 165 | 1.3983 | 0.66404 | 60 | 0.42321 | 1.3136 | 1.269 | 2.229 |  | 1.0907 | 1.0406 | 2.1401 |  | 0.057237 | 1.1979 |  | 1.0689 | [11 empty] | 5.3007 | 5.5834 | 4.5848 | 4.2628 | 3.067
| 172.5 |  | 7.987 | 8.109 | 0.011972 | 1.5274 | 0.66358 | 180 | 1.5898 | 0.66345 | 75 | 0.3461 | 1.3062 | 0.39805 | 2.4544 | 1.8203 | 1.1763 | 0.10541 | 2.4848 | 1.7662 | 0.91079 | 1.4435 | 0.69795 | 2.138 | 1.4733 | 0.7155 |  | 1.3417 | [7 empty] | 1.6683 | 1.5792 | 0.59188 | 1.3336 | 0.73318 | 0.05602 | 0.69549 | 0.82316 | 1.1544 | 0.8446 | 1.6375 | 1.3741
| 187.5 |  | 7.99 | 8.122 | 0.0063246 | 1.6523 | 0.66333 | 195 | 1.4807 | 0.66348 | 90 | 0.65911 | 1.3509 | 0.39804 | 2.2918 | 1.8916 | 0.86772 | 0.22034 | 1.9948 | 1.5295 | 0.95641 | 1.2213 | 0.73188 | 1.9698 | 1.5459 | 0.48247 |  | 0.10145 | [7 empty] | 0.40639 | 1.114 | 1.7756 | 0.80989 | 2.6553 | 2.2746 | 0.25721 | 0.48048 | 0.72543 | 0.4045 | 1.1388 | 0.9613
| 202.5 |  | 7.9865 | 8.091 | 0.011005 | 1.309 | 0.66362 | 210 | 0.72889 | 0.6631 | 105 | 0.93498 | 1.1799 | 0.34844 | 2.0518 | 1.5633 | 0.34177 | 0.56519 | 1.4518 | 0.86549 | 0.83396 | 1.1212 | 0.54015 | 1.7703 | 1.2572 | 0.55755 |  | 0.13915 | [7 empty] | 0.34815 | 1.3456 | 1.5749 | 0.70905 | 2.402 | 1.9256 | 0.0085605 | 0.30117 | 0.39711 | 0.11942 | 0.78091 | 0.57277
| 217.5 |  | 7.9991 | 8.011 | 0.015239 | 0.14874 | 0.66257 | 225 |  |  | 120 | 1.0068 | 1.1272 | 0.15707 | 2.0011 | 1.3398 | 0.19502 | 0.86367 | 1.3318 | 0.53198 | 0.99318 | 1.1282 | 0.34471 | 1.9051 | 1.2212 | 0.7419 | 1.6709 | 0.63163 | 0.019091 | 1.0159 | 0.14519 | 0.027834 | 0.88634 | 1.0726 | 0.34269 | 0.63593 | 1.7472 | 1.8374 | 0.81235 | 2.6177 | 1.9891 | 0.39233 | 0.065885 | 0.010008 | 0.3362 | 0.38905 | 0.1051
[8 empty] | 240 |  |  | 135 | 1.1665 | 1.0192 | 0.049974 | 2.0527 | 1.2041 | 0.12029 | 1.155 | 1.3051 | 0.21086 | 1.0009 | 1.3515 | 0.30347 | 2.0553 | 1.19 | 1.0271 | 1.4317 | 1.1288 | 0.61364 | 0.73411 | 0.35256 | 0.23397 | 0.70564 | 1.4582 | 0.49279 | 0.66345 | 1.9698 | 1.7834 | 0.73483 | 2.6604 | 1.9018 | 0.8804 | 0.51008 | 0.54193 | 0.86559 | 0.0034011 | 0.43452
[8 empty] | 255 |  |  | 150 | 1.1274 | 1.1735 | 0.08799 |  | 1.0168 | 0.11131 | 1.0699 |  | 0.032695 | 1.1184 |  | 0.062682 |  | 0.95847 |  | 1.3324 | 1.4402 | 0.96233 | 0.7298 | 0.31589 | 0.3974 | 0.66213 |  | 0.30352 | 0.56724 | 1.8279 | 1.8435 | 0.67522 |  | 1.6194 | 1.2266 | 0.87558 | 0.83971 | 1.1992 |  | 0.85105
[8 empty] | 270 |  |  | 165 | 0.97018 | 1.0384 | 0.31201 |  | 0.89802 | 0.13018 | 0.64423 |  | 0.05486 | 0.72769 |  | 0.05739 |  | 0.60828 |  | 0.58894 | 1.3653 | 1.1543 | 0.55092 | 0.36334 | 0.45903 | 0.24679 |  | 0.36429 | 0.029939 | 1.0212 | 1.0876 | 0.34695 |  | 0.9399 | 1.3518 | 1.0552 | 1.0131 | 1.2425 |  | 1.0207
[8 empty] | 285 |  |  | 180 | 0.79607 | 0.8468 | 0.56969 |  | 0.79981 | 0.1005 | 0.16144 |  | 0.1184 | 0.24065 |  | 0.028512 |  | 0.24504 |  | 0.13362 | 1.4835 | 1.4375 | 0.067945 | 0.74463 | 0.7981 | 0.51527 |  | 0.7521 | 0.063088 | 0.88316 | 0.93048 | 0.6424 |  | 0.87445 | 1.4574 | 1.2205 | 1.1861 | 1.2538 |  | 1.1631
[8 empty] | 300 |  |  | 195 | 0.38942 | 0.65385 | 0.55137 |  | 0.7976 | 0.33843 | 0.24013 |  | 0.53105 | 0.061031 |  | 0.22272 |  | 0.2613 |  | 0.11983 | 1.5022 | 1.5466 | 0.26325 | 0.71026 | 0.96847 | 0.83748 |  | 1.0859 | 0.14508 | 0.56897 | 0.83175 | 0.71241 |  | 0.96125 | 1.4445 | 1.3422 | 1.2394 | 1.2476 |  | 1.151
[8 empty] | 315 |  |  | 210 | 0.44284 | 0.019635 | 0.087691 |  | 0.54705 | 0.4889 | 0.56291 |  | 1.0888 | 0.11841 |  | 0.62974 |  | 0.47432 |  | 0.16761 | 1.4828 | 1.5468 | 0.8574 | 0.54426 | 0.93706 | 0.97322 |  | 1.4308 | 0.68273 | 0.34463 | 0.74003 | 0.79151 |  | 1.2458 | 1.2261 | 1.3956 | 1.2496 | 1.1987 |  | 1.0253
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.70081 | 0.5475 |  | 1.2364 | 0.078244 |  | 0.59212 |  | 0.61826 |  | 0.3276 | 1.4517 | 1.3372 |  | 0.023301 | 0.6223 | 0.50191 |  | 1.1117 |  | 0.39059 | 0.94863 | 0.81124 |  | 1.3948 |  | 1.3496 | 1.1435 | 1.1583 |  | 0.93295
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.46519 | 0.28727 |  | 1.2348 | 0.12474 |  | 0.7895 |  | 0.84289 |  | 0.1402 | 1.2868 | 1.2396 |  | 0.37602 | 0.75821 | 0.58428 |  | 1.4177 |  | 0.52269 | 0.88855 | 0.71064 |  | 1.5231 |  | 1.1377 | 0.99974 | 1.0311 |  | 0.72497
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.50077 | 0.1816 |  | 1.2982 | 0.25632 |  | 0.82025 |  | 0.99091 |  | 0.040959 | 1.2163 | 1.2071 |  | 0.48701 | 0.89523 | 0.5892 |  | 1.5712 |  | 0.5105 | 0.90133 | 0.60868 |  | 1.5558 |  | 1.0693 | 0.92156 | 0.99696 |  | 0.63742
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.054034 |  | 0.96049 |  | 0.9366 |  | 0.23356 | 1.1433 | 1.0605 |  |  | 0.69786 | 0.59402 |  | 1.5191 |  |  | 0.93787 | 0.81867 |  | 1.7261 |  |  | 0.83867 | 0.84838 |  | 0.50871
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.54285 |  | 0.48159 |  | 0.93835 |  | 0.056103 | 0.89524 | 0.87919 |  |  | 1.2077 | 0.61096 |  | 1.5431 |  |  | 1.2133 | 0.64317 |  | 1.54 |  |  | 0.48788 | 0.66161 |  | 0.32166
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.47259 |  | 0.44635 |  | 0.84195 |  | 0.28383 | 0.82919 | 0.72808 |  |  | 0.98124 | 0.46932 |  | 1.304 |  |  | 1.2451 | 0.73599 |  | 1.5374 |  |  | 0.40928 | 0.56067 |  | 0.25571
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.37883 |  | 0.46865 |  | 0.77732 |  | 0.36945 | 0.83038 | 0.6979 |  |  | 0.99449 | 0.56759 |  | 1.3368 |  |  | 1.3628 | 0.92213 |  | 1.667 |  |  | 0.37455 | 0.49547 |  | 0.21395
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.47431 |  | 0.45382 |  | 0.85039 |  | 0.50122 | 0.88774 | 0.70489 |  |  | 1.143 | 0.62058 |  | 1.4568 |  |  | 1.6493 | 1.1075 |  | 1.9146 |  |  | 0.32995 | 0.48268 |  | 0.17479
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.55939 |  | 0.64375 |  | 1.1032 |  | 0.11944 | 0.79769 | 0.75728 |  |  | 1.2097 | 0.599 |  | 1.6945 |  |  | 1.2903 | 0.69793 |  | 1.756 |  |  | 0.36337 | 0.54375 |  | 0.14448
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.41161 | [5 empty] | 0.097207 | 0.77571 | 0.74404 |  |  | 1.2101 | 0.73584 |  |  |  |  | 1.254 | 0.80129 |  |  |  |  | 0.34978 | 0.48171
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.32785 | [5 empty] | 0.048106 |  |  |  |  | 1.1962 | 0.80038 |  |  |  |  | 1.0926 | 0.72258
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.28591 | [5 empty] | 0.032772 |  |  |  |  | 1.5428 | 1.159 |  |  |  |  | 1.5158 | 1.1517
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.4676 | [5 empty] | 0.17606 |  |  |  |  | 1.5582 | 1.0056 |  |  |  |  | 1.2841 | 0.78254
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.36681 | [5 empty] | 1.2059 | [5 empty] | 0.81323
[11 empty] | 435 | [27 empty] | 0.84055
[11 empty] | 450 | [27 empty] | 0.95252
45
PTB | 55 | 8.009 | 8.0488 | 7.7552 | 0.013097 | -3.6474 | 0.43 | 60 | -3.3156 | 0.43
| 70 | 8.009 | 8.0414 | 7.8282 | 0.0079501 | -2.6519 | 0.43 | 75 | -2.2506 | 0.43
| 85 | 8.009 | 8.0342 | 7.9178 | 0.012777 | -1.4481 | 0.43 | 90 | -1.2274 | 0.43
| 100 | 8.009 | 8.0283 | 7.9652 | 0.014419 | -0.7861 | 0.43 | 105 | -0.68008 | 0.43
| 115 | 8.009 | 8.0232 | 7.9857 | 0.0078775 | -0.46804 | 0.43 | 120 | -0.43457 | 0.43
| 130 | 8.009 | 8.0172 | 7.9877 | 0.012318 | -0.36762 | 0.43 | 135 | -0.29422 | 0.43
| 145 | 8.009 | 8.0106 | 7.9988 | 0.01227 | -0.14742 | 0.43 | 150 | 0.047033 | 0.43
| 160 | 8.009 | 8.009 | 8.0439 | 0.0089133 | 0.43595 | 0.43 | 165 | 0.55854 | 0.43
| 175 | 8.009 | 8.0102 | 8.0745 | 0.0073855 | 0.80374 | 0.43 | 180 | 0.90118 | 0.43
| 190 | 8.009 | 8.0142 | 8.1021 | 0.012151 | 1.096 | 0.43 | 195 | 1.1438 | 0.43
| 205 | 8.009 | 8.0185 | 8.1179 | 0.0093153 | 1.2392 | 0.43 | 210 | 1.1119 | 0.43
| 220 | 8.009 | 8.0217 | 8.0904 | 0.013015 | 0.8571 | 0.43 | 225 | 0.62609 | 0.43
| 235 | 8.009 | 8.0329 | 8.0461 | 0.013052 | 0.16408 | 0.43 | 240 | 0.048961 | 0.43
| 250 | 8.009 | 8.0375 | 8.0229 | 0.019667 | -0.18127 | 0.43 | 255 | -0.40038 | 0.43
| 265 | 8.009 | 8.0501 | 7.9826 | 0.017637 | -0.83859 | 0.43 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 8.0908 | 7.656 | 0.010463 | -5.3742 | 0.5 | 45 | -5.3742 | 0.5
| 60 |  | 8.0844 | 7.7505 | 0.0088704 | -4.1304 | 0.5 | 60 | -4.1304 | 0.5
| 75 |  | 8.0819 | 7.829 | 0.0078807 | -3.1295 | 0.5 | 75 | -3.1295 | 0.5
| 90 |  | 8.0887 | 7.937 | 0.012183 | -1.8757 | 0.5 | 90 | -1.8757 | 0.5
| 105 |  | 8.0841 | 8.0085 | 0.0098809 | -0.93541 | 0.5 | 105 | -0.93541 | 0.5
| 120 |  | 8.0859 | 8.039 | 0.0071818 | -0.58027 | 0.5 | 120 | -0.58027 | 0.5
| 135 |  | 8.084 | 8.0675 | 0.0071635 | -0.20435 | 0.5 | 135 | -0.20435 | 0.5
| 150 |  | 8.0824 | 8.0795 | 0.0094451 | -0.036128 | 0.5 | 150 | -0.036128 | 0.5
| 165 |  | 8.0817 | 8.119 | 0.0071818 | 0.46129 | 0.5 | 165 | 0.46129 | 0.5
| 180 |  | 8.0778 | 8.1445 | 0.0082558 | 0.82609 | 0.5 | 180 | 0.82609 | 0.5
| 195 |  | 8.0736 | 8.1455 | 0.0068633 | 0.89093 | 0.5 | 195 | 0.89093 | 0.5
| 210 |  | 8.0807 | 8.141 | 0.010712 | 0.7466 | 0.5 | 210 | 0.7466 | 0.5
| 225 | 8.0772 | 8.0772 | 8.0855 | 0.013169 | 0.10251 | 0.5 | 225 | 0.10251 | 0.5
| 240 |  | 8.0781 | 8.054 | 0.010463 | -0.29858 | 0.5 | 240 | -0.29858 | 0.5
| 255 |  | 8.0851 | 8.0225 | 0.016819 | -0.77451 | 0.5 | 255 | -0.77451 | 0.5
| 270 |  | 8.0802 | 7.996 | 0.012312 | -1.0423 | 0.5 | 270 | -1.0423 | 0.5
| 285 |  | 8.0823 | 7.9355 | 0.01191 | -1.8166 | 0.5 | 285 | -1.8166 | 0.5
| 300 |  | 8.0829 | 7.915 | 0.013572 | -2.0775 | 0.5 | 300 | -2.0775 | 0.5
| 315 |  | 8.0838 | 7.905 | 0.01 | -2.2121 | 0.5 | 315 | -2.2121 | 0.5
| 330 |  | 8.0849 | 7.8905 | 0.011459 | -2.4047 | 0.5 | 330 | -2.4047 | 0.5
| 345 |  | 8.0868 | 7.8795 | 0.016051 | -2.5637 | 0.5 | 345 | -2.5637 | 0.5
| 360 |  | 8.0881 | 7.8835 | 0.017852 | -2.5299 | 0.5 | 360 | -2.5299 | 0.5
| 375 |  | 8.0955 | 7.8885 | 0.015985 | -2.5572 | 0.5 | 375 | -2.5572 | 0.5
| 390 |  | 8.0931 | 7.888 | 0.017652 | -2.5345 | 0.5 | 390 | -2.5345 | 0.5
| 405 |  | 8.0804 | 7.8595 | 0.01191 | -2.734 | 0.5 | 405 | -2.734 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 7.9944 | 7.95 | 0.04 | -0.55594 | 0.625 | 120 | -0.55594 | 0.625
| 135 |  | 7.99 | 7.99 | 0.03 | -0.00046811 | 0.62135 | 135 | -0.00046811 | 0.62135
| 150 |  | 7.9952 | 8.02 | 0.03 | 0.31058 | 0.62282 | 150 | 0.31058 | 0.62282
| 165 |  | 7.9912 | 8.06 | 0.02 | 0.86124 | 0.62122 | 165 | 0.86124 | 0.62122
| 180 |  | 7.9982 | 8.12 | 0.02 | 1.5235 | 0.6247 | 180 | 1.5235 | 0.6247
| 195 |  | 8.0009 | 8.14 | 0.02 | 1.7379 | 0.62578 | 195 | 1.7379 | 0.62578
| 210 |  | 8.0044 | 8.13 | 0.03 | 1.5687 | 0.62967 | 210 | 1.5687 | 0.62967
| 225 |  | 8.0182 | 8.07 | 0.03 | 0.64554 | 0.62304 | 225 | 0.64554 | 0.62304
| 240 |  | 8.011 | 8.04 | 0.03 | 0.36242 | 0.6219 | 240 | 0.36242 | 0.6219
| 255 |  | 7.9996 | 8 | 0.03 | 0.0051452 | 0.62064 | 255 | 0.0051452 | 0.62064
| 270 | 8.004 | 8.0046 | 7.97 | 0.04 | -0.43239 | 0.62494 | 270 | -0.43239 | 0.62494
| 285 |  | 8.0015 | 7.94 | 0.03 | -0.76891 | 0.61583 | 285 | -0.76891 | 0.61583
| 300 |  | 7.9978 | 7.9 | 0.04 | -1.2228 | 0.62078 | 300 | -1.2228 | 0.62078
| 315 |  | 7.9977 | 7.89 | 0.04 | -1.3464 | 0.62005 | 315 | -1.3464 | 0.62005
| 330 |  | 7.992 | 7.88 | 0.05 | -1.402 | 0.62955 | 330 | -1.402 | 0.62955
| 345 |  | 8.0009 | 7.88 | 0.04 | -1.5117 | 0.61882 | 345 | -1.5117 | 0.61882
| 360 |  | 7.9982 | 7.88 | 0.04 | -1.4772 | 0.61924 | 360 | -1.4772 | 0.61924
| 375 |  | 7.9912 | 7.87 | 0.04 | -1.5164 | 0.61955 | 375 | -1.5164 | 0.61955
| 390 |  | 7.9952 | 7.9 | 0.04 | -1.1903 | 0.62117 | 390 | -1.1903 | 0.62117
| 405 |  | 7.99 | 7.88 | 0.04 | -1.3772 | 0.62046 | 405 | -1.3772 | 0.62046
| 420 |  | 7.9944 | 7.9 | 0.06 | -1.1814 | 0.64193 | 420 | -1.1814 | 0.64193
435
450
465
480
495
510
525
45
DTI | 60 |  | 8.0198 | 7.3205 | 0.029465 | -8.7193 | 0.79279 | 60 | -8.7193 | 0.79279
| 75 |  | 8.0329 | 7.7365 | 0.024979 | -3.6895 | 0.72245 | 75 | -3.6895 | 0.72245
| 90 |  | 8.0241 | 8.0045 | 0.022821 | -0.24377 | 0.68664 | 90 | -0.24377 | 0.68664
| 105 |  | 8.0373 | 8.0755 | 0.017313 | 0.47549 | 0.65522 | 105 | 0.47549 | 0.65522
| 120 |  | 8.0245 | 8.1105 | 0.01905 | 1.0719 | 0.65971 | 120 | 1.0719 | 0.65971
| 135 |  | 8.0184 | 8.128 | 0.012814 | 1.3672 | 0.63504 | 135 | 1.3672 | 0.63504
| 150 |  | 8.0101 | 8.1365 | 0.009333 | 1.5774 | 0.62668 | 150 | 1.5774 | 0.62668
| 165 |  | 8.0234 | 8.138 | 0.016092 | 1.4279 | 0.64544 | 165 | 1.4279 | 0.64544
| 180 |  | 8.0249 | 8.1575 | 0.016182 | 1.6522 | 0.64423 | 180 | 1.6522 | 0.64423
| 195 |  | 8.0233 | 8.1535 | 0.013089 | 1.6232 | 0.6339 | 195 | 1.6232 | 0.6339
| 210 |  | 8.0202 | 8.133 | 0.01689 | 1.4062 | 0.64891 | 210 | 1.4062 | 0.64891
| 225 |  | 8.0238 | 8.1015 | 0.022775 | 0.96841 | 0.67818 | 225 | 0.96841 | 0.67818
| 240 |  | 8.0136 | 8.0535 | 0.017252 | 0.49845 | 0.65677 | 240 | 0.49845 | 0.65677
| 255 | 8.0035 | 8.0154 | 8.019 | 0.020494 | 0.04533 | 0.67386 | 255 | 0.04533 | 0.67386
| 270 |  | 8.0136 | 7.997 | 0.013803 | -0.20661 | 0.64862 | 270 | -0.20661 | 0.64862
| 285 |  | 8.0238 | 7.9665 | 0.01927 | -0.71408 | 0.67262 | 285 | -0.71408 | 0.67262
| 300 |  | 8.0202 | 7.9445 | 0.019324 | -0.94415 | 0.67473 | 300 | -0.94415 | 0.67473
| 315 |  | 8.0233 | 7.944 | 0.015355 | -0.98796 | 0.65842 | 315 | -0.98796 | 0.65842
| 330 |  | 8.0249 | 7.952 | 0.018525 | -0.90855 | 0.67054 | 330 | -0.90855 | 0.67054
| 345 |  | 8.0234 | 7.9115 | 0.017252 | -1.3951 | 0.66855 | 345 | -1.3951 | 0.66855
| 360 |  | 8.0101 | 7.8995 | 0.020125 | -1.3814 | 0.6823 | 360 | -1.3814 | 0.6823
| 375 |  | 8.0184 | 7.893 | 0.018093 | -1.5635 | 0.67367 | 375 | -1.5635 | 0.67367
| 390 |  | 8.0098 | 7.917 | 0.018093 | -1.1582 | 0.67163 | 390 | -1.1582 | 0.67163
| 405 |  | 8.0199 | 7.8955 | 0.021392 | -1.5517 | 0.6888 | 405 | -1.5517 | 0.6888
| 420 |  | 8.0188 | 7.895 | 0.016384 | -1.5444 | 0.66644 | 420 | -1.5444 | 0.66644
| 435 |  | 8.0075 | 7.884 | 0.025005 | -1.5426 | 0.70908 | 435 | -1.5426 | 0.70908
| 450 |  | 8.0196 | 7.903 | 0.020287 | -1.454 | 0.68277 | 450 | -1.454 | 0.68277
465
480
495
510
525
45
DWG | 50 |  | 8.0687 | 7.6045 | 0.044885 | -5.7527 | 0.62071 | 60 | -4.1802 | 0.61812
| 65 |  | 8.1107 | 7.8354 | 0.025825 | -3.3939 | 0.61683 | 75 | -2.3382 | 0.61767
| 80 |  | 8.0907 | 7.9442 | 0.017844 | -1.8103 | 0.6181 | 90 | -1.0443 | 0.61834
| 95 |  | 8.0847 | 8.0312 | 0.017333 | -0.66133 | 0.61847 | 105 | -0.21027 | 0.61871
| 110 |  | 8.0797 | 8.0809 | 0.0072833 | 0.015265 | 0.61884 | 120 | 0.28357 | 0.61905
| 125 |  | 8.0757 | 8.1094 | 0.013118 | 0.41772 | 0.61915 | 135 | 0.66626 | 0.61921
| 140 |  | 8.0747 | 8.1385 | 0.0076537 | 0.79054 | 0.61924 | 150 | 0.93692 | 0.6194
| 155 |  | 8.0717 | 8.1532 | 0.0081701 | 1.0101 | 0.61948 | 165 | 1.0945 | 0.61954
| 170 |  | 8.0707 | 8.1624 | 0.01052 | 1.1366 | 0.61957 | 180 | 1.0354 | 0.61931
| 185 |  | 8.0757 | 8.1552 | 0.0099463 | 0.98485 | 0.61917 | 195 | 0.94403 | 0.61948
| 200 |  | 8.0697 | 8.1442 | 0.012467 | 0.92362 | 0.61963 | 210 | 0.64356 | 0.61962
| 215 |  | 8.0697 | 8.1103 | 0.019691 | 0.50353 | 0.61961 | 225 | 0.1706 | 0.61966
| 230 |  | 8.0687 | 8.069 | 0.013054 | 0.0041312 | 0.61968 | 240 | -0.19004 | 0.61968
| 245 |  | 8.0687 | 8.0455 | 0.015757 | -0.28712 | 0.61968 | 255 | -0.55146 | 0.61964
| 260 |  | 8.0697 | 8.0145 | 0.018719 | -0.68363 | 0.61962 | 270 | -0.99527 | 0.61974
| 275 |  | 8.0677 | 7.9748 | 0.012308 | -1.1511 | 0.6198 | 285 | -1.3442 | 0.61969
| 290 |  | 8.0702 | 7.9539 | 0.017409 | -1.4407 | 0.61963 | 300 | -1.6662 | 0.61965
| 305 |  | 8.0702 | 7.9266 | 0.012384 | -1.779 | 0.61966 | 315 | -1.8824 | 0.61973
| 320 |  | 8.0692 | 7.9131 | 0.012161 | -1.9341 | 0.61976 | 330 | -1.9919 | 0.61977
| 335 |  | 8.0692 | 7.9061 | 0.014369 | -2.0209 | 0.61977 | 345 | -2.0769 | 0.61973
| 350 |  | 8.0702 | 7.9003 | 0.011541 | -2.1049 | 0.6197 | 360 | -2.1717 | 0.61969
| 365 |  | 8.0707 | 7.8927 | 0.011625 | -2.2051 | 0.61968 | 375 | -2.2719 | 0.61966
| 380 |  | 8.0712 | 7.8851 | 0.010494 | -2.3053 | 0.61965 | 390 | -2.2857 | 0.6197
| 395 |  | 8.0702 | 7.8865 | 0.011904 | -2.2759 | 0.61973 | 405 | -2.3271 | 0.61973
| 410 |  | 8.0702 | 7.8803 | 0.01473 | -2.3527 | 0.61974 | 420 | -2.3376 | 0.61968
| 425 |  | 8.0712 | 7.8831 | 0.012866 | -2.3301 | 0.61966 | 435 | -2.3874 | 0.61974
| 440 |  | 8.0697 | 7.8747 | 0.014257 | -2.416 | 0.61979 | 450 | -2.3938 | 0.61981
| 455 |  | 8.0692 | 7.8769 | 0.01519 | -2.3827 | 0.61982 | 465 | -2.4027 | 0.61985
| 470 |  | 8.0687 | 7.874 | 0.011878 | -2.4126 | 0.61986 | 480 | -2.4399 | 0.61987
| 485 |  | 8.0687 | 7.8707 | 0.012844 | -2.4535 | 0.61987 | 495 | -2.461 | 0.61987
| 500 | 8.069 | 8.0687 | 7.8698 | 0.015121 | -2.4647 | 0.61987 | 510 | -2.4153 | 0.61991
| 515 |  | 8.0677 | 7.8748 | 0.013308 | -2.3906 | 0.61993 | 525 | -2.4239 | 0.61999
| 530 |  | 8.0667 | 7.8698 | 0.014057 | -2.4405 | 0.62002
UL-LMPS | 42.5 | 8.1951 | 8.2041 | 7.6553 | 0.010417 | -6.6887 | 0.69813 | 45 | -6.4689 | 0.70253
| 57.5 | 8.1219 | 8.0975 | 7.6627 | 0.014126 | -5.3696 | 0.72456 | 60 | -5.262 | 0.72041
| 72.5 | 8.0635 | 8.0471 | 7.667 | 0.014657 | -4.724 | 0.6997 | 75 | -4.4706 | 0.70002
| 87.5 | 8.0015 | 7.9821 | 7.7263 | 0.018659 | -3.2041 | 0.70166 | 90 | -3.0117 | 0.70136
| 102.5 | 7.9313 | 7.9374 | 7.7747 | 0.015025 | -2.0497 | 0.69986 | 105 | -1.9773 | 0.7001
| 117.5 | 7.8806 | 7.8847 | 7.7573 | 0.017991 | -1.6149 | 0.7013 | 120 | -1.6465 | 0.72108
| 132.5 | 7.8304 | 7.8408 | 7.6993 | 0.095337 | -1.8045 | 0.81998 | 135 | -1.6346 | 0.86407
| 147.5 | 7.7719 | 7.7841 | 7.723 | 0.017251 | -0.78542 | 1.0845 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | 64 |  | 8.1355 | 8.0965 | 0.022542 | -0.47913 | 2.55 | 75 | -0.10241 | 2.55
| 79 |  | 8.1467 | 8.1495 | 0.025021 | 0.034582 | 2.55 | 90 | 0.026508 | 2.55
| 94 |  | 8.1521 | 8.154 | 0.027028 | 0.023573 | 2.55 | 105 | 0.10589 | 2.55
| 109 |  | 8.1255 | 8.1365 | 0.027004 | 0.13582 | 2.55 | 120 | -0.60651 | 2.55
| 124 |  | 8.106 | 8.035 | 0.040066 | -0.87645 | 2.55 | 135 | -1.6254 | 2.55
| 139 |  | 8.0987 | 7.945 | 0.027625 | -1.8977 | 2.55 | 150 | -2.238 | 2.55
| 154 |  | 8.0819 | 7.891 | 0.023147 | -2.3617 | 2.55 | 165 | -2.1953 | 2.55
| 169 |  | 8.0698 | 7.8975 | 0.039454 | -2.1347 | 2.55 | 180 | -2.2841 | 2.55
| 184 |  | 8.0718 | 7.883 | 0.030105 | -2.3384 | 2.55 | 195 | -2.359 | 2.55
| 199 |  | 8.0771 | 7.886 | 0.023486 | -2.3665 | 2.55 | 210 | -2.5302 | 2.55
| 214 |  | 8.0895 | 7.88 | 0.025131 | -2.5897 | 2.55 | 225 | -2.896 | 2.55
| 229 |  | 8.1181 | 7.874 | 0.035895 | -3.0074 | 2.55 | 240 | -2.9201 | 2.55
| 244 |  | 8.1406 | 7.9055 | 0.019861 | -2.8883 | 2.55 | 255 | -3.191 | 2.55
| 259 |  | 8.1464 | 7.8775 | 0.032907 | -3.301 | 2.55 | 270 | -3.2414 | 2.55
| 274 |  | 8.1499 | 7.8875 | 0.026926 | -3.2198 | 2.55 | 285 | -3.0959 | 2.55
| 289 |  | 8.1393 | 7.891 | 0.026137 | -3.0508 | 2.55 | 300 | -3.1507 | 2.55
| 304 |  | 8.1182 | 7.8595 | 0.035015 | -3.187 | 2.55 | 315 | -3.1942 | 2.55
| 319 | 8.101 | 8.101 | 7.842 | 0.023079 | -3.1968 | 2.55 | 330 | -3.2699 | 2.55
| 334 |  | 8.07 | 7.804 | 0.022804 | -3.2965 | 2.55 | 345 | -3.5165 | 2.55
| 349 |  | 8.05 | 7.7605 | 0.020641 | -3.5965 | 2.55 | 360 | -3.4471 | 2.55
| 364 |  | 8.0703 | 7.7965 | 0.025189 | -3.3927 | 2.55 | 375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | 68 |  | 8.091 | 7.751 | 0.008 | -4.202 | 0.64 | 75 | -3.747 | 0.64
| 83 |  | 8.088 | 7.827 | 0.012 | -3.227 | 0.64 | 90 | -2.5233 | 0.64
| 98 |  | 8.086 | 7.947 | 0.01 | -1.719 | 0.64 | 105 | -1.4133 | 0.64
| 113 |  | 8.08 | 7.994 | 0.01 | -1.064 | 0.64 | 120 | -0.88527 | 0.64
| 128 |  | 8.082 | 8.027 | 0.011 | -0.681 | 0.64 | 135 | -0.47287 | 0.64
| 143 |  | 8.077 | 8.058 | 0.008 | -0.235 | 0.64 | 150 | 0.019333 | 0.64
| 158 |  | 8.076 | 8.101 | 0.011 | 0.31 | 0.64 | 165 | 0.51207 | 0.64
| 173 |  | 8.076 | 8.136 | 0.006 | 0.743 | 0.64 | 180 | 0.80087 | 0.64
| 188 |  | 8.074 | 8.144 | 0.009 | 0.867 | 0.64 | 195 | 0.69387 | 0.64
| 203 |  | 8.068 | 8.108 | 0.018 | 0.496 | 0.64 | 210 | 0.1894 | 0.64
| 218 |  | 8.062 | 8.049 | 0.015 | -0.161 | 0.64 | 225 | -0.4214 | 0.64
| 233 |  | 8.064 | 8.006 | 0.015 | -0.719 | 0.64 | 240 | -0.99713 | 0.64
| 248 |  | 8.062 | 7.956 | 0.015 | -1.315 | 0.64 | 255 | -1.5003 | 0.64
| 263 |  | 8.061 | 7.923 | 0.009 | -1.712 | 0.64 | 270 | -1.8921 | 0.64
| 278 |  | 8.057 | 7.888 | 0.013 | -2.098 | 0.64 | 285 | -2.2427 | 0.64
| 293 |  | 8.058 | 7.864 | 0.011 | -2.408 | 0.64 | 300 | -2.4724 | 0.64
| 308 |  | 8.053 | 7.848 | 0.007 | -2.546 | 0.64 | 315 | -2.6267 | 0.64
| 323 |  | 8.055 | 7.836 | 0.014 | -2.719 | 0.64 | 330 | -2.8063 | 0.64
| 338 |  | 8.053 | 7.819 | 0.009 | -2.906 | 0.64 | 345 | -3.1333 | 0.64
| 353 |  | 8.046 | 7.773 | 0.016 | -3.393 | 0.64 | 360
| 368 |  | 8.05 | 7.615 | 0.02 | -5.404 | 0.64 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-S 12
| 12 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.42458
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 11.9369 | 11.529 | 0.008756 | -3.4171 | 0.63668 | 45 | -2.6971 | 0.64103
| 52.5 |  | 11.9309 | 11.695 | 0.010801 | -1.977 | 0.64538 | 60 | -1.4976 | 0.64518 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 11.9386 | 11.817 | 0.010593 | -1.0183 | 0.64497 | 75 | -0.39433 | 0.64473 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 11.9476 | 11.975 | 0.011785 | 0.22963 | 0.64448 | 90 | 0.81722 | 0.64879 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 11.9432 | 12.111 | 0.0099443 | 1.4048 | 0.65309 | 105 | 1.5214 | 0.64871 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 11.9503 | 12.146 | 0.0069921 | 1.6379 | 0.64434 | 120 | 1.8111 | 0.64474 |  | 0.34205 | 2.3546 | 4.0417 | 3.8811 | 1.4588 | 2.3953 | 4.3451 | 4.0708 | 1.3355 | 1.9538 | 1.9562 | 0.96539 | 0.19567 | 2.8543 | 2.7693 | 0.20068 |  |  | 2.3188 | 2.3273 | 0.09313 | 1.751 | 1.7819 | 0.99864 | 2.2001 | 2.2029 | 0.11663 | 2.0191 | 2.0193 | 0.82606
| 127.5 | 11.9437 | 11.9352 | 12.172 | 0.0078881 | 1.9843 | 0.64515 | 135 | 2.1063 | 0.64471 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 11.9517 | 12.218 | 0.0078881 | 2.2284 | 0.64426 | 150 | 2.3975 | 0.64488 | 45 |  | 1.6535 |  | 2.5281 | [6 empty] | 1.1561
| 157.5 |  | 11.9288 | 12.235 | 0.0052705 | 2.5665 | 0.64549 | 165 | 2.5606 | 0.64527 | 60 | 0.37953 | 1.6016 | 2.5816 | 2.6261 |  | 1.4635 | 2.6018 | 2.6133 |  | 1.1439 | 1.3142 |  | 0.2764 | [11 empty] | 7.2575 | 7.917 | 6.3386 | 5.1011 | 4.4139
| 172.5 |  | 11.9371 | 12.242 | 0.0063246 | 2.5546 | 0.64505 | 180 | 2.3783 | 0.64902 | 75 | 0.34368 | 1.5374 | 1.3114 | 2.8033 | 1.4395 | 1.4277 | 1.1589 | 2.8476 | 1.3106 | 0.1957 | 1.5686 | 0.088595 | 1.6925 | 0.10784 | 1.6281 | [9 empty] | 4.1046 | 4.3763 | 2.9305 | 3.0131 | 1.1046 | 2.9757
| 187.5 |  | 11.945 | 12.208 | 0.013166 | 2.2021 | 0.65299 | 195 | 1.6739 | 0.65331 | 90 | 0.58903 | 1.3366 | 0.9447 | 2.5617 | 1.4308 | 0.9164 | 0.46758 | 2.3439 | 1.0309 | 0.38754 | 1.5024 | 0.12254 | 1.7949 | 0.50011 | 1.3801 | [9 empty] | 1.4869 | 1.1038 | 0.22569 | 0.58739 | 1.2496 | 0.10604
| 202.5 |  | 11.9333 | 12.07 | 0.018257 | 1.1456 | 0.65363 | 210 | 0.44896 | 0.64833 | 105 | 0.83133 | 1.2391 | 0.72439 | 2.1959 | 1.1673 | 0.5265 | 0.051959 | 1.6954 | 0.45047 | 0.52658 | 1.1971 | 0.063755 | 1.6195 | 0.45818 | 1.2391 | [9 empty] | 0.74246 | 0.054236 | 0.54365 | 0.00040102 | 1.6583 | 0.47268
| 217.5 |  | 11.9747 | 11.945 | 0.016499 | -0.24772 | 0.64302 | 225 |  |  | 120 | 0.88694 | 1.0628 | 0.53983 | 2.0654 | 0.82348 | 0.25424 | 0.32624 | 1.4911 | 0.016412 | 0.53995 | 1.223 | 0.2521 | 1.6559 | 0.29046 | 1.4277 | 0.75554 |  |  | 1.2086 | 0.46704 | 0.21855 | 0.72046 | 0.97569 | 0.45302 | 0.52801 | 0.35826 | 0.57431 | 0.022294 | 1.7018 | 0.31917
[8 empty] | 240 |  |  | 135 | 1.0088 | 0.85556 | 0.57516 | 2.1103 | 0.71897 | 0.128 | 0.41628 | 1.4374 | 0.27453 | 0.27902 | 1.483 | 0.1406 | 1.677 | 0.14019 | 1.5861 | 0.27573 |  |  | 0.88584 | 0.033005 | 0.081884 | 0.34171 | 1.3453 | 0.21283 | 0.66519 | 0.34631 | 0.20717 | 0.078523 | 1.6507 | 0.064429
[8 empty] | 255 |  |  | 150 | 1.0944 | 0.92019 | 0.77553 |  | 0.62742 | 0.14828 | 0.27475 |  | 0.47719 | 0.12638 |  | 0.31346 |  | 0.17639 |  | 0.028253 |  |  | 0.67456 | 0.36344 | 0.21518 | 0.088919 |  | 0.080724 | 0.72925 | 0.35181 | 0.19577 | 0.064032 |  | 0.11442
[8 empty] | 270 |  |  | 165 | 0.94503 | 0.73576 | 1.0254 |  | 0.063766 | 0.1976 | 0.1776 |  | 0.9743 | 0.34753 |  | 0.73784 |  | 1.0515 |  | 0.42367 |  |  | 0.45565 | 0.46112 | 0.26458 | 0.58132 |  | 0.4291 | 0.88905 | 0.017896 | 0.17 | 0.18498 |  | 0.9065
[8 empty] | 285 |  |  | 180 | 0.67102 | 0.39016 | 1.0148 |  | 0.035181 | 0.2976 | 0.4614 |  | 0.78886 | 0.70745 |  | 0.46936 |  | 1.1487 |  | 0.64156 |  |  | 0.24507 | 0.9866 | 0.67726 | 1.3158 |  | 0.23044 | 0.3592 | 0.32493 | 0.029976 | 0.7277 |  | 0.43384
[8 empty] | 300 |  |  | 195 | 0.24802 | 0.32671 | 0.57678 |  | 0.48665 | 0.10647 | 0.40754 |  | 0.29844 | 0.29027 |  | 0.18326 |  | 0.11034 |  | 0.39053 |  |  | 0.85426 | 1.2519 | 1.2855 | 1.5109 |  | 1.4396 | 0.52831 | 0.90704 | 0.95713 | 1.2074 |  | 1.1256
[8 empty] | 315 |  |  | 210 | 0.27048 | 0.17378 | 0.29279 |  | 0.17683 | 0.10033 | 0.050041 |  | 0.50949 | 0.13837 |  | 0.38863 |  | 0.50972 |  | 0.40093 |  |  | 1.1115 | 1.0053 | 1.0452 | 0.87879 |  | 1.3939 | 1.5448 | 1.5178 | 1.5298 | 1.3344 |  | 1.8887
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.29066 | 0.2823 |  | 0.43859 | 0.53117 |  | 0.14445 |  | 0.66271 |  | 0.13814 |  |  |  | 0.98114 | 1.1907 | 0.64711 |  | 1.311 |  | 1.1719 | 1.3767 | 0.80796 |  | 1.4978
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.32852 | 0.3394 |  | 0.59863 | 0.61964 |  | 0.26062 |  | 0.86154 |  | 0.15564 |  |  |  | 1.0429 | 1.2833 | 0.65189 |  | 1.5119 |  | 0.93133 | 1.1878 | 0.52588 |  | 1.429
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.27715 | 0.36407 |  | 0.72004 | 0.59651 |  | 0.42328 |  | 0.99491 |  | 0.10402 |  |  |  | 0.91876 | 1.1195 | 0.51637 |  | 1.503 |  | 0.84073 | 1.0514 | 0.42823 |  | 1.4493
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.53391 |  | 0.3821 |  | 0.8936 |  | 0.19708 |  |  |  |  | 1.1314 | 0.58736 |  | 1.4756 |  |  | 0.96422 | 0.40633 |  | 1.3228
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.79818 |  | 0.0040933 |  | 0.79132 |  | 0.26289 |  |  |  |  | 1.2513 | 0.45685 |  | 1.2391 |  |  | 1.0371 | 0.20963 |  | 1.0264
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.77602 |  | 0.046349 |  | 0.72102 |  | 0.29897 |  |  |  |  | 1.2176 | 0.44481 |  | 1.1585 |  |  | 0.95721 | 0.15811 |  | 0.8983
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.8006 |  | 0.045398 |  | 0.74619 |  | 0.060018 |  |  |  |  | 1.2549 | 0.45151 |  | 1.1957 |  |  | 1.2284 | 0.40279 |  | 1.1674
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.64144 |  | 0.21564 |  | 0.83985 |  | 0.1459 |  |  |  |  | 1.254 | 0.59953 |  | 1.4415 |  |  | 1.1509 | 0.47401 |  | 1.3466
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.56957 |  | 1.1278 |  | 1.645 |  | 0.22246 |  |  |  |  | 0.9417 | 0.36763 |  | 1.995 |  |  | 0.74725 | 0.15631 |  | 1.8473
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.52163 | [5 empty] | 0.33111 |  |  |  |  | 1.0326 | 0.5007 |  |  |  |  | 0.70405 | 0.17381
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.52842 | [5 empty] | 0.41667 |  |  |  |  | 1.0106 | 0.47294 |  |  |  |  | 0.60796 | 0.066
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.69182 | [5 empty] | 0.15481 |  |  |  |  | 1.2908 | 0.59375 |  |  |  |  | 1.1769 | 0.45696
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.88761 | [5 empty] | 0.64771 |  |  |  |  | 1.6934 | 0.79446 |  |  |  |  | 1.0512 | 0.15319
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.40723 | [5 empty] | 0.6964 | [5 empty] | 0.30072
[11 empty] | 435 | [27 empty] | 0.32937
[11 empty] | 450 | [27 empty] | 0.31821
45
PTB | 55 | 11.995 | 12.0575 | 11.7922 | 0.016225 | -2.2008 | 0.42 | 60 | -1.8313 | 0.42
| 70 | 11.995 | 12.0444 | 11.9128 | 0.016713 | -1.0923 | 0.42 | 75 | -0.69637 | 0.42
| 85 | 11.995 | 12.03 | 12.0415 | 0.014884 | 0.095589 | 0.42 | 90 | 0.29781 | 0.42
| 100 | 11.995 | 12.0212 | 12.1056 | 0.017098 | 0.70224 | 0.42 | 105 | 0.78832 | 0.42
| 115 | 11.995 | 12.0184 | 12.1338 | 0.0085215 | 0.96047 | 0.42 | 120 | 1.0316 | 0.42
| 130 | 11.995 | 12.0056 | 12.1465 | 0.0079711 | 1.1739 | 0.42 | 135 | 1.2197 | 0.42
| 145 | 11.995 | 11.9995 | 12.1569 | 0.0088405 | 1.3115 | 0.42 | 150 | 1.4355 | 0.42
| 160 | 11.995 | 11.995 | 12.1969 | 0.014358 | 1.6836 | 0.42 | 165 | 1.7297 | 0.42
| 175 | 11.995 | 12.0033 | 12.222 | 0.0086603 | 1.8219 | 0.42 | 180 | 1.7865 | 0.42
| 190 | 11.995 | 12.0103 | 12.2164 | 0.018682 | 1.7158 | 0.42 | 195 | 1.4543 | 0.42
| 205 | 11.995 | 12.0139 | 12.1258 | 0.029553 | 0.93142 | 0.42 | 210 | 0.68738 | 0.42
| 220 | 11.995 | 12.0269 | 12.0509 | 0.016593 | 0.1993 | 0.42 | 225 | 0.044306 | 0.42
| 235 | 11.995 | 12.0376 | 12.0056 | 0.015022 | -0.26569 | 0.42 | 240 | -0.3605 | 0.42
| 250 | 11.995 | 12.0518 | 11.9855 | 0.019695 | -0.55011 | 0.42 | 255 | -0.75768 | 0.42
| 265 | 11.995 | 12.0636 | 11.9221 | 0.01444 | -1.1728 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 12.1239 | 11.613 | 0.020026 | -4.2136 | 0.5 | 45 | -4.2136 | 0.5
| 60 |  | 12.1315 | 11.771 | 0.015526 | -2.9712 | 0.5 | 60 | -2.9712 | 0.5
| 75 |  | 12.1242 | 11.905 | 0.019331 | -1.8084 | 0.5 | 75 | -1.8084 | 0.5
| 90 |  | 12.1285 | 12.078 | 0.015079 | -0.41596 | 0.5 | 90 | -0.41596 | 0.5
| 105 |  | 12.1221 | 12.168 | 0.010563 | 0.37823 | 0.5 | 105 | 0.37823 | 0.5
| 120 |  | 12.1225 | 12.2235 | 0.013485 | 0.83358 | 0.5 | 120 | 0.83358 | 0.5
| 135 |  | 12.1112 | 12.271 | 0.012524 | 1.3194 | 0.5 | 135 | 1.3194 | 0.5
| 150 |  | 12.1082 | 12.296 | 0.010954 | 1.551 | 0.5 | 150 | 1.551 | 0.5
| 165 |  | 12.1099 | 12.338 | 0.013992 | 1.8836 | 0.5 | 165 | 1.8836 | 0.5
| 180 |  | 12.1115 | 12.356 | 0.014654 | 2.0183 | 0.5 | 180 | 2.0183 | 0.5
| 195 |  | 12.1153 | 12.2815 | 0.019808 | 1.3714 | 0.5 | 195 | 1.3714 | 0.5
| 210 |  | 12.1052 | 12.179 | 0.028079 | 0.60924 | 0.5 | 210 | 0.60924 | 0.5
| 225 | 12.1096 | 12.1096 | 12.0875 | 0.018317 | -0.18209 | 0.5 | 225 | -0.18209 | 0.5
| 240 |  | 12.1192 | 12.0445 | 0.01905 | -0.61638 | 0.5 | 240 | -0.61638 | 0.5
| 255 |  | 12.1205 | 12.0025 | 0.015174 | -0.97356 | 0.5 | 255 | -0.97356 | 0.5
| 270 |  | 12.11 | 11.9575 | 0.014824 | -1.2593 | 0.5 | 270 | -1.2593 | 0.5
| 285 |  | 12.1122 | 11.902 | 0.013992 | -1.7354 | 0.5 | 285 | -1.7354 | 0.5
| 300 |  | 12.1217 | 11.887 | 0.01689 | -1.9362 | 0.5 | 300 | -1.9362 | 0.5
| 315 |  | 12.1215 | 11.8685 | 0.016944 | -2.0872 | 0.5 | 315 | -2.0872 | 0.5
| 330 |  | 12.1194 | 11.8725 | 0.015517 | -2.0372 | 0.5 | 330 | -2.0372 | 0.5
| 345 |  | 12.1202 | 11.8705 | 0.018771 | -2.0602 | 0.5 | 345 | -2.0602 | 0.5
| 360 |  | 12.1286 | 11.8715 | 0.018715 | -2.1198 | 0.5 | 360 | -2.1198 | 0.5
| 375 |  | 12.1315 | 11.8665 | 0.018994 | -2.184 | 0.5 | 375 | -2.184 | 0.5
| 390 |  | 12.1299 | 11.849 | 0.01944 | -2.3154 | 0.5 | 390 | -2.3154 | 0.5
| 405 |  | 12.1345 | 11.822 | 0.022148 | -2.5749 | 0.5 | 405 | -2.5749 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 11.993 | 12.07 | 0.02 | 0.64176 | 0.58315 | 120 | 0.64176 | 0.58315
| 135 |  | 11.9904 | 12.14 | 0.02 | 1.2474 | 0.58676 | 135 | 1.2474 | 0.58676
| 150 |  | 11.9912 | 12.2 | 0.03 | 1.7411 | 0.59202 | 150 | 1.7411 | 0.59202
| 165 |  | 11.996 | 12.25 | 0.02 | 2.1173 | 0.5915 | 165 | 2.1173 | 0.5915
| 180 |  | 11.996 | 12.31 | 0.02 | 2.6178 | 0.59438 | 180 | 2.6178 | 0.59438
| 195 |  | 11.9984 | 12.3 | 0.04 | 2.5137 | 0.59948 | 195 | 2.5137 | 0.59948
| 210 |  | 12.0154 | 12.2 | 0.05 | 1.5366 | 0.59744 | 210 | 1.5366 | 0.59744
| 225 |  | 12.0161 | 12.12 | 0.03 | 0.86507 | 0.58578 | 225 | 0.86507 | 0.58578
| 240 |  | 12.0187 | 12.08 | 0.03 | 0.51036 | 0.58363 | 240 | 0.51036 | 0.58363
| 255 |  | 12.0088 | 12.01 | 0.04 | 0.010214 | 0.58472 | 255 | 0.010214 | 0.58472
| 270 | 12.004 | 12.005 | 11.9733 | 0.046667 | -0.26401 | 0.5862 | 270 | -0.26401 | 0.5862
| 285 |  | 12.0057 | 11.93 | 0.06 | -0.63058 | 0.59108 | 285 | -0.63058 | 0.59108
| 300 |  | 11.9933 | 11.89 | 0.06 | -0.86166 | 0.5904 | 300 | -0.86166 | 0.5904
| 315 |  | 11.9885 | 11.87 | 0.04 | -0.9884 | 0.58002 | 315 | -0.9884 | 0.58002
| 330 |  | 11.993 | 11.88 | 0.03 | -0.94211 | 0.57656 | 330 | -0.94211 | 0.57656
| 345 |  | 11.9984 | 11.85 | 0.04 | -1.2368 | 0.57812 | 345 | -1.2368 | 0.57812
| 360 |  | 11.996 | 11.85 | 0.04 | -1.2168 | 0.57836 | 360 | -1.2168 | 0.57836
| 375 |  | 11.996 | 11.84 | 0.04 | -1.3005 | 0.57788 | 375 | -1.3005 | 0.57788
| 390 |  | 11.9912 | 11.85 | 0.06 | -1.1777 | 0.58874 | 390 | -1.1777 | 0.58874
| 405 |  | 11.9904 | 11.86 | 0.05 | -1.0878 | 0.58384 | 405 | -1.0878 | 0.58384
| 420 |  | 11.993 | 11.85 | 0.06 | -1.1926 | 0.58856 | 420 | -1.1926 | 0.58856
435
450
465
480
495
510
525
45
DTI | 60 |  | 12.0306 | 10.967 | 0.039216 | -8.8406 | 0.65359 | 60 | -8.8406 | 0.65359
| 75 |  | 12.0368 | 11.503 | 0.036577 | -4.4351 | 0.61089 | 75 | -4.4351 | 0.61089
| 90 |  | 12.0311 | 11.9575 | 0.031768 | -0.61176 | 0.56777 | 90 | -0.61176 | 0.56777
| 105 |  | 12.0315 | 12.1315 | 0.016631 | 0.83103 | 0.51323 | 105 | 0.83103 | 0.51323
| 120 |  | 12.0399 | 12.1985 | 0.023458 | 1.3172 | 0.52812 | 120 | 1.3172 | 0.52812
| 135 |  | 12.0291 | 12.2085 | 0.015985 | 1.4914 | 0.5086 | 135 | 1.4914 | 0.5086
| 150 |  | 12.0198 | 12.226 | 0.023486 | 1.7157 | 0.52701 | 150 | 1.7157 | 0.52701
| 165 |  | 12.0371 | 12.247 | 0.010809 | 1.7436 | 0.4978 | 165 | 1.7436 | 0.4978
| 180 |  | 12.039 | 12.285 | 0.021151 | 2.0434 | 0.51786 | 180 | 2.0434 | 0.51786
| 195 |  | 12.0344 | 12.295 | 0.017014 | 2.165 | 0.50724 | 195 | 2.165 | 0.50724
| 210 |  | 12.0315 | 12.26 | 0.024921 | 1.8993 | 0.52993 | 210 | 1.8993 | 0.52993
| 225 |  | 12.0418 | 12.161 | 0.027511 | 0.99001 | 0.54277 | 225 | 0.99001 | 0.54277
| 240 |  | 12.0357 | 12.0805 | 0.015381 | 0.37257 | 0.51273 | 240 | 0.37257 | 0.51273
| 255 | 12.0075 | 12.0295 | 12.0195 | 0.023278 | -0.083337 | 0.53544 | 255 | -0.083337 | 0.53544
| 270 |  | 12.0357 | 11.9825 | 0.023368 | -0.44167 | 0.53737 | 270 | -0.44167 | 0.53737
| 285 |  | 12.0418 | 11.9375 | 0.016819 | -0.86603 | 0.52199 | 285 | -0.86603 | 0.52199
| 300 |  | 12.0315 | 11.8955 | 0.018489 | -1.1303 | 0.5278 | 300 | -1.1303 | 0.5278
| 315 |  | 12.0344 | 11.909 | 0.023598 | -1.0424 | 0.54139 | 315 | -1.0424 | 0.54139
| 330 |  | 12.039 | 11.91 | 0.016859 | -1.0715 | 0.52329 | 330 | -1.0715 | 0.52329
| 345 |  | 12.0371 | 11.8645 | 0.015035 | -1.4341 | 0.52135 | 345 | -1.4341 | 0.52135
| 360 |  | 12.0198 | 11.8375 | 0.025726 | -1.5164 | 0.55149 | 360 | -1.5164 | 0.55149
| 375 |  | 12.0291 | 11.828 | 0.017652 | -1.6718 | 0.52877 | 375 | -1.6718 | 0.52877
| 390 |  | 12.0138 | 11.8555 | 0.021392 | -1.3175 | 0.5373 | 390 | -1.3175 | 0.5373
| 405 |  | 12.024 | 11.8225 | 0.024895 | -1.6754 | 0.54946 | 405 | -1.6754 | 0.54946
| 420 |  | 12.0229 | 11.835 | 0.023952 | -1.5624 | 0.54587 | 420 | -1.5624 | 0.54587
| 435 |  | 12.0115 | 11.824 | 0.026036 | -1.5612 | 0.55316 | 435 | -1.5612 | 0.55316
| 450 |  | 12.0236 | 11.769 | 0.028636 | -2.1176 | 0.5649 | 450 | -2.1176 | 0.5649
465
480
495
510
525
45
DWG | 50 |  | 11.9613 | 11.262 | 0.10231 | -5.8466 | 0.56107 | 60 | -3.9553 | 0.55656
| 65 |  | 12.0923 | 11.7284 | 0.06223 | -3.0096 | 0.55431 | 75 | -1.6403 | 0.55395
| 80 |  | 12.0993 | 11.9837 | 0.034944 | -0.9557 | 0.55377 | 90 | -0.083164 | 0.55425
| 95 |  | 12.0833 | 12.126 | 0.016704 | 0.3531 | 0.55449 | 105 | 0.83068 | 0.55496
| 110 |  | 12.0683 | 12.1974 | 0.027842 | 1.0695 | 0.5552 | 120 | 1.2977 | 0.55546
| 125 |  | 12.0603 | 12.2306 | 0.015985 | 1.4118 | 0.55559 | 135 | 1.5593 | 0.55558
| 140 |  | 12.0613 | 12.2583 | 0.018399 | 1.633 | 0.55557 | 150 | 1.6597 | 0.55578
| 155 |  | 12.0543 | 12.256 | 0.018619 | 1.673 | 0.55589 | 165 | 1.5848 | 0.55579
| 170 |  | 12.0573 | 12.2431 | 0.021049 | 1.5407 | 0.55574 | 180 | 1.4101 | 0.5557
| 185 |  | 12.0583 | 12.2205 | 0.017235 | 1.3449 | 0.55568 | 195 | 1.1218 | 0.55585
| 200 |  | 12.0523 | 12.1741 | 0.01626 | 1.0103 | 0.55594 | 210 | 0.72821 | 0.55592
| 215 |  | 12.0523 | 12.1231 | 0.021334 | 0.58716 | 0.55592 | 225 | 0.27464 | 0.55594
| 230 |  | 12.0513 | 12.0656 | 0.011495 | 0.11838 | 0.55596 | 240 | -0.083552 | 0.55602
| 245 |  | 12.0493 | 12.0271 | 0.017885 | -0.18452 | 0.55605 | 255 | -0.46061 | 0.55605
| 260 |  | 12.0493 | 11.9772 | 0.015061 | -0.59865 | 0.55606 | 270 | -0.80014 | 0.55613
| 275 |  | 12.0473 | 11.9388 | 0.01935 | -0.90089 | 0.55616 | 285 | -1.049 | 0.55615
| 290 |  | 12.0481 | 11.9128 | 0.020957 | -1.1231 | 0.55614 | 300 | -1.2689 | 0.5561
| 305 |  | 12.0498 | 11.8881 | 0.014828 | -1.3417 | 0.55608 | 315 | -1.3987 | 0.55608
| 320 |  | 12.0498 | 11.8778 | 0.014419 | -1.4272 | 0.55608 | 330 | -1.4856 | 0.55607
| 335 |  | 12.0504 | 11.8679 | 0.017099 | -1.5148 | 0.55606 | 345 | -1.5704 | 0.55597
| 350 |  | 12.0534 | 11.8608 | 0.020562 | -1.5983 | 0.55593 | 360 | -1.6713 | 0.55593
| 365 |  | 12.0534 | 11.8476 | 0.01729 | -1.7078 | 0.55594 | 375 | -1.7296 | 0.55592
| 380 |  | 12.0541 | 11.8443 | 0.021656 | -1.7406 | 0.55591 | 390 | -1.7205 | 0.55592
| 395 |  | 12.0538 | 11.8476 | 0.015813 | -1.7105 | 0.55592 | 405 | -1.8117 | 0.55593
| 410 |  | 12.0538 | 11.8293 | 0.018169 | -1.8623 | 0.55594 | 420 | -1.8293 | 0.55592
| 425 |  | 12.0541 | 11.8356 | 0.020859 | -1.8128 | 0.55592 | 435 | -1.855 | 0.55598
| 440 |  | 12.0521 | 11.826 | 0.015268 | -1.8761 | 0.55601 | 450 | -1.8315 | 0.556
| 455 |  | 12.0524 | 11.8344 | 0.020875 | -1.8091 | 0.55599 | 465 | -1.8452 | 0.55599
| 470 |  | 12.0528 | 11.8282 | 0.02019 | -1.8633 | 0.55598 | 480 | -1.9113 | 0.556
| 485 |  | 12.0524 | 11.8192 | 0.014438 | -1.9352 | 0.556 | 495 | -1.8866 | 0.55608
| 500 | 12.05 | 12.05 | 11.8256 | 0.018922 | -1.8622 | 0.55611 | 510 | -1.9021 | 0.55604
| 515 |  | 12.0524 | 11.8208 | 0.021364 | -1.922 | 0.556 | 525 | -1.8924 | 0.55602
| 530 |  | 12.0518 | 11.8255 | 0.021611 | -1.8775 | 0.55603
UL-LMPS | 42.5 | 12.3684 | 12.4039 | 11.7203 | 0.017434 | -5.5112 | 0.6965 | 45 | -5.3197 | 0.69642
| 57.5 | 12.2452 | 12.2392 | 11.7053 | 0.015381 | -4.3618 | 0.69605 | 60 | -4.2282 | 0.69611
| 72.5 | 12.135 | 12.1326 | 11.7007 | 0.016944 | -3.56 | 0.69639 | 75 | -3.3093 | 0.69663
| 87.5 | 12.0391 | 12.0279 | 11.7807 | 0.022618 | -2.0556 | 0.69787 | 90 | -1.8546 | 0.69759
| 102.5 | 11.9329 | 11.9408 | 11.8393 | 0.016051 | -0.84977 | 0.69619 | 105 | -0.76687 | 0.69626
| 117.5 | 11.8332 | 11.8544 | 11.8127 | 0.017852 | -0.35235 | 0.6966 | 120 | -0.33672 | 0.69671
| 132.5 | 11.7563 | 11.7968 | 11.7663 | 0.02059 | -0.25859 | 0.69729 | 135 | -0.13651 | 0.7305
| 147.5 | 11.6515 | 11.6863 | 11.7417 | 0.016694 | 0.47392 | 0.89653 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | 68 |  | 12.18 | 11.896 | 0.013 | -2.332 | 0.52 | 75 | -1.7342 | 0.52
| 83 |  | 12.174 | 12.046 | 0.013 | -1.051 | 0.52 | 90 | -0.51853 | 0.52
| 98 |  | 12.178 | 12.189 | 0.015 | 0.09 | 0.52 | 105 | 0.4316 | 0.52
| 113 |  | 12.167 | 12.267 | 0.012 | 0.822 | 0.52 | 120 | 1.0446 | 0.52
| 128 |  | 12.167 | 12.325 | 0.013 | 1.299 | 0.52 | 135 | 1.4371 | 0.52
| 143 |  | 12.165 | 12.359 | 0.007 | 1.595 | 0.52 | 150 | 1.8134 | 0.52
| 158 |  | 12.166 | 12.417 | 0.014 | 2.063 | 0.52 | 165 | 2.5012 | 0.52
| 173 |  | 12.158 | 12.523 | 0.457 | 3.002 | 0.52 | 180 | 2.4112 | 0.52
| 188 |  | 12.153 | 12.364 | 0.023 | 1.736 | 0.52 | 195 | 1.218 | 0.52
| 203 |  | 12.147 | 12.223 | 0.02 | 0.626 | 0.52 | 210 | 0.28393 | 0.52
| 218 |  | 12.148 | 12.135 | 0.017 | -0.107 | 0.52 | 225 | -0.303 | 0.52
| 233 |  | 12.136 | 12.072 | 0.019 | -0.527 | 0.52 | 240 | -0.83453 | 0.52
| 248 |  | 12.142 | 11.998 | 0.015 | -1.186 | 0.52 | 255 | -1.3279 | 0.52
| 263 |  | 12.147 | 11.966 | 0.017 | -1.49 | 0.52 | 270 | -1.5791 | 0.52
| 278 |  | 12.138 | 11.934 | 0.02 | -1.681 | 0.52 | 285 | -1.7389 | 0.52
| 293 |  | 12.131 | 11.912 | 0.015 | -1.805 | 0.52 | 300 | -1.8974 | 0.52
| 308 |  | 12.133 | 11.89 | 0.015 | -2.003 | 0.52 | 315 | -2.0492 | 0.52
| 323 |  | 12.131 | 11.876 | 0.013 | -2.102 | 0.52 | 330 | -2.2177 | 0.52
| 338 |  | 12.13 | 11.845 | 0.027 | -2.35 | 0.52 | 345 | -3.0043 | 0.52
| 353 |  | 12.127 | 11.672 | 0.018 | -3.752 | 0.52 | 360
| 368 |  | 12.123 | 11.158 | 0.066 | -7.96 | 0.52 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-S 20
| 20 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.44907
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 19.8674 | 19.439 | 0.0073786 | -2.1565 | 0.62917 | 45 | -1.4326 | 0.63163
| 52.5 |  | 19.8708 | 19.73 | 0.012472 | -0.70858 | 0.6341 | 60 | 0.0017742 | 0.63678 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 19.8606 | 20.002 | 0.018738 | 0.71213 | 0.63946 | 75 | 1.2878 | 0.63914 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 19.8805 | 20.251 | 0.012867 | 1.8635 | 0.63882 | 90 | 2.1354 | 0.63888 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 19.8765 | 20.355 | 0.0052705 | 2.4074 | 0.63895 | 105 | 2.4905 | 0.63884 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 19.8833 | 20.395 | 0.0084984 | 2.5736 | 0.63873 | 120 | 2.7024 | 0.63882 |  | 1.0407 | 5.1545 | 5.2042 | 2.9475 | 4.4418 | 4.8279 | 4.8706 | 2.3094 | 4.001 | 0.22155 | 1.7563 | 0.82312 | 1.9022 | 1.0163 | 1.0538 | 0.358 | 0.73227 | 0.56631 | 4.8654 | 4.481 | 0.15463 | 0.36379 | 1.5745 | 0.6383 | 4.5632 | 4.1348 | 0.52624 | 0.7244 | 1.2617 | 0.27096 | 2.926 | 2.5276 | 0.50041 | 0.3974 | 1.3425 | 0.85787
| 127.5 | 19.8733 | 19.8772 | 20.44 | 0.0066667 | 2.8311 | 0.63892 | 135 | 2.789 | 0.63871 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 19.8906 | 20.437 | 0.0067495 | 2.7469 | 0.63849 | 150 | 2.8524 | 0.63908 | 45 |  | 1.7701 |  | 2.029 | [6 empty] | 0.49565
| 157.5 |  | 19.8537 | 20.441 | 0.0073786 | 2.9579 | 0.63968 | 165 | 2.5351 | 0.64186 | 60 | 0.12279 | 1.6984 | 3.6527 | 2.4477 |  | 1.8338 | 4.0638 | 2.6236 |  | 2.2027 | 1.0167 |  | 0.93834 | [11 empty] | 11.0996 | 12.429 | 10.3661 | 8.0425 | 8.2513
| 172.5 |  | 19.8742 | 20.294 | 0.030984 | 2.1123 | 0.64405 | 180 | 1.2994 | 0.63892 | 75 | 0.36577 | 1.6218 | 2.5489 | 2.5408 | 1.7329 | 1.4765 | 2.541 | 2.4989 | 1.6066 | 1.0703 | 1.1926 | 0.11274 | 0.22695 | 0.96855 | 1.1015 |  | 5.4631 | [7 empty] | 10.447 | 11.4924 | 9.7491 | 8.4825 | 7.5039 | 9.7042 | 0.13406 | 0.3094 | 0.91978 | 1.3934 | 1.4742 | 0.96988
| 187.5 |  | 19.8803 | 19.977 | 0.020028 | 0.48656 | 0.63379 | 195 | -0.24789 | 0.63675 | 90 | 0.52089 | 1.2666 | 1.726 | 2.1493 | 1.507 | 0.89009 | 1.4265 | 1.9115 | 1.1685 | 0.54475 | 1.1091 | 0.25832 | 0.60414 | 0.29904 | 0.89163 |  | 4.0988 | [7 empty] | 8.7595 | 9.3934 | 8.2646 | 7.5353 | 6.3046 | 8.0795 | 0.40488 | 0.17332 | 0.19783 | 0.44075 | 0.73312 | 0.31007
| 202.5 |  | 19.853 | 19.658 | 0.034577 | -0.98234 | 0.6397 | 210 | -1.5137 | 0.63373 | 105 | 0.72798 | 1.2305 | 1.2386 | 1.7787 | 1.2789 | 0.61666 | 0.64312 | 1.2994 | 0.66962 | 0.047821 | 0.74703 | 0.04579 | 0.68577 | 0.0039658 | 0.71155 |  | 1.97 | [7 empty] | 6.3516 | 6.4921 | 5.6757 | 5.4573 | 4.2696 | 5.6682 | 1.221 | 0.91727 | 0.65308 | 0.62711 | 0.26697 | 0.63411
| 217.5 |  | 19.9122 | 19.505 | 0.022236 | -2.0451 | 0.62776 | 225 |  |  | 120 | 0.8112 | 1.1085 | 1.0789 | 1.7179 | 1.0451 | 0.38185 | 0.37003 | 1.1617 | 0.30356 | 0.0027735 | 0.81628 | 0.077375 | 0.79322 | 0.077409 | 0.88546 | 3.3961 | 0.016344 | 1.6115 | 1.0619 | 0.35765 | 0.0047247 | 0.0072267 | 0.79306 | 0.069005 | 4.2962 | 4.0706 | 3.5431 | 3.4249 | 2.2263 | 3.6407 | 2.1145 | 1.7989 | 1.6244 | 1.6118 | 1.1579 | 1.6601
[8 empty] | 240 |  |  | 135 | 0.91655 | 0.89538 | 1.1021 | 1.4371 | 0.88884 | 0.016517 | 0.28359 | 0.7882 | 0.0029854 | 0.25589 | 0.75269 | 0.013003 | 0.53134 | 0.27013 | 0.76625 | 2.5657 | 1.0153 | 2.2056 | 0.5953 | 0.28755 | 0.29017 | 0.52526 | 0.96262 | 0.27968 | 3.0469 | 2.5203 | 2.3921 | 2.0634 | 1.1681 | 2.4208 | 2.474 | 2.12 | 2.0944 | 1.9649 | 1.5892 | 2.1028
[8 empty] | 255 |  |  | 150 | 1.0724 | 1.0398 | 1.2991 |  | 0.92929 | 0.010249 | 0.34291 |  | 0.12546 | 0.31835 |  | 0.12924 |  | 0.4451 |  | 2.3878 | 1.4617 | 2.5571 | 0.31735 | 0.77059 | 0.74682 | 1.0268 |  | 0.62878 | 2.5966 | 1.8292 | 1.7363 | 1.3653 |  | 1.8748 | 2.6852 | 2.2646 | 2.2403 | 2.0818 |  | 2.2993
[8 empty] | 270 |  |  | 165 | 0.94215 | 0.76792 | 1.231 |  | 0.91721 | 0.16077 | 0.405 |  | 0.007032 | 0.53486 |  | 0.16092 |  | 0.38323 |  | 3.1846 | 1.0759 | 2.5519 | 0.24296 | 1.2773 | 1.077 | 1.5527 |  | 1.2368 | 2.7789 | 2.1878 | 2.2385 | 1.6417 |  | 2.0969 | 2.403 | 2.035 | 2.0837 | 1.8291 |  | 2.0167
[8 empty] | 285 |  |  | 180 | 0.65554 | 0.30246 | 0.48632 |  | 0.66006 | 0.38096 | 0.14453 |  | 0.030473 | 0.2119 |  | 0.39444 |  | 0.16703 |  | 2.8395 | 1.3098 | 2.6346 | 1.5814 | 2.5053 | 2.0532 | 2.1895 |  | 2.441 | 1.1025 | 0.56352 | 0.88938 | 0.65459 |  | 0.51407 | 1.8267 | 1.5752 | 1.7194 | 1.6123 |  | 1.5508
[8 empty] | 300 |  |  | 195 | 0.51378 | 0.31927 | 0.40967 |  | 0.4289 | 0.20261 | 1.0216 |  | 0.078905 | 0.79036 |  | 0.11926 |  | 0.9106 |  | 1.6781 | 1.4289 | 2.2271 | 2.1736 | 3.0255 | 2.7141 | 1.8765 |  | 2.8443 | 0.57206 | 1.1973 | 0.96238 | 0.17733 |  | 1.0819 | 1.1322 | 0.92605 | 1.002 | 1.3496 |  | 0.95164
[8 empty] | 315 |  |  | 210 | 0.096868 | 0.099038 | 1.0964 |  | 0.10819 | 0.22245 | 1.1534 |  | 0.23415 | 1.3069 |  | 0.0094691 |  | 1.3247 |  | 0.6026 | 1.2128 | 1.4846 | 1.8682 | 2.0332 | 2.1476 | 0.82824 |  | 2.1705 | 1.3358 | 1.4323 | 1.5746 | 0.23965 |  | 1.5946 | 0.55301 | 0.61445 | 0.51664 | 1.0995 |  | 0.51323
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.24484 | 0.8082 |  | 0.5314 | 0.99746 |  | 0.27091 |  | 1.2652 |  | 0.081932 | 1.0245 | 0.98667 |  | 1.0975 | 1.2733 | 0.27168 |  | 1.5415 |  | 1.183 | 1.3547 | 0.35326 |  | 1.6227 |  | 0.52252 | 0.41622 | 0.86097 |  | 0.29935
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.38397 | 0.57838 |  | 0.4177 | 0.90558 |  | 0.029228 |  | 0.93976 |  | 0.05521 | 0.84009 | 0.86664 |  | 1.0202 | 1.3265 | 0.41192 |  | 1.3632 |  | 0.95637 | 1.2645 | 0.35508 |  | 1.3007 |  | 0.43412 | 0.27083 | 0.67539 |  | 0.25853
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.34082 | 0.57875 |  | 0.25028 | 0.86629 |  | 0.088723 |  | 0.7865 |  | 0.18603 | 0.74419 | 0.82836 |  | 1.0795 | 1.3446 | 0.46369 |  | 1.2681 |  | 0.89007 | 1.166 | 0.28197 |  | 1.0875 |  | 0.3706 | 0.2258 | 0.61304 |  | 0.26456
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.71734 |  | 0.081918 |  | 0.64313 |  | 0.088947 | 0.52793 | 0.5682 |  |  | 0.84463 | 0.12258 |  | 0.77122 |  |  | 0.75841 | 0.034469 |  | 0.68395 |  |  | 0.1906 | 0.51125 |  | 0.22637
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.86611 |  | 0.46027 |  | 0.42826 |  | 0.49133 | 0.43329 | 0.65737 |  |  | 1.3541 | 0.47362 |  | 0.92043 |  |  | 0.86693 | 0.011346 |  | 0.42272 |  |  | 0.048785 | 0.43725 |  | 0.24836
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.72805 |  | 0.38286 |  | 0.36384 |  | 0.82331 | 0.2784 | 0.65575 |  |  | 1.3264 | 0.5769 |  | 0.96542 |  |  | 0.48744 | 0.24006 |  | 0.11765 |  |  | 0.061783 | 0.38815 |  | 0.22782
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.76749 |  | 0.23428 |  | 0.54675 |  | 0.53541 | 0.38849 | 0.63526 |  |  | 1.3779 | 0.59183 |  | 1.1616 |  |  | 0.83491 | 0.060218 |  | 0.61317 |  |  | 0.015809 | 0.36019 |  | 0.11738
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.59767 |  | 0.049814 |  | 0.64964 |  | 0.48296 | 0.41779 | 0.6428 |  |  | 1.1453 | 0.53048 |  | 1.2008 |  |  | 0.64445 | 0.046507 |  | 0.69661 |  |  | 0.12899 | 0.39634 |  | 0.10757
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.55818 |  | 0.68403 |  | 1.2213 |  | 0.33731 | 0.37926 | 0.53343 |  |  | 0.85371 | 0.28507 |  | 1.5196 |  |  | 0.51311 | 0.049327 |  | 1.1832 |  |  | 0.15167 | 0.40112 |  | 0.14463
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.78851 | [5 empty] | 0.6264 | 0.30671 | 0.59428 |  |  | 1.1434 | 0.3414 |  |  |  |  | 0.50388 | 0.28165 |  |  |  |  | 0.082552 | 0.43587
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.053 | [5 empty] | 0.8345 |  |  |  |  | 1.3559 | 0.29048 |  |  |  |  | 0.51395 | 0.53814
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.92011 | [5 empty] | 0.55237 |  |  |  |  | 1.2481 | 0.32049 |  |  |  |  | 0.69364 | 0.2292
[8 empty] | 510 |  |  | 405 | [9 empty] | 1.3203 | [5 empty] | 0.72271 |  |  |  |  | 1.524 | 0.19375 |  |  |  |  | 0.77073 | 0.5282
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.61988 | [5 empty] | 0.19447 | [5 empty] | 0.42313
[11 empty] | 435 | [27 empty] | 0.33085
[11 empty] | 450 | [27 empty] | 0.78259
45
PTB | 55 | 19.894 | 19.9713 | 19.8757 | 0.012479 | -0.47845 | 0.43 | 60 | -0.1075 | 0.43
| 70 | 19.894 | 19.979 | 20.1057 | 0.021676 | 0.63439 | 0.43 | 75 | 0.96166 | 0.43
| 85 | 19.894 | 19.9393 | 20.2616 | 0.01599 | 1.6162 | 0.43 | 90 | 1.6711 | 0.43
| 100 | 19.894 | 19.9263 | 20.2811 | 0.0145 | 1.7808 | 0.43 | 105 | 1.8416 | 0.43
| 115 | 19.894 | 19.9171 | 20.3081 | 0.013272 | 1.963 | 0.43 | 120 | 1.9792 | 0.43
| 130 | 19.894 | 19.8989 | 20.2993 | 0.012687 | 2.0117 | 0.43 | 135 | 1.972 | 0.43
| 145 | 19.894 | 19.8935 | 20.27 | 0.01271 | 1.8927 | 0.43 | 150 | 1.8962 | 0.43
| 160 | 19.894 | 19.894 | 20.2727 | 0.013728 | 1.9033 | 0.43 | 165 | 1.6932 | 0.43
| 175 | 19.894 | 19.9019 | 20.1552 | 0.036413 | 1.2729 | 0.43 | 180 | 0.71499 | 0.43
| 190 | 19.894 | 19.9054 | 19.8256 | 0.047 | -0.40078 | 0.43 | 195 | -0.70513 | 0.43
| 205 | 19.894 | 19.9238 | 19.6621 | 0.03349 | -1.3138 | 0.43 | 210 | -1.4277 | 0.43
| 220 | 19.894 | 19.932 | 19.6021 | 0.023953 | -1.6555 | 0.43 | 225 | -1.7533 | 0.43
| 235 | 19.894 | 19.9549 | 19.566 | 0.021213 | -1.9489 | 0.43 | 240 | -2.0036 | 0.43
| 250 | 19.894 | 19.9699 | 19.5479 | 0.024668 | -2.113 | 0.43 | 255 | -2.2273 | 0.43
| 265 | 19.894 | 19.9889 | 19.498 | 0.018708 | -2.456 | 0.43 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 20.0949 | 19.479 | 0.028451 | -3.0651 | 0.5 | 45 | -3.0651 | 0.5
| 60 |  | 20.0839 | 19.7685 | 0.020333 | -1.5706 | 0.5 | 60 | -1.5706 | 0.5
| 75 |  | 20.0769 | 20.0335 | 0.015985 | -0.21637 | 0.5 | 75 | -0.21637 | 0.5
| 90 |  | 20.0756 | 20.2685 | 0.01927 | 0.96092 | 0.5 | 90 | 0.96092 | 0.5
| 105 |  | 20.0815 | 20.3525 | 0.021491 | 1.3496 | 0.5 | 105 | 1.3496 | 0.5
| 120 |  | 20.0773 | 20.4135 | 0.023232 | 1.6746 | 0.5 | 120 | 1.6746 | 0.5
| 135 |  | 20.0684 | 20.4615 | 0.023232 | 1.9589 | 0.5 | 135 | 1.9589 | 0.5
| 150 |  | 20.0637 | 20.4425 | 0.023592 | 1.888 | 0.5 | 150 | 1.888 | 0.5
| 165 |  | 20.0616 | 20.427 | 0.022965 | 1.8214 | 0.5 | 165 | 1.8214 | 0.5
| 180 |  | 20.066 | 20.2705 | 0.047181 | 1.0189 | 0.5 | 180 | 1.0189 | 0.5
| 195 |  | 20.0635 | 19.9545 | 0.040324 | -0.54347 | 0.5 | 195 | -0.54347 | 0.5
| 210 |  | 20.0686 | 19.7465 | 0.022775 | -1.6052 | 0.5 | 210 | -1.6052 | 0.5
| 225 | 20.0522 | 20.0522 | 19.6615 | 0.028704 | -1.9486 | 0.5 | 225 | -1.9486 | 0.5
| 240 |  | 20.067 | 19.6035 | 0.017852 | -2.31 | 0.5 | 240 | -2.31 | 0.5
| 255 |  | 20.0675 | 19.566 | 0.026238 | -2.4993 | 0.5 | 255 | -2.4993 | 0.5
| 270 |  | 20.0523 | 19.5415 | 0.018432 | -2.5475 | 0.5 | 270 | -2.5475 | 0.5
| 285 |  | 20.0707 | 19.5155 | 0.02481 | -2.7664 | 0.5 | 285 | -2.7664 | 0.5
| 300 |  | 20.0687 | 19.512 | 0.019894 | -2.7742 | 0.5 | 300 | -2.7742 | 0.5
| 315 |  | 20.0776 | 19.4995 | 0.023725 | -2.8795 | 0.5 | 315 | -2.8795 | 0.5
| 330 |  | 20.0713 | 19.505 | 0.017918 | -2.8216 | 0.5 | 330 | -2.8216 | 0.5
| 345 |  | 20.0642 | 19.5045 | 0.021392 | -2.7897 | 0.5 | 345 | -2.7897 | 0.5
| 360 |  | 20.0696 | 19.4835 | 0.020333 | -2.9205 | 0.5 | 360 | -2.9205 | 0.5
| 375 |  | 20.0861 | 19.461 | 0.026338 | -3.1123 | 0.5 | 375 | -3.1123 | 0.5
| 390 |  | 20.0495 | 19.4575 | 0.022913 | -2.9529 | 0.5 | 390 | -2.9529 | 0.5
| 405 |  | 20.0533 | 19.389 | 0.022688 | -3.3129 | 0.5 | 405 | -3.3129 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 19.9648 | 20.3 | 0.04 | 1.6787 | 0.56521 | 120 | 1.6787 | 0.56521
| 135 |  | 19.9679 | 20.41 | 0.04 | 2.2142 | 0.56807 | 135 | 2.2142 | 0.56807
| 150 |  | 19.9521 | 20.46 | 0.03 | 2.5457 | 0.56906 | 150 | 2.5457 | 0.56906
| 165 |  | 19.9668 | 20.52 | 0.03 | 2.7704 | 0.56987 | 165 | 2.7704 | 0.56987
| 180 |  | 19.9653 | 20.53 | 0.03 | 2.8286 | 0.57024 | 180 | 2.8286 | 0.57024
| 195 |  | 19.9692 | 20.34 | 0.08 | 1.8568 | 0.57486 | 195 | 1.8568 | 0.57486
| 210 |  | 19.9837 | 20.04 | 0.08 | 0.28167 | 0.56592 | 210 | 0.28167 | 0.56592
| 225 |  | 19.9873 | 19.82 | 0.07 | -0.83718 | 0.55697 | 225 | -0.83718 | 0.55697
| 240 |  | 20.0003 | 19.77 | 0.08 | -1.1515 | 0.55771 | 240 | -1.1515 | 0.55771
| 255 |  | 19.9756 | 19.71 | 0.06 | -1.3297 | 0.55216 | 255 | -1.3297 | 0.55216
| 270 | 19.978 | 19.9826 | 19.62 | 0.06 | -1.8144 | 0.54933 | 270 | -1.8144 | 0.54933
| 285 |  | 19.967 | 19.65 | 0.07 | -1.5875 | 0.55345 | 285 | -1.5875 | 0.55345
| 300 |  | 19.9542 | 19.63 | 0.04 | -1.6245 | 0.54733 | 300 | -1.6245 | 0.54733
| 315 |  | 19.9553 | 19.62 | 0.07 | -1.6801 | 0.55328 | 315 | -1.6801 | 0.55328
| 330 |  | 19.9542 | 19.59 | 0.07 | -1.8252 | 0.55252 | 330 | -1.8252 | 0.55252
| 345 |  | 19.9692 | 19.56 | 0.06 | -2.0492 | 0.54841 | 345 | -2.0492 | 0.54841
| 360 |  | 19.9653 | 19.58 | 0.05 | -1.9297 | 0.54707 | 360 | -1.9297 | 0.54707
| 375 |  | 19.9668 | 19.58 | 0.05 | -1.9374 | 0.54699 | 375 | -1.9374 | 0.54699
| 390 |  | 19.9521 | 19.58 | 0.08 | -1.8649 | 0.55519 | 390 | -1.8649 | 0.55519
| 405 |  | 19.9679 | 19.57 | 0.05 | -1.9926 | 0.54666 | 405 | -1.9926 | 0.54666
| 420 |  | 19.9648 | 19.55 | 0.06 | -2.0779 | 0.54837 | 420 | -2.0779 | 0.54837
435
450
465
480
495
510
525
45
DTI | 60 |  | 20.0451 | 17.8075 | 0.0532 | -11.1628 | 0.63608 | 60 | -11.1628 | 0.63608
| 75 |  | 20.0565 | 18.224 | 0.052955 | -9.1367 | 0.62092 | 75 | -9.1367 | 0.62092
| 90 |  | 20.0554 | 18.7115 | 0.065396 | -6.7008 | 0.63856 | 90 | -6.7008 | 0.63856
| 105 |  | 20.0567 | 19.306 | 0.056326 | -3.7431 | 0.59449 | 105 | -3.7431 | 0.59449
| 120 |  | 20.0471 | 19.7635 | 0.045338 | -1.4148 | 0.55556 | 120 | -1.4148 | 0.55556
| 135 |  | 20.0382 | 20.0175 | 0.041023 | -0.1034 | 0.53996 | 135 | -0.1034 | 0.53996
| 150 |  | 20.0173 | 20.098 | 0.035924 | 0.40332 | 0.52869 | 150 | 0.40332 | 0.52869
| 165 |  | 20.0259 | 20.005 | 0.03927 | -0.10425 | 0.53704 | 165 | -0.10425 | 0.53704
| 180 |  | 20.0467 | 20.0965 | 0.045454 | 0.24849 | 0.54659 | 180 | 0.24849 | 0.54659
| 195 |  | 20.0141 | 20.075 | 0.055298 | 0.30412 | 0.56922 | 195 | 0.30412 | 0.56922
| 210 |  | 20.0262 | 19.9745 | 0.034409 | -0.25806 | 0.52945 | 210 | -0.25806 | 0.52945
| 225 |  | 20.0177 | 19.865 | 0.049151 | -0.76269 | 0.56092 | 225 | -0.76269 | 0.56092
| 240 |  | 20.016 | 19.7755 | 0.048718 | -1.2018 | 0.56249 | 240 | -1.2018 | 0.56249
| 255 | 20.0065 | 20.0215 | 19.722 | 0.035333 | -1.4957 | 0.53777 | 255 | -1.4957 | 0.53777
| 270 |  | 20.016 | 19.637 | 0.035258 | -1.8937 | 0.53997 | 270 | -1.8937 | 0.53997
| 285 |  | 20.0177 | 19.6125 | 0.028814 | -2.0241 | 0.53062 | 285 | -2.0241 | 0.53062
| 300 |  | 20.0262 | 19.5545 | 0.031031 | -2.3553 | 0.53545 | 300 | -2.3553 | 0.53545
| 315 |  | 20.0141 | 19.582 | 0.035034 | -2.1591 | 0.54111 | 315 | -2.1591 | 0.54111
| 330 |  | 20.0467 | 19.5935 | 0.041836 | -2.2607 | 0.55324 | 330 | -2.2607 | 0.55324
| 345 |  | 20.0259 | 19.5555 | 0.031031 | -2.3488 | 0.53542 | 345 | -2.3488 | 0.53542
| 360 |  | 20.0173 | 19.5195 | 0.032196 | -2.4867 | 0.53821 | 360 | -2.4867 | 0.53821
| 375 |  | 20.0382 | 19.5025 | 0.023368 | -2.6735 | 0.52657 | 375 | -2.6735 | 0.52657
| 390 |  | 20.0128 | 19.541 | 0.029895 | -2.3574 | 0.53412 | 390 | -2.3574 | 0.53412
| 405 |  | 20.023 | 19.494 | 0.040575 | -2.6418 | 0.5536 | 405 | -2.6418 | 0.5536
| 420 |  | 20.0219 | 19.4955 | 0.030689 | -2.6289 | 0.53655 | 420 | -2.6289 | 0.53655
| 435 |  | 20.0105 | 19.502 | 0.027261 | -2.5414 | 0.53148 | 435 | -2.5414 | 0.53148
| 450 |  | 20.0226 | 19.4545 | 0.02305 | -2.8374 | 0.5275 | 450 | -2.8374 | 0.5275
465
480
495
510
525
45
DWG | 50 |  | 19.7123 | 18.6739 | 0.11085 | -5.2679 | 0.5588 | 60 | -3.4864 | 0.55206
| 65 |  | 20.0543 | 19.5338 | 0.064089 | -2.5956 | 0.54869 | 75 | -1.1446 | 0.54819
| 80 |  | 20.0753 | 19.9912 | 0.037436 | -0.41909 | 0.54794 | 90 | 0.48837 | 0.54846
| 95 |  | 20.0473 | 20.2362 | 0.030288 | 0.9421 | 0.54873 | 105 | 1.308 | 0.54917
| 110 |  | 20.0243 | 20.3229 | 0.028024 | 1.491 | 0.54939 | 120 | 1.6722 | 0.54963
| 125 |  | 20.0123 | 20.3651 | 0.020842 | 1.7627 | 0.54975 | 135 | 1.7367 | 0.54962
| 140 |  | 20.0193 | 20.3644 | 0.024625 | 1.7237 | 0.54955 | 150 | 1.6116 | 0.54961
| 155 |  | 20.0153 | 20.3267 | 0.027478 | 1.5556 | 0.54965 | 165 | 1.357 | 0.54974
| 170 |  | 20.0093 | 20.261 | 0.027041 | 1.2577 | 0.54979 | 180 | 0.83494 | 0.54977
| 185 |  | 20.0093 | 20.1341 | 0.03926 | 0.62354 | 0.54975 | 195 | 0.14277 | 0.54975
| 200 |  | 20.0093 | 19.9898 | 0.043604 | -0.097621 | 0.54974 | 210 | -0.47031 | 0.54992
| 215 |  | 20.0003 | 19.869 | 0.032786 | -0.65666 | 0.55 | 225 | -1.0824 | 0.55004
| 230 |  | 19.9993 | 19.7403 | 0.031484 | -1.2952 | 0.55006 | 240 | -1.5234 | 0.55016
| 245 |  | 19.9953 | 19.6679 | 0.022779 | -1.6375 | 0.5502 | 255 | -1.7468 | 0.55019
| 260 |  | 19.9963 | 19.6361 | 0.028867 | -1.8015 | 0.55019 | 270 | -1.9245 | 0.55028
| 275 |  | 19.9923 | 19.5953 | 0.021349 | -1.9859 | 0.55032 | 285 | -2.014 | 0.55045
| 290 |  | 19.9854 | 19.5801 | 0.01826 | -2.0281 | 0.55051 | 300 | -2.1418 | 0.55027
| 305 |  | 19.9994 | 19.5597 | 0.027532 | -2.1986 | 0.55015 | 315 | -2.2129 | 0.55021
| 320 |  | 19.9959 | 19.552 | 0.025321 | -2.22 | 0.55025 | 330 | -2.3025 | 0.55021
| 335 |  | 19.9984 | 19.5297 | 0.021378 | -2.3438 | 0.55019 | 345 | -2.305 | 0.55008
| 350 |  | 20.0044 | 19.5472 | 0.025538 | -2.2856 | 0.55002 | 360 | -2.2357 | 0.55005
| 365 |  | 20.0024 | 19.5602 | 0.033075 | -2.2108 | 0.55007 | 375 | -2.1977 | 0.55013
| 380 |  | 19.9989 | 19.5607 | 0.022219 | -2.1912 | 0.55016 | 390 | -2.1538 | 0.55013
| 395 |  | 20.0004 | 19.5734 | 0.017738 | -2.135 | 0.55011 | 405 | -2.1662 | 0.55007
| 410 |  | 20.0029 | 19.5665 | 0.023715 | -2.1818 | 0.55005 | 420 | -2.2523 | 0.55004
| 425 |  | 20.0039 | 19.5463 | 0.021028 | -2.2876 | 0.55004 | 435 | -2.2479 | 0.5501
| 440 |  | 20.0004 | 19.5548 | 0.024819 | -2.228 | 0.55013 | 450 | -2.1451 | 0.55021
| 455 |  | 19.9954 | 19.5748 | 0.025191 | -2.1036 | 0.55025 | 465 | -2.2095 | 0.55017
| 470 |  | 20.0004 | 19.5479 | 0.025382 | -2.2625 | 0.55013 | 480 | -2.2606 | 0.55015
| 485 |  | 19.9994 | 19.5475 | 0.019328 | -2.2596 | 0.55016 | 495 | -2.1806 | 0.55032
| 500 | 19.99 | 19.99 | 19.562 | 0.019493 | -2.1411 | 0.5504 | 510 | -2.14 | 0.55027
| 515 |  | 19.9969 | 19.5691 | 0.021337 | -2.1394 | 0.55021 | 525 | -2.1678 | 0.55021
| 530 |  | 19.9969 | 19.5606 | 0.019356 | -2.1819 | 0.55022
UL-LMPS | 42.5 | 20.668 | 20.6121 | 19.8487 | 0.023154 | -3.704 | 0.6947 | 45 | -3.5442 | 0.69464
| 57.5 | 20.4643 | 20.3651 | 19.806 | 0.019582 | -2.7451 | 0.69434 | 60 | -2.5533 | 0.69459
| 72.5 | 20.2594 | 20.1383 | 19.8173 | 0.032049 | -1.5939 | 0.69588 | 75 | -1.37 | 0.69575
| 87.5 | 20.0628 | 19.9386 | 19.8887 | 0.026618 | -0.25049 | 0.69511 | 90 | -0.11135 | 0.69496
| 102.5 | 19.8626 | 19.744 | 19.8593 | 0.017991 | 0.58432 | 0.69419 | 105 | 0.62419 | 0.70083
| 117.5 | 19.6962 | 19.6264 | 19.788 | 0.022499 | 0.82354 | 0.73403 | 120 | 0.8631 | 0.73253
| 132.5 | 19.5439 | 19.5087 | 19.7157 | 0.021121 | 1.0609 | 0.72503 | 135 | 1.1596 | 0.82208
| 147.5 | 19.3403 | 19.2783 | 19.597 | 0.024233 | 1.653 | 1.3073 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | 64 |  | 20.3899 | 20.685 | 0.048504 | 1.4472 | 1.8 | 75 | 1.5509 | 1.8
| 79 |  | 20.4073 | 20.7315 | 0.044042 | 1.5886 | 1.8 | 90 | 1.341 | 1.8
| 94 |  | 20.4052 | 20.6605 | 0.043827 | 1.251 | 1.8 | 105 | 0.094748 | 1.8
| 109 |  | 20.3804 | 20.314 | 0.17142 | -0.32571 | 1.8 | 120 | -1.4465 | 1.8
| 124 |  | 20.3508 | 19.9735 | 0.051939 | -1.854 | 1.8 | 135 | -2.0652 | 1.8
| 139 |  | 20.3366 | 19.901 | 0.037683 | -2.142 | 1.8 | 150 | -2.4163 | 1.8
| 154 |  | 20.3187 | 19.8075 | 0.043392 | -2.5161 | 1.8 | 165 | -2.1821 | 1.8
| 169 |  | 20.3085 | 19.89 | 0.045073 | -2.0607 | 1.8 | 180 | -2.2847 | 1.8
| 184 |  | 20.3003 | 19.82 | 0.035244 | -2.3662 | 1.8 | 195 | -2.4687 | 1.8
| 199 |  | 20.31 | 19.801 | 0.042661 | -2.5059 | 1.8 | 210 | -2.5978 | 1.8
| 214 |  | 20.3202 | 19.7855 | 0.035759 | -2.6313 | 1.8 | 225 | -2.7483 | 1.8
| 229 |  | 20.3592 | 19.791 | 0.038237 | -2.7909 | 1.8 | 240 | -2.8303 | 1.8
| 244 |  | 20.3782 | 19.7985 | 0.042087 | -2.8447 | 1.8 | 255 | -2.9331 | 1.8
| 259 |  | 20.3906 | 19.786 | 0.042227 | -2.9653 | 1.8 | 270 | -2.9137 | 1.8
| 274 |  | 20.4073 | 19.8165 | 0.04171 | -2.895 | 1.8 | 285 | -2.8601 | 1.8
| 289 |  | 20.3839 | 19.8035 | 0.036314 | -2.8475 | 1.8 | 300 | -2.8929 | 1.8
| 304 |  | 20.3593 | 19.767 | 0.047473 | -2.9094 | 1.8 | 315 | -2.9099 | 1.8
| 319 | 20.3209 | 20.3209 | 19.7295 | 0.046957 | -2.9101 | 1.8 | 330 | -3.0695 | 1.8
| 334 |  | 20.2658 | 19.632 | 0.036649 | -3.1274 | 1.8 | 345 | -3.0811 | 1.8
| 349 |  | 20.2201 | 19.6005 | 0.037343 | -3.0643 | 1.8 | 360 | -3.0791 | 1.8
| 364 |  | 20.2702 | 19.645 | 0.036491 | -3.0845 | 1.8 | 375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | 68 |  | 20.233 | 20.034 | 0.022 | -0.984 | 0.49 | 75 | -0.31013 | 0.49
| 83 |  | 20.223 | 20.316 | 0.025 | 0.46 | 0.49 | 90 | 0.74607 | 0.49
| 98 |  | 20.216 | 20.433 | 0.02 | 1.073 | 0.49 | 105 | 1.3115 | 0.49
| 113 |  | 20.197 | 20.517 | 0.011 | 1.584 | 0.49 | 120 | 1.7389 | 0.49
| 128 |  | 20.2 | 20.587 | 0.012 | 1.916 | 0.49 | 135 | 1.9697 | 0.49
| 143 |  | 20.187 | 20.597 | 0.011 | 2.031 | 0.49 | 150 | 1.9955 | 0.49
| 158 |  | 20.201 | 20.596 | 0.025 | 1.955 | 0.49 | 165 | 1.6876 | 0.49
| 173 |  | 20.184 | 20.463 | 0.042 | 1.382 | 0.49 | 180 | 0.69087 | 0.49
| 188 |  | 20.192 | 20.172 | 0.052 | -0.099 | 0.49 | 195 | -0.64267 | 0.49
| 203 |  | 20.17 | 19.915 | 0.033 | -1.264 | 0.49 | 210 | -1.6131 | 0.49
| 218 |  | 20.174 | 19.768 | 0.025 | -2.012 | 0.49 | 225 | -2.1739 | 0.49
| 233 |  | 20.219 | 19.742 | 0.02 | -2.359 | 0.49 | 240 | -2.3343 | 0.49
| 248 |  | 20.166 | 19.701 | 0.019 | -2.306 | 0.49 | 255 | -2.4255 | 0.49
| 263 |  | 20.18 | 19.663 | 0.032 | -2.562 | 0.49 | 270 | -2.4794 | 0.49
| 278 |  | 20.167 | 19.686 | 0.016 | -2.385 | 0.49 | 285 | -2.3836 | 0.49
| 293 |  | 20.154 | 19.674 | 0.018 | -2.382 | 0.49 | 300 | -2.4557 | 0.49
| 308 |  | 20.157 | 19.645 | 0.02 | -2.54 | 0.49 | 315 | -2.6847 | 0.49
| 323 |  | 20.213 | 19.637 | 0.015 | -2.85 | 0.49 | 330 | -2.8631 | 0.49
| 338 |  | 20.152 | 19.572 | 0.02 | -2.878 | 0.49 | 345 | -3.3587 | 0.49
| 353 |  | 20.139 | 19.352 | 0.02 | -3.908 | 0.49 | 360
| 368 |  | 20.13 | 18.298 | 0.089 | -9.101 | 0.49 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 0.5
| 0.5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 16.3544
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 0.50715 | 0.458 | 0.010328 | -9.6912 | 2.169 | 45
| 52.5 |  | 0.50642 | 0.464 | 0.0096609 | -8.3757 | 2.1721 | 60 | -7.9761 | 2.1722 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 0.50636 | 0.468 | 0.011353 | -7.5765 | 2.1723 | 75 | -6.3501 | 2.1733 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 0.50592 | 0.48 | 0.010541 | -5.1237 | 2.1742 | 90 | -5.3524 | 2.1735 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 0.50625 | 0.478 | 0.010328 | -5.5812 | 2.1728 | 105 | -5.9957 | 2.0736 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 0.50647 | 0.474 | 0.0084327 | -6.4103 | 1.9745 | 120 | -6.2377 | 2.0726 |  | 0.096974 | 0.17999 |  |  |  | 0.27845 | [12 empty] | 0.12803 | 0.2259 | 0.051678
| 127.5 | 0.50684 | 0.50673 | 0.476 | 0.0096609 | -6.0651 | 2.1708 | 135 | -5.0914 | 2.175 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 0.50479 | 0.484 | 0.0096609 | -4.1178 | 2.1791 | 150 | -4.2026 | 2.1772 | 45
| 157.5 |  | 0.50568 | 0.484 | 0.0096609 | -4.2875 | 2.1753 | 165 | -4.0613 | 2.077 | 60 | 0.099737 | 0.05602 |  |  |  | 0.15625
| 172.5 |  | 0.50538 | 0.486 | 0.0084327 | -3.8351 | 1.9787 | 180 | -3.7805 | 1.9758 | 75 | 0.21049 | 0.048944 |  |  |  | 0.25991
| 187.5 |  | 0.50689 | 0.488 | 0.0063246 | -3.7259 | 1.9728 | 195 | -4.3664 | 2.0704 | 90 | 0.26395 | 0.013134 |  |  |  | 0.27728
| 202.5 |  | 0.50741 | 0.482 | 0.010328 | -5.007 | 2.1679 | 210 |  |  | 105 | 0.17989 | 0.093033 |  |  |  | 0.27354
| 217.5 |  | 0.50845 | 0.478 | 0.010328 | -5.9897 | 2.1634 | 225 |  |  | 120 | 0.092168 | 0.11569 |  |  |  | 0.20867 | [12 empty] | 0.10061 | 0.19334 | 0.014911
[8 empty] | 240 |  |  | 135 | 0.15984 | 0.064269 |  |  |  | 0.22468 | [12 empty] | 0.031091 | 0.19101 | 0.033099
[8 empty] | 255 |  |  | 150 | 0.23585 | 0.012955 |  |  |  | 0.249 | [12 empty] | 0.026269 | 0.20912 | 0.039219
[8 empty] | 270 |  |  | 165 | 0.21148 | 0.020141 |  |  |  | 0.23168 | [12 empty] | 0.014998 | 0.19596 | 0.035093
[8 empty] | 285 |  |  | 180 | 0.20348 | 0.0049657 |  |  |  | 0.20823 | [12 empty] | 0.026654 | 0.17609 | 0.031575
[8 empty] | 300 |  |  | 195 | 0.16201 | 0.049511 |  |  |  | 0.21181 | [12 empty] | 0.014123 | 0.1759 | 0.035294
[8 empty] | 315 |  |  | 210 | [5 empty] | 0.24692 | [13 empty] | 0.18198 | 0.064034
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.17998 | [13 empty] | 0.11143 | 0.06773
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.18944 | [13 empty] | 0.15592 | 0.032952
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.20084 | [13 empty] | 0.16013 | 0.040066
[8 empty] | 375 |  |  | 270 | [20 empty] | 0.037123
[8 empty] | 390 |  |  | 285 | [20 empty] | 0.045688
[8 empty] | 405 |  |  | 300 | [20 empty] | 0.035573
[8 empty] | 420 |  |  | 315 | [20 empty] | 0.020178
[8 empty] | 435 |  |  | 330 | [20 empty] | 0.0011145
[8 empty] | 450 |  |  | 345 | [20 empty] | 0.01593
[8 empty] | 465 |  |  | 360 | [20 empty] | 0.0031023
[8 empty] | 480 |  |  | 375 | [20 empty] | 0.00097888
[8 empty] | 495 |  |  | 390 | [20 empty] | 0.011068
[8 empty] | 510 |  |  | 405
[8 empty] | 525 |  |  | 420
435
450
45
PTB | 55 | 0.49548 | 0.49827 | 0.45083 | 0.0071728 | -9.5205 | 0.47 | 60 | -9.6222 | 0.44
| 70 | 0.49548 | 0.49811 | 0.44917 | 0.0040825 | -9.8256 | 0.38 | 75 | -9.8239 | 0.41
| 85 | 0.49548 | 0.49808 | 0.44917 | 0.0071728 | -9.8205 | 0.47 | 90 | -9.7086 | 0.44
| 100 | 0.49548 | 0.49807 | 0.45083 | 0.0040825 | -9.4848 | 0.38 | 105 | -8.9623 | 0.43333
| 115 | 0.49548 | 0.49593 | 0.45667 | 0.0096309 | -7.9172 | 0.54 | 120 | -7.7579 | 0.54667
| 130 | 0.49548 | 0.49607 | 0.45917 | 0.01018 | -7.4395 | 0.56 | 135 | -7.73 | 0.54667
| 145 | 0.49548 | 0.49624 | 0.455 | 0.0088465 | -8.3109 | 0.52 | 150 | -8.0959 | 0.53
| 160 | 0.49548 | 0.49548 | 0.4575 | 0.0098907 | -7.6659 | 0.55 | 165 | -7.5497 | 0.55333
| 175 | 0.49548 | 0.49722 | 0.46083 | 0.01018 | -7.3173 | 0.56 | 180 | -7.1343 | 0.55667
| 190 | 0.49548 | 0.49608 | 0.4625 | 0.0098907 | -6.7683 | 0.55 | 195 | -7.0387 | 0.55333
| 205 | 0.49548 | 0.49772 | 0.46 | 0.010215 | -7.5794 | 0.56 | 210 | -6.9799 | 0.51
| 220 | 0.49548 | 0.49707 | 0.46833 | 0.0056466 | -5.7808 | 0.41 | 225 | -6.0403 | 0.44
| 235 | 0.49548 | 0.49786 | 0.4652 | 0.0087178 | -6.5594 | 0.5 | 240 | -6.5986 | 0.5
| 250 | 0.49548 | 0.49897 | 0.46565 | 0.0084348 | -6.6771 | 0.5 | 255 | -6.5681 | 0.48333
| 265 | 0.49548 | 0.4992 | 0.4675 | 0.0067566 | -6.35 | 0.45 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 0.54521 | 0.5035 | 0.015313 | -7.6503 | 2.24 | 45
| 60 |  | 0.54811 | 0.5095 | 0.014681 | -7.0442 | 2.1323 | 60 | -7.0442 | 2.1323
| 75 |  | 0.54571 | 0.5155 | 0.01099 | -5.5359 | 2.1323 | 75 | -5.5359 | 2.1323
| 90 |  | 0.54656 | 0.5185 | 0.0067082 | -5.1339 | 2.1323 | 90 | -5.1339 | 2.1323
| 105 |  | 0.54526 | 0.521 | 0.0055251 | -4.4493 | 2.1323 | 105 | -4.4493 | 2.1323
| 120 |  | 0.54606 | 0.5225 | 0.007864 | -4.3145 | 2.1323 | 120 | -4.3145 | 2.1323
| 135 |  | 0.54596 | 0.524 | 0.0082078 | -4.0223 | 2.1323 | 135 | -4.0223 | 2.1323
| 150 |  | 0.54576 | 0.524 | 0.0082078 | -3.9871 | 2.1323 | 150 | -3.9871 | 2.1323
| 165 |  | 0.54636 | 0.526 | 0.0094032 | -3.7265 | 2.1323 | 165 | -3.7265 | 2.1323
| 180 |  | 0.54516 | 0.525 | 0.0088852 | -3.698 | 2.1323 | 180 | -3.698 | 2.1323
| 195 |  | 0.54636 | 0.527 | 0.0097872 | -3.5435 | 2.1323 | 195 | -3.5435 | 2.1323
| 210 |  | 0.54586 | 0.53 | 0.01026 | -2.9055 | 2.1323 | 210 | -2.9055 | 2.1323
| 225 | 0.54576 | 0.54576 | 0.529 | 0.010208 | -3.0709 | 2.1323 | 225 | -3.0709 | 2.1323
| 240 |  | 0.54596 | 0.527 | 0.0097872 | -3.4728 | 2.1323 | 240 | -3.4728 | 2.1323
| 255 |  | 0.54576 | 0.528 | 0.010052 | -3.2542 | 2.1323 | 255 | -3.2542 | 2.1323
| 270 |  | 0.54506 | 0.527 | 0.0097872 | -3.3134 | 2.1323 | 270 | -3.3134 | 2.1323
| 285 |  | 0.54506 | 0.528 | 0.010052 | -3.1299 | 2.1323 | 285 | -3.1299 | 2.1323
| 300 |  | 0.54516 | 0.527 | 0.0097872 | -3.3311 | 2.1323 | 300 | -3.3311 | 2.1323
| 315 |  | 0.54486 | 0.526 | 0.0094032 | -3.4614 | 2.1323 | 315 | -3.4614 | 2.1323
| 330 |  | 0.54616 | 0.525 | 0.0088852 | -3.8743 | 2.1323 | 330 | -3.8743 | 2.1323
| 345 |  | 0.54496 | 0.521 | 0.0044721 | -4.3967 | 2.1323 | 345 | -4.3967 | 2.1323
| 360 |  | 0.54636 | 0.523 | 0.007327 | -4.2756 | 2.1323 | 360 | -4.2756 | 2.1323
| 375 |  | 0.54561 | 0.522 | 0.0061559 | -4.3273 | 2.1323 | 375 | -4.3273 | 2.1323
| 390 |  | 0.54741 | 0.521 | 0.0030779 | -4.8245 | 2.1323 | 390 | -4.8245 | 2.1323
| 405 |  | 0.55411 | 0.5215 | 0.0048936 | -5.8851 | 2.1323 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 0.50295 | 0.48 | 0.02 | -4.5629 | 2.3085 | 120 | -4.5629 | 2.3085
| 135 |  | 0.503 | 0.48 | 0.02 | -4.5735 | 2.308 | 135 | -4.5735 | 2.308
| 150 |  | 0.50336 | 0.48 | 0.02 | -4.6402 | 2.3052 | 150 | -4.6402 | 2.3052
| 165 |  | 0.50162 | 0.48 | 0.02 | -4.311 | 2.3192 | 165 | -4.311 | 2.3192
| 180 |  | 0.50117 | 0.48 | 0.02 | -4.2239 | 2.3229 | 180 | -4.2239 | 2.3229
| 195 |  | 0.50068 | 0.48 | 0.02 | -4.1313 | 2.3268 | 195 | -4.1313 | 2.3268
| 210 |  | 0.49985 | 0.48 | 0.02 | -3.9721 | 2.3336 | 210 | -3.9721 | 2.3336
| 225 |  | 0.50104 | 0.48 | 0.02 | -4.199 | 2.3239 | 225 | -4.199 | 2.3239
| 240 |  | 0.50011 | 0.48 | 0.02 | -4.0217 | 2.3315 | 240 | -4.0217 | 2.3315
| 255 |  | 0.49959 | 0.48 | 0.02 | -3.9216 | 2.3357 | 255 | -3.9216 | 2.3357
| 270 | 0.5 | 0.49964 | 0.48 | 0.02 | -3.9318 | 2.3353 | 270 | -3.9318 | 2.3353
| 285 |  | 0.49943 | 0.48 | 0.02 | -3.891 | 2.3371 | 285 | -3.891 | 2.3371
| 300 |  | 0.4996 | 0.48 | 0.02 | -3.9237 | 2.3357 | 300 | -3.9237 | 2.3357
| 315 |  | 0.49895 | 0.48 | 0.02 | -3.7976 | 2.341 | 315 | -3.7976 | 2.341
| 330 |  | 0.49925 | 0.48 | 0.02 | -3.8558 | 2.3386 | 330 | -3.8558 | 2.3386
| 345 |  | 0.50068 | 0.48 | 0.02 | -4.1313 | 2.3268 | 345 | -4.1313 | 2.3268
| 360 |  | 0.50117 | 0.48 | 0.02 | -4.2239 | 2.3229 | 360 | -4.2239 | 2.3229
| 375 |  | 0.50162 | 0.48 | 0.02 | -4.311 | 2.3192 | 375 | -4.311 | 2.3192
| 390 |  | 0.50336 | 0.48 | 0.02 | -4.6402 | 2.3052 | 390 | -4.6402 | 2.3052
| 405 |  | 0.503 | 0.48 | 0.02 | -4.5735 | 2.308 | 405 | -4.5735 | 2.308
| 420 |  | 0.50295 | 0.48 | 0.02 | -4.5629 | 2.3085 | 420 | -4.5629 | 2.3085
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 2
| 2 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.58788
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 1.9412 | 1.8 | 0.010541 | -7.2723 | 0.87576 | 45
| 52.5 |  | 1.9396 | 1.804 | 0.0096609 | -6.9925 | 0.87646 | 60 | -6.9431 | 0.84975 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 1.944 | 1.81 | 0 | -6.8937 | 0.82304 | 75 | -6.613 | 0.8485 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 1.9452 | 1.822 | 0.010328 | -6.3324 | 0.87396 | 90 | -6.1371 | 0.84816 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 1.9456 | 1.83 | 4.6811e-16 | -5.9418 | 0.82237 | 105 | -5.7726 | 0.84773 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 1.9471 | 1.838 | 0.010328 | -5.6035 | 0.87309 | 120 | -5.4319 | 0.87325 |  | 0.80725 | 0.28267 | 0.30289 |  | 0.57854 | 1.1858 | 0.06143 |  | 1.2227 | 0.4513 |  | 0.37316 |  | 0.6394 |  | 0.22699 | 0.088614 | 0.0045461 | 0.48224 | 1.3787 | 0.21607 | 0.55729 |  | 0.20321 | 0.46193 | 0.82482 | 0.33123 | 0.58358 |  | 0.09303 | 0.10756 | 0.25597 | 0.049183 | 0.21784 |  | 0.050364
| 127.5 | 1.947 | 1.9464 | 1.844 | 0.0096609 | -5.2604 | 0.87341 | 135 | -5.136 | 0.84747 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 1.9476 | 1.85 | 0 | -5.0116 | 0.82152 | 150 | -4.7022 | 0.84768 | 45
| 157.5 |  | 1.9455 | 1.86 | 0.010541 | -4.3929 | 0.87383 | 165 | -4.4569 | 0.87371 | 60 | 1.0416 | 0.30293 |  |  | 1.0094 | 0.71978 |  |  | 0.35423 |  |  | 0.79341
| 172.5 |  | 1.946 | 1.858 | 0.010328 | -4.5209 | 0.8736 | 180 | -4.679 | 0.84787 | 75 | 1.304 | 0.10045 |  |  | 0.62881 | 1.2293 |  |  | 0.24837 |  |  | 0.56128
| 187.5 |  | 1.9461 | 1.852 | 0.0063246 | -4.8371 | 0.82214 | 195 | -5.014 | 0.8481 | 90 | 1.4521 | 0.38324 | 0.67298 |  | 0.36113 | 1.0484 | 0.030476 |  | 0.64602 | 0.48205 |  | 0.074749 |  | 0.39406 | [16 empty] | 0.20505 | 0.46765 | 0.28364 | 0.45064 |  | 0.29892
| 202.5 |  | 1.945 | 1.844 | 0.0096609 | -5.191 | 0.87405 | 210 |  |  | 105 | 1.2786 | 0.14488 | 0.55897 |  | 0.1266 | 1.15 | 0.0080115 |  | 0.77723 | 0.48869 |  | 0.018203 |  | 0.43464 |  |  | 0.098557 | [7 empty] | 0.26405 | 0.80812 | 0.33297 | 0.61097 |  | 0.31827 | 0.22203 | 0.45343 | 0.25188 | 0.42314 |  | 0.25269
| 217.5 |  | 1.9462 | 1.832 | 0.0063246 | -5.8671 | 0.82212 | 225 |  |  | 120 | 1.3951 | 0.25646 | 0.49309 |  | 0.18574 | 1.1592 | 0.14044 |  | 1.2197 | 0.36541 |  | 0.38717 |  | 0.57 |  | 0.89342 | 0.08321 | 0.48319 | 1.2551 | 0.093302 | 1.0448 | 0.17693 |  | 1.1681 | 0.2796 | 0.88581 | 0.40326 | 0.57779 |  | 0.15277 | 0.21228 | 0.46919 | 0.26572 | 0.38939 |  | 0.15902
[8 empty] | 240 |  |  | 135 | 1.3964 | 0.18357 | 0.37299 |  | 0.44222 | 1.2263 | 0.25675 |  | 1.482 | 0.28183 |  | 0.58788 |  | 0.61628 |  | 0.59019 | 0.043952 | 0.30713 | 0.65499 | 0.59064 | 0.4912 | 0.024333 |  | 0.94353 | 0.27308 | 0.87076 | 0.36048 | 0.48432 |  | 0.0015106 | 0.16685 | 0.41969 | 0.20455 | 0.30001 |  | 0.0457
[8 empty] | 255 |  |  | 150 | 1.3639 | 0.36678 | 0.3334 |  | 0.26177 | 0.97635 | 0.28026 |  | 1.2615 | 0.14903 |  | 0.54426 |  | 0.46802 |  | 0.5204 | 0.025759 | 0.2011 | 0.61964 | 0.59538 | 0.27754 | 0.0036579 |  | 0.73861 | 0.21881 | 0.80214 | 0.39232 | 0.41338 |  | 0.051679 | 0.068038 | 0.31388 | 0.1431 | 0.19169 |  | 0.0032222
[8 empty] | 270 |  |  | 165 | 1.5639 | 0.47054 | 0.35925 |  | 0.10293 | 1.0814 | 0.35685 |  | 1.2521 | 0.12045 |  | 0.46755 |  | 0.39507 |  | 0.36261 | 0.016716 | 0.14147 | 0.34504 | 1.1214 | 0.10438 | 0.17269 |  | 0.3732 | 0.19228 | 0.87205 | 0.41795 | 0.41269 |  | 0.11903 | 0.066298 | 0.35329 | 0.16387 | 0.20045 |  | 0.037521
[8 empty] | 285 |  |  | 180 | 1.4905 | 0.30661 | 0.22045 |  | 0.16354 | 1.179 | 0.45711 |  | 1.2463 | 0.065976 |  | 0.39895 |  | 0.30356 |  | 0.49625 | 0.031246 | 0.18443 | 0.27304 | 1.1158 | 0.017858 | 0.07426 |  | 0.37284 | 0.36491 | 1.0086 | 0.51114 | 0.44598 |  | 0.24406 | 0.12581 | 0.39545 | 0.18864 | 0.20314 |  | 0.080053
[8 empty] | 300 |  |  | 195 | 1.0757 | 0.11916 | 0.052839 |  | 0.37387 | 1.2522 | 0.43935 |  | 1.1777 | 0.11376 |  | 0.2885 |  | 0.27983 |  | 0.53641 | 0.063493 | 0.16702 | 0.0010814 | 1.0309 | 0.11651 | 0.051827 |  | 0.36796 | 0.53976 | 1.0139 | 0.48645 | 0.461 |  | 0.28384 | 0.16701 | 0.36163 | 0.14281 | 0.17802 |  | 0.064187
[8 empty] | 315 |  |  | 210 | [5 empty] | 0.99575 | 0.4888 |  | 1.1696 | 0.044811 |  | 0.44625 |  | 0.23133 |  | 0.48818 | 0.03394 | 0.17745 |  | 0.83937 | 0.076227 | 0.083578 |  | 0.49028 |  | 0.87848 | 0.45718 | 0.32356 |  | 0.16644 |  | 0.33603 | 0.16179 | 0.13572 |  | 0.041109
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.0775 | 0.49204 |  | 1.1591 | 0.012467 |  | 0.38377 |  | 0.22284 |  | 0.58044 | 0.0905 | 0.15736 |  | 0.77948 | 0.19675 | 0.11403 |  | 0.52142 |  | 0.94771 | 0.4919 | 0.37409 |  | 0.23429 |  | 0.30449 | 0.11624 | 0.10501 |  | 0.012817
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.2166 | 0.50184 |  | 1.1886 | 0.037906 |  | 0.32159 |  | 0.23082 |  | 0.27585 | 0.079117 | 0.035984 |  | 1.4161 | 0.22342 | 0.15135 |  | 0.14702 |  | 0.89248 | 0.38414 | 0.32684 |  | 0.17107 |  | 0.29452 | 0.082199 | 0.091996 |  | 0.0042233
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.2023 | 0.53191 |  | 1.3941 | 0.002976 |  | 0.52177 |  | 0.32101 |  | 0.19983 | 0.049493 | 0.034547 |  | 1.6506 | 0.44244 | 0.22825 |  | 0.17815 |  | 0.91439 | 0.41144 | 0.32304 |  | 0.082316 |  | 0.33637 | 0.12605 | 0.12013 |  | 0.014051
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.053078 |  | 0.59024 |  | 0.31154 |  | 0.25609 | 0.077965 | 0.028651 |  |  | 0.55656 | 0.23057 |  | 0.15964 |  |  | 0.52221 | 0.36889 |  | 0.14542 |  |  | 0.14366 | 0.11545 |  | 0.01485
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.017245 |  | 0.38815 |  | 0.25253 |  | 0.1445 | 0.038438 | 0.02235 |  |  | 0.46628 | 0.2544 |  | 0.029519 |  |  | 0.36594 | 0.29825 |  | 0.11745 |  |  | 0.11868 | 0.11859 |  | 0.01406
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.070742 |  | 0.2702 |  | 0.10034 |  | 0.28351 | 0.035031 | 0.087455 |  |  | 0.45163 | 0.16416 |  | 0.081958 |  |  | 0.50592 | 0.34407 |  | 0.31357 |  |  | 0.18311 | 0.14602 |  | 0.10869
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.11035 |  | 0.38272 |  | 0.13291 |  | 0.28658 | 0.069718 | 0.051627 |  |  | 0.61639 | 0.21052 |  | 0.098385 |  |  | 0.58906 | 0.38037 |  | 0.32613 |  |  | 0.18202 | 0.12997 |  | 0.077834
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.17525 |  | 0.39328 |  | 0.080088 |  | 0.27569 |  |  |  |  | 0.85506 | 0.27043 |  | 0.27015 |  |  | 0.69437 | 0.41565 |  | 0.41728
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.23536 | [5 empty] | 0.23326 |  |  |  |  | 0.89725 | 0.23239 |  |  |  |  | 0.66971 | 0.35404
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.18455 | [5 empty] | 0.0345 |  |  |  |  | 0.86256 | 0.26447 |  |  |  |  | 0.45006 | 0.21979
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.35493 | [5 empty] | 0.014441 |  |  |  |  | 1.3263 | 0.33701 |  |  |  |  | 0.62341 | 0.23367
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.31536 | [5 empty] | 0.094064 |  |  |  |  | 1.1227 | 0.26996 |  |  |  |  | 0.63496 | 0.27088
[8 empty] | 510 |  |  | 405 | [15 empty] | 0.054311 | [5 empty] | 0.31185 | [5 empty] | 0.1834
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.033395 | [5 empty] | 0.27807 | [5 empty] | 0.17547
[11 empty] | 435 | [27 empty] | 0.23913
450
45
PTB | 55 | 1.9873 | 1.994 | 1.8317 | 0.0057621 | -8.1352 | 0.49 | 60 | -8.1342 | 0.49
| 70 | 1.9873 | 1.9929 | 1.8309 | 0.0073318 | -8.1322 | 0.49 | 75 | -8.1031 | 0.49
| 85 | 1.9873 | 1.9918 | 1.8316 | 0.0055377 | -8.0449 | 0.49 | 90 | -7.7959 | 0.49
| 100 | 1.9873 | 1.9891 | 1.8439 | 0.0094094 | -7.2981 | 0.49 | 105 | -7.2329 | 0.49
| 115 | 1.9873 | 1.9896 | 1.8483 | 0.0083406 | -7.1025 | 0.49 | 120 | -7.0538 | 0.49333
| 130 | 1.9873 | 1.9883 | 1.85 | 0.010445 | -6.9564 | 0.5 | 135 | -6.7346 | 0.49667
| 145 | 1.9873 | 1.9872 | 1.8622 | 0.0099802 | -6.2908 | 0.49 | 150 | -6.2598 | 0.49
| 160 | 1.9873 | 1.9873 | 1.8642 | 0.0092861 | -6.1979 | 0.49 | 165 | -6.2756 | 0.49333
| 175 | 1.9873 | 1.9883 | 1.8604 | 0.010215 | -6.4312 | 0.5 | 180 | -6.3857 | 0.49667
| 190 | 1.9873 | 1.9873 | 1.8622 | 0.0099802 | -6.2947 | 0.49 | 195 | -6.2428 | 0.49
| 205 | 1.9873 | 1.9879 | 1.8658 | 0.01018 | -6.1392 | 0.49 | 210 | -6.1038 | 0.49
| 220 | 1.9873 | 1.9883 | 1.8683 | 0.008165 | -6.0329 | 0.49 | 225 | -6.0781 | 0.49
| 235 | 1.9873 | 1.9883 | 1.8657 | 0.0084348 | -6.1686 | 0.49 | 240 | -6.2107 | 0.49
| 250 | 1.9873 | 1.991 | 1.8657 | 0.0084348 | -6.295 | 0.49 | 255 | -6.4406 | 0.49333
| 265 | 1.9873 | 1.9938 | 1.8596 | 0.010215 | -6.732 | 0.5 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 1.996 | 1.826 | 0.012312 | -8.5166 | 0.91 | 45
| 60 |  | 2.003 | 1.856 | 0.0094032 | -7.3385 | 0.79771 | 60 | -7.3385 | 0.79771
| 75 |  | 2.0031 | 1.868 | 0.0089443 | -6.7441 | 0.79771 | 75 | -6.7441 | 0.79771
| 90 |  | 2.0083 | 1.875 | 0.0088852 | -6.637 | 0.79771 | 90 | -6.637 | 0.79771
| 105 |  | 2.0077 | 1.888 | 0.0061559 | -5.9616 | 0.79771 | 105 | -5.9616 | 0.79771
| 120 |  | 2.01 | 1.894 | 0.0082078 | -5.7707 | 0.79771 | 120 | -5.7707 | 0.79771
| 135 |  | 2.009 | 1.901 | 0.010208 | -5.3753 | 0.79771 | 135 | -5.3753 | 0.79771
| 150 |  | 2.0112 | 1.907 | 0.007327 | -5.1805 | 0.79771 | 150 | -5.1805 | 0.79771
| 165 |  | 2.0122 | 1.91 | 0.0064889 | -5.0785 | 0.79771 | 165 | -5.0785 | 0.79771
| 180 |  | 2.0101 | 1.908 | 0.0089443 | -5.0789 | 0.79771 | 180 | -5.0789 | 0.79771
| 195 |  | 2.0086 | 1.911 | 0.0044721 | -4.8586 | 0.79771 | 195 | -4.8586 | 0.79771
| 210 |  | 2.0106 | 1.91 | 0.0064889 | -5.003 | 0.79771 | 210 | -5.003 | 0.79771
| 225 | 2.0092 | 2.0092 | 1.911 | 0.0078807 | -4.887 | 0.79771 | 225 | -4.887 | 0.79771
| 240 |  | 2.0077 | 1.91 | 0.0064889 | -4.8658 | 0.79771 | 240 | -4.8658 | 0.79771
| 255 |  | 2.0097 | 1.907 | 0.007327 | -5.1097 | 0.79771 | 255 | -5.1097 | 0.79771
| 270 |  | 2.0071 | 1.9 | 0.01026 | -5.3356 | 0.79771 | 270 | -5.3356 | 0.79771
| 285 |  | 2.0028 | 1.898 | 0.010052 | -5.2346 | 0.79771 | 285 | -5.2346 | 0.79771
| 300 |  | 2.005 | 1.894 | 0.0082078 | -5.538 | 0.79771 | 300 | -5.538 | 0.79771
| 315 |  | 2.0053 | 1.89 | 0.011239 | -5.7516 | 0.79771 | 315 | -5.7516 | 0.79771
| 330 |  | 2.0058 | 1.886 | 0.0082078 | -5.9746 | 0.79771 | 330 | -5.9746 | 0.79771
| 345 |  | 2.0058 | 1.881 | 0.010208 | -6.2238 | 0.79771 | 345 | -6.2238 | 0.79771
| 360 |  | 2.0035 | 1.879 | 0.010208 | -6.216 | 0.79771 | 360 | -6.216 | 0.79771
| 375 |  | 2.0102 | 1.875 | 0.0088852 | -6.7252 | 0.79771 | 375 | -6.7252 | 0.79771
| 390 |  | 2.0014 | 1.87 | 2.2781e-16 | -6.5649 | 0.79771 | 390 | -6.5649 | 0.79771
| 405 |  | 1.9822 | 1.874 | 0.0082078 | -5.4581 | 0.94 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 2.0035 | 1.86 | 0.02 | -7.1636 | 0.89193 | 120 | -7.1636 | 0.89193
| 135 |  | 2.0007 | 1.88 | 0.02 | -6.0338 | 0.90292 | 135 | -6.0338 | 0.90292
| 150 |  | 2.0011 | 1.89 | 0.02 | -5.5533 | 0.90688 | 150 | -5.5533 | 0.90688
| 165 |  | 1.9987 | 1.9 | 0.02 | -4.9379 | 0.9133 | 165 | -4.9379 | 0.9133
| 180 |  | 2.0012 | 1.9 | 0.02 | -5.0548 | 0.91117 | 180 | -5.0548 | 0.91117
| 195 |  | 2.0003 | 1.9 | 0.02 | -5.0155 | 0.91189 | 195 | -5.0155 | 0.91189
| 210 |  | 2.0022 | 1.9 | 0.02 | -5.1056 | 0.91025 | 210 | -5.1056 | 0.91025
| 225 |  | 2.0032 | 1.9 | 0.02 | -5.1517 | 0.90941 | 225 | -5.1517 | 0.90941
| 240 |  | 2.0015 | 1.91 | 0 | -4.5714 | 0.86854 | 240 | -4.5714 | 0.86854
| 255 |  | 2.0006 | 1.91 | 0 | -4.5265 | 0.86936 | 255 | -4.5265 | 0.86936
| 270 | 2.002 | 2.0021 | 1.91 | 0 | -4.6024 | 0.86798 | 270 | -4.6024 | 0.86798
| 285 |  | 2.0025 | 1.91 | 0 | -4.6204 | 0.86765 | 285 | -4.6204 | 0.86765
| 300 |  | 1.9985 | 1.9 | 0.02 | -4.9294 | 0.91346 | 300 | -4.9294 | 0.91346
| 315 |  | 1.9983 | 1.9 | 0.02 | -4.9209 | 0.91362 | 315 | -4.9209 | 0.91362
| 330 |  | 1.9962 | 1.9 | 0.02 | -4.821 | 0.91544 | 330 | -4.821 | 0.91544
| 345 |  | 2.0003 | 1.9 | 0.02 | -5.0155 | 0.91189 | 345 | -5.0155 | 0.91189
| 360 |  | 2.0012 | 1.9 | 0.02 | -5.0548 | 0.91117 | 360 | -5.0548 | 0.91117
| 375 |  | 1.9987 | 1.9 | 0.02 | -4.9379 | 0.9133 | 375 | -4.9379 | 0.9133
| 390 |  | 2.0011 | 1.9 | 0.02 | -5.0536 | 0.91119 | 390 | -5.0536 | 0.91119
| 405 |  | 2.0007 | 1.9 | 0.02 | -5.0342 | 0.91155 | 405 | -5.0342 | 0.91155
| 420 |  | 2.0035 | 1.9 | 0.02 | -5.1671 | 0.90912 | 420 | -5.1671 | 0.90912
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
| 105 |  | 2.0522 | 1.949 | 0.0044721 | -5.0301 | 2.6162 | 105 | -5.0301 | 2.6162
| 120 |  | 2.0502 | 1.955 | 0.0088852 | -4.6455 | 2.6082 | 120 | -4.6455 | 2.6082
| 135 |  | 2.0506 | 1.961 | 0.010208 | -4.3716 | 2.6023 | 135 | -4.3716 | 2.6023
| 150 |  | 2.0509 | 1.967 | 0.007327 | -4.0917 | 2.5923 | 150 | -4.0917 | 2.5923
| 165 |  | 2.0483 | 1.968 | 0.0061559 | -3.9192 | 2.591 | 165 | -3.9192 | 2.591
| 180 |  | 2.0449 | 1.97 | 4.5563e-16 | -3.6622 | 2.5883 | 180 | -3.6622 | 2.5883
| 195 |  | 2.048 | 1.976 | 0.0094032 | -3.514 | 2.5805 | 195 | -3.514 | 2.5805
| 210 |  | 2.0528 | 1.976 | 0.0094032 | -3.7393 | 2.5805 | 210 | -3.7393 | 2.5805
| 225 |  | 2.0482 | 1.976 | 0.0094032 | -3.5273 | 2.5805 | 225 | -3.5273 | 2.5805
| 240 |  | 2.0468 | 1.969 | 0.0078807 | -3.8007 | 2.5896 | 240 | -3.8007 | 2.5896
| 255 | 2.0457 | 2.0472 | 1.966 | 0.0082078 | -3.9674 | 2.5936 | 255 | -3.9674 | 2.5936
| 270 |  | 2.0465 | 1.967 | 0.007327 | -3.8863 | 2.5923 | 270 | -3.8863 | 2.5923
| 285 |  | 2.0452 | 1.959 | 0.010208 | -4.2147 | 2.605 | 285 | -4.2147 | 2.605
| 300 |  | 2.0465 | 1.962 | 0.010052 | -4.1306 | 2.5994 | 300 | -4.1306 | 2.5994
| 315 |  | 2.0472 | 1.963 | 0.0097872 | -4.1139 | 2.5976 | 315 | -4.1139 | 2.5976
| 330 |  | 2.0468 | 1.964 | 0.0094032 | -4.0449 | 2.5962 | 330 | -4.0449 | 2.5962
| 345 |  | 2.0482 | 1.959 | 0.010208 | -4.3573 | 2.605 | 345 | -4.3573 | 2.605
| 360 |  | 2.0528 | 1.951 | 0.010208 | -4.9572 | 2.6157 | 360 | -4.9572 | 2.6157
| 375 |  | 2.048 | 1.946 | 0.0082078 | -4.9788 | 2.6203 | 375 | -4.9788 | 2.6203
| 390 |  | 2.0449 | 1.947 | 0.007327 | -4.787 | 2.6189 | 390 | -4.787 | 2.6189
| 405 |  | 2.0483 | 1.942 | 0.010052 | -5.1885 | 2.6262 | 405 | -5.1885 | 2.6262
| 420 |  | 2.0509 | 1.943 | 0.0097872 | -5.262 | 2.6243 | 420 | -5.262 | 2.6243
| 435 |  | 2.0506 | 1.946 | 0.0082078 | -5.103 | 2.6203 | 435 | -5.103 | 2.6203
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
| 80 |  | 2.1237 | 1.9447 | 0.0088278 | -8.4272 | 2.3628 | 90 | -7.8716 | 2.3617
| 95 |  | 2.1237 | 1.9624 | 0.011171 | -7.5938 | 2.3612 | 105 | -7.213 | 2.3613
| 110 |  | 2.1227 | 1.9736 | 0.0076723 | -7.0226 | 2.3613 | 120 | -6.7055 | 2.3586
| 125 |  | 2.1257 | 1.9865 | 0.0076157 | -6.547 | 2.3572 | 135 | -6.097 | 2.361
| 140 |  | 2.1197 | 1.9952 | 0.0088058 | -5.872 | 2.3629 | 150 | -5.5628 | 2.3662
| 155 |  | 2.1147 | 2.0003 | 0.010005 | -5.4083 | 2.3679 | 165 | -5.3876 | 2.3671
| 170 |  | 2.1157 | 2.0019 | 0.0098183 | -5.3773 | 2.3667 | 180 | -5.2481 | 2.3666
| 185 |  | 2.1157 | 2.006 | 0.0080067 | -5.1836 | 2.3665 | 195 | -5.1504 | 2.3657
| 200 |  | 2.1167 | 2.008 | 0.0060142 | -5.1339 | 2.3653 | 210 | -4.888 | 2.3665
| 215 |  | 2.1147 | 2.0139 | 0.0079457 | -4.7651 | 2.3671 | 225 | -4.8551 | 2.365
| 230 |  | 2.1177 | 2.0139 | 0.0079048 | -4.9 | 2.3639 | 240 | -4.963 | 2.3655
| 245 |  | 2.1157 | 2.01 | 0.0039189 | -4.9945 | 2.3663 | 255 | -5.1174 | 2.3664
| 260 |  | 2.1157 | 2.0061 | 0.0089692 | -5.1788 | 2.3665 | 270 | -5.1994 | 2.3673
| 275 |  | 2.1147 | 2.0045 | 0.0089211 | -5.2096 | 2.3676 | 285 | -5.2788 | 2.3662
| 290 |  | 2.1167 | 2.0042 | 0.010012 | -5.3134 | 2.3655 | 300 | -5.3566 | 2.3652
| 305 |  | 2.1172 | 2.0033 | 0.0094321 | -5.3783 | 2.3651 | 315 | -5.4687 | 2.3648
| 320 |  | 2.1177 | 2.0009 | 0.0099707 | -5.5139 | 2.3647 | 330 | -5.5251 | 2.3654
| 335 |  | 2.1167 | 1.9996 | 0.010001 | -5.5307 | 2.3658 | 345 | -5.6203 | 2.3652
| 350 |  | 2.1177 | 1.9977 | 0.0097452 | -5.665 | 2.3649 | 360 | -5.7428 | 2.3646
| 365 |  | 2.1182 | 1.9957 | 0.0090404 | -5.7817 | 2.3645 | 375 | -5.815 | 2.3653
| 380 |  | 2.1172 | 1.9937 | 0.0077775 | -5.8317 | 2.3657 | 390 | -5.7561 | 2.3656
| 395 |  | 2.1172 | 1.9961 | 0.0092029 | -5.7183 | 2.3655 | 405 | -5.8456 | 2.3649
| 410 |  | 2.1182 | 1.993 | 0.0071194 | -5.9092 | 2.3646 | 420 | -5.8903 | 2.3646
| 425 |  | 2.1182 | 1.9936 | 0.0076897 | -5.8809 | 2.3646 | 435 | -5.9586 | 2.3643
| 440 |  | 2.1187 | 1.9916 | 0.0054263 | -5.9975 | 2.3642 | 450 | -5.9436 | 2.3653
| 455 |  | 2.1172 | 1.9919 | 0.0059372 | -5.9167 | 2.3658 | 465 | -5.9536 | 2.3655
| 470 |  | 2.1177 | 1.9912 | 0.0061999 | -5.972 | 2.3653 | 480 | -6.0381 | 2.3654
| 485 |  | 2.1177 | 1.9891 | 0.007953 | -6.0711 | 2.3654 | 495 | -6.0537 | 2.3664
| 500 | 2.116 | 2.1163 | 1.9884 | 0.0054263 | -6.045 | 2.3669 | 510 | -6.1288 | 2.3649
| 515 |  | 2.1192 | 1.9884 | 0.0067519 | -6.1707 | 2.3639 | 525 | -6.115 | 2.3642
| 530 |  | 2.1187 | 1.9897 | 0.0074674 | -6.0872 | 2.3643
| 545 |  | 2.1187 | 1.9884 | 0.0053743 | -6.1485 | 2.3644
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
| 87 |  | 2.029 | 1.929 | 0.010208 | -4.9299 | 6.3 | 90 | -4.8281 | 6.3
| 102 |  | 2.0245 | 1.935 | 0.0088852 | -4.4209 | 6.3 | 105 | -4.3553 | 6.3
| 117 |  | 2.0207 | 1.938 | 0.010052 | -4.0927 | 6.3 | 120 | -4.0761 | 6.3
| 132 |  | 2.0148 | 1.934 | 0.010463 | -4.0097 | 6.3 | 135 | -4.0708 | 6.3
| 147 |  | 2.0087 | 1.922 | 0.010052 | -4.3155 | 6.3 | 150 | -4.2679 | 6.3
| 162 |  | 2.0037 | 1.922 | 0.010052 | -4.0775 | 6.3 | 165 | -4.0334 | 6.3
| 177 |  | 2.0012 | 1.924 | 0.0094032 | -3.8574 | 6.3 | 180 | -3.8758 | 6.3
| 192 |  | 2 | 1.921 | 0.010208 | -3.9496 | 6.3 | 195 | -3.9478 | 6.3
| 207 |  | 2.0008 | 1.922 | 0.010052 | -3.9407 | 6.3 | 210 | -3.9712 | 6.3
| 222 |  | 2.003 | 1.921 | 0.010208 | -4.0932 | 6.3 | 225 | -4.1457 | 6.3
| 237 |  | 2.0064 | 1.919 | 0.010208 | -4.356 | 6.3 | 240 | -4.3416 | 6.3
| 252 |  | 2.0091 | 1.923 | 0.011743 | -4.2837 | 6.3 | 255 | -4.3059 | 6.3
| 267 |  | 2.0103 | 1.922 | 0.010052 | -4.3946 | 6.3 | 270 | -4.4194 | 6.3
| 282 |  | 2.0088 | 1.918 | 0.010052 | -4.5186 | 6.3 | 285 | -4.4777 | 6.3
| 297 |  | 2.0055 | 1.919 | 0.010208 | -4.314 | 6.3 | 300 | -4.3703 | 6.3
| 312 |  | 2.002 | 1.91 | 0.0091766 | -4.5952 | 6.3 | 315 | -4.5908 | 6.3
| 327 | 1.9973 | 1.9973 | 1.906 | 0.0082078 | -4.5734 | 6.3 | 330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 42 |  | 2.03 | 1.83 | 0.012 | -9.852 | 1.4 | 45
| 57 |  | 2.023 | 1.843 | 0.01 | -8.898 | 1.4 | 60 | -8.6994 | 1.4
LEI | 72 |  | 2.024 | 1.864 | 0.01 | -7.905 | 1.4 | 75 | -7.7068 | 1.4
| 87 |  | 2.025 | 1.885 | 0.013 | -6.914 | 1.4 | 90 | -6.7652 | 1.4
| 102 |  | 2.026 | 1.901 | 0.01 | -6.17 | 1.4 | 105 | -5.9928 | 1.4
| 117 |  | 2.025 | 1.918 | 0.01 | -5.284 | 1.4 | 120 | -5.1066 | 1.4
| 132 |  | 2.024 | 1.935 | 0.013 | -4.397 | 1.4 | 135 | -4.367 | 1.4
| 147 |  | 2.025 | 1.939 | 0.01 | -4.247 | 1.4 | 150 | -4.247 | 1.4
| 162 |  | 2.025 | 1.939 | 0.01 | -4.247 | 1.4 | 165 | -4.2766 | 1.4
| 177 |  | 2.025 | 1.936 | 0.012 | -4.395 | 1.4 | 180 | -4.3946 | 1.4
| 192 |  | 2.026 | 1.937 | 0.01 | -4.393 | 1.4 | 195 | -4.3638 | 1.4
| 207 |  | 2.025 | 1.939 | 0.01 | -4.247 | 1.4 | 210 | -4.2376 | 1.4
| 222 |  | 2.024 | 1.939 | 0.01 | -4.2 | 1.4 | 225 | -4.2288 | 1.4
| 237 |  | 2.026 | 1.938 | 0.01 | -4.344 | 1.4 | 240 | -4.3142 | 1.4
| 252 |  | 2.026 | 1.941 | 0.01 | -4.195 | 1.4 | 255 | -4.2148 | 1.4
| 267 |  | 2.026 | 1.939 | 0.01 | -4.294 | 1.4 | 270 | -4.3232 | 1.4
| 282 |  | 2.027 | 1.937 | 0.012 | -4.44 | 1.4 | 285 | -4.5688 | 1.4
| 297 |  | 2.026 | 1.923 | 0.01 | -5.084 | 1.4 | 300 | -5.0746 | 1.4
| 312 |  | 2.025 | 1.923 | 0.012 | -5.037 | 1.4 | 315 | -5.0952 | 1.4
| 327 |  | 2.027 | 1.919 | 0.01 | -5.328 | 1.4 | 330 | -5.3 | 1.4
| 342 |  | 2.024 | 1.919 | 0.01 | -5.188 | 1.4 | 345
| 357 |  | 2.027 | 1.889 | 0.012 | -6.808 | 1.4 | 360
| 372 |  | 2.028 | 1.82 | 0.012 | -10.256 | 1.4 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 5
| 5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.4654
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 4.9215 | 4.634 | 0.018974 | -5.8421 | 0.71116 | 45
| 52.5 |  | 4.926 | 4.65 | 0.010541 | -5.6033 | 0.69021 | 60 | -5.3455 | 0.69032 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 4.9245 | 4.674 | 0.0096609 | -5.0876 | 0.69042 | 75 | -4.9192 | 0.69002 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 4.9302 | 4.696 | 0.0084327 | -4.7507 | 0.68962 | 90 | -4.3192 | 0.70003 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 4.9265 | 4.735 | 0.015811 | -3.8876 | 0.71044 | 105 | -3.7158 | 0.69971 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 4.9349 | 4.76 | 0.010541 | -3.5439 | 0.68897 | 120 | -3.2328 | 0.68945 |  | 0.47941 | 0.36116 | 0.82088 |  | 0.20652 | 0.9686 | 0.52492 |  | 0.66303 | 1.1612 |  | 0.10677 |  | 0.94985 |  | 0.60918 |  |  | 0.20246 | 0.72553 | 0.14761 | 0.99794 |  | 0.023453 | 0.76479 | 1.2066 | 0.52099 | 1.3777 |  | 0.5506
| 127.5 | 4.9278 | 4.928 | 4.784 | 0.0096609 | -2.9217 | 0.68994 | 135 | -2.7873 | 0.6896 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 4.9329 | 4.802 | 0.010328 | -2.653 | 0.68925 | 150 | -2.4989 | 0.69006 | 45
| 157.5 |  | 4.9214 | 4.806 | 0.0084327 | -2.3449 | 0.69086 | 165 | -2.4595 | 0.69048 | 60 | 1.5868 | 0.47813 |  |  | 1.7038 | 1.1919 |  |  | 0.50307 |  |  | 1.3812
| 172.5 |  | 4.9268 | 4.8 | 0.010541 | -2.574 | 0.6901 | 180 | -2.8694 | 0.68944 | 75 | 1.5093 | 0.48885 |  |  | 1.1645 | 1.0992 |  |  | 0.036754 |  |  | 0.80005
| 187.5 |  | 4.9362 | 4.78 | 0.010541 | -3.1648 | 0.68879 | 195 | -3.2481 | 0.68891 | 90 | 1.2762 | 0.45547 | 0.5366 |  | 1.0382 | 0.8878 | 0.42885 |  | 0.026176 | 0.19501 |  | 0.69874 |  | 0.38577
| 202.5 |  | 4.9344 | 4.77 | 0.0094281 | -3.3315 | 0.68904 | 210 |  |  | 105 | 1.1231 | 0.49171 | 0.48784 |  | 1.0072 | 0.67114 | 0.35971 |  | 0.12933 | 0.11352 |  | 0.6319 |  | 0.40492 | [10 empty] | 0.024095 | 0.79156 | 0.34393 | 0.38168 |  | 0.78331
| 217.5 |  | 4.9563 | 4.726 | 0.012649 | -4.647 | 0.68599 | 225 |  |  | 120 | 1.1294 | 0.2411 | 0.36418 |  | 0.84433 | 0.96484 | 0.49593 |  | 0.057912 | 0.18661 |  | 0.67986 |  | 0.37939 |  | 0.96557 |  |  | 0.99551 | 0.0007437 | 0.82955 | 0.45661 |  | 0.052765 | 0.19049 | 1.0376 | 0.38219 | 0.47266 |  | 0.86213
[8 empty] | 240 |  |  | 135 | 1.1383 | 0.011164 | 0.35558 |  | 0.65833 | 1.2578 | 0.51259 |  | 0.27611 | 0.36631 |  | 0.69348 |  | 0.22837 |  | 0.52353 |  |  | 0.53987 | 0.53522 | 0.57504 | 0.088798 |  | 0.17282 | 0.1022 | 0.94941 | 0.11563 | 0.3873 |  | 0.62891
[8 empty] | 255 |  |  | 150 | 1.1058 | 0.062258 | 0.21242 |  | 0.49843 | 1.3107 | 0.64497 |  | 0.42785 | 0.2751 |  | 0.58959 |  | 0.22645 |  | 0.51356 |  |  | 0.52009 | 0.52414 | 0.62883 | 0.21755 |  | 0.028296 | 0.10071 | 0.9402 | 0.057737 | 0.26862 |  | 0.50408
[8 empty] | 270 |  |  | 165 | 0.95895 | 0.19453 | 0.011799 |  | 0.2051 | 1.3075 | 0.77588 |  | 0.62799 | 0.14593 |  | 0.39623 |  | 0.18666 |  | 0.2908 |  |  | 0.11068 | 0.84913 | 0.31839 | 0.10402 |  | 0.10605 | 0.20019 | 0.94891 | 0.058778 | 0.1669 |  | 0.35923
[8 empty] | 285 |  |  | 180 | 0.53861 | 0.62836 | 0.43678 |  | 0.19285 | 1.3521 | 0.91005 |  | 0.7004 | 0.050734 |  | 0.36419 |  | 0.24635 |  | 0.35991 |  |  | 0.40563 | 1.0057 | 0.18691 | 0.096518 |  | 0.1809 | 0.68725 | 1.1596 | 0.23292 | 0.23684 |  | 0.49096
[8 empty] | 300 |  |  | 195 | 0.21201 | 0.88451 | 0.74984 |  | 0.49876 | 1.2947 | 0.99561 |  | 0.75318 | 0.072601 |  | 0.26862 |  | 0.27947 |  | 0.49104 |  |  | 0.8276 | 1.1799 | 0.00045369 | 0.069037 |  | 0.25444 | 1.1467 | 1.4187 | 0.51263 | 0.37099 |  | 0.66889
[8 empty] | 315 |  |  | 210 | [5 empty] | 1.0903 | 0.91226 |  | 0.69946 | 0.12917 |  | 0.16695 |  | 0.24904 |  | 0.68359 |  |  |  | 0.7609 | 0.2213 | 0.29334 |  | 0.028708 |  | 1.3064 | 0.54136 | 0.35118 |  | 0.61656
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.3023 | 0.96223 |  | 0.63038 | 0.035552 |  | 0.39082 |  | 0.34411 |  | 1.085 |  |  |  | 0.85941 | 0.30857 | 0.26978 |  | 0.11015 |  | 1.8193 | 0.89245 | 0.71932 |  | 1.1048
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.4212 | 1.0354 |  | 0.72083 | 0.023638 |  | 0.39637 |  | 0.33815 |  | 0.89865 |  |  |  | 1.2059 | 0.087134 | 0.089277 |  | 0.30215 |  | 1.8709 | 0.8687 | 0.71071 |  | 1.0879
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.6426 | 1.1821 |  | 0.87788 | 0.012717 |  | 0.41635 |  | 0.34451 |  | 0.58719 |  |  |  | 1.7265 | 0.20999 | 0.14858 |  | 0.57192 |  | 1.9352 | 0.77444 | 0.63879 |  | 1.02
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.029463 |  | 0.36263 |  | 0.26374 |  | 0.46917 |  |  |  |  | 0.29672 | 0.25515 |  | 0.59401 |  |  | 0.71544 | 0.6252 |  | 0.92352
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.030536 |  | 0.319 |  | 0.28242 |  | 0.54091 |  |  |  |  | 0.26006 | 0.17166 |  | 0.51895 |  |  | 0.7741 | 0.62453 |  | 0.94342
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.17928 |  | 0.14417 |  | 0.27602 |  | 0.5959 |  |  |  |  | 0.39924 | 0.13432 |  | 0.47559 |  |  | 0.92332 | 0.63408 |  | 0.94453
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.24346 |  | 0.27562 |  | 0.43909 |  | 0.82213 |  |  |  |  | 0.56702 | 0.20172 |  | 0.74329 |  |  | 1.292 | 0.88824 |  | 1.3788
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.29777 |  | 0.68631 |  | 0.81487 |  | 0.90557 |  |  |  |  | 0.62563 | 0.19079 |  | 1.1817 |  |  | 1.4157 | 0.94854 |  | 1.8051
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.54663 | [5 empty] | 0.93483 |  |  |  |  | 0.58673 | 0.07634 |  |  |  |  | 1.4151 | 0.75381
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.57249 | [5 empty] | 0.41006 |  |  |  |  | 0.77992 | 0.047776 |  |  |  |  | 1.0071 | 0.39963
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.71114 | [5 empty] | 0.55592 |  |  |  |  | 1.0809 | 0.14621 |  |  |  |  | 1.3961 | 0.60893
[8 empty] | 495 |  |  | 390 | [9 empty] | 1.0409 | [5 empty] | 0.33906 |  |  |  |  | 1.6738 | 0.3131 |  |  |  |  | 1.6533 | 0.56477
[8 empty] | 510 |  |  | 405 | [15 empty] | 0.25145 | [5 empty] | 0.29638 | [5 empty] | 0.46573
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.31793 | [5 empty] | 0.36542 | [5 empty] | 0.57945
[11 empty] | 435 | [27 empty] | 0.48283
450
45
PTB | 55 | 5.0146 | 5.0405 | 4.6904 | 0.010198 | -6.9456 | 0.42 | 60 | -6.8251 | 0.42
| 70 | 5.0146 | 5.0368 | 4.7052 | 0.0089315 | -6.5841 | 0.42 | 75 | -6.3262 | 0.42
| 85 | 5.0146 | 5.0303 | 4.738 | 0.015 | -5.8105 | 0.42 | 90 | -5.5184 | 0.42
| 100 | 5.0146 | 5.0263 | 4.7783 | 0.011672 | -4.9342 | 0.42 | 105 | -4.7709 | 0.42
| 115 | 5.0146 | 5.0237 | 4.8004 | 0.010198 | -4.4443 | 0.42 | 120 | -4.2852 | 0.42
| 130 | 5.0146 | 5.0179 | 4.8188 | 0.010132 | -3.9671 | 0.42 | 135 | -3.8482 | 0.42
| 145 | 5.0146 | 5.0152 | 4.8342 | 0.008297 | -3.6104 | 0.42 | 150 | -3.5298 | 0.42
| 160 | 5.0146 | 5.0146 | 4.8457 | 0.0084348 | -3.3686 | 0.42 | 165 | -3.3538 | 0.42
| 175 | 5.0146 | 5.0159 | 4.8491 | 0.0094931 | -3.3241 | 0.42 | 180 | -3.3713 | 0.42
| 190 | 5.0146 | 5.0183 | 4.8444 | 0.0091652 | -3.4657 | 0.42 | 195 | -3.4456 | 0.42
| 205 | 5.0146 | 5.0201 | 4.8492 | 0.0071728 | -3.4053 | 0.42 | 210 | -3.4643 | 0.42
| 220 | 5.0146 | 5.0221 | 4.8422 | 0.0099802 | -3.5823 | 0.42 | 225 | -3.6467 | 0.42
| 235 | 5.0146 | 5.0311 | 4.8412 | 0.011662 | -3.7754 | 0.42 | 240 | -3.8836 | 0.42
| 250 | 5.0146 | 5.0339 | 4.8275 | 0.0089685 | -4.0999 | 0.42 | 255 | -4.2541 | 0.42
| 265 | 5.0146 | 5.0399 | 4.81 | 0.008165 | -4.5625 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 5.0887 | 4.707 | 0.011743 | -7.5002 | 0.57846 | 45
| 60 |  | 5.0863 | 4.79 | 0.0091766 | -5.8247 | 0.55816 | 60 | -5.8247 | 0.55816
| 75 |  | 5.0829 | 4.808 | 0.0089443 | -5.4076 | 0.55271 | 75 | -5.4076 | 0.55271
| 90 |  | 5.0797 | 4.837 | 0.0097872 | -4.7771 | 0.55151 | 90 | -4.7771 | 0.55151
| 105 |  | 5.082 | 4.868 | 0.01105 | -4.2102 | 0.55222 | 105 | -4.2102 | 0.55222
| 120 |  | 5.0816 | 4.905 | 0.01539 | -3.4745 | 0.55949 | 120 | -3.4745 | 0.55949
| 135 |  | 5.0781 | 4.936 | 0.010463 | -2.7985 | 0.55091 | 135 | -2.7985 | 0.55091
| 150 |  | 5.0767 | 4.953 | 0.0097872 | -2.4368 | 0.5499 | 150 | -2.4368 | 0.5499
| 165 |  | 5.0719 | 4.957 | 0.007327 | -2.2656 | 0.54744 | 165 | -2.2656 | 0.54744
| 180 |  | 5.0758 | 4.962 | 0.01105 | -2.2422 | 0.55172 | 180 | -2.2422 | 0.55172
| 195 |  | 5.0782 | 4.958 | 0.0089443 | -2.3672 | 0.54845 | 195 | -2.3672 | 0.54845
| 210 |  | 5.0766 | 4.947 | 0.011743 | -2.5531 | 0.55275 | 210 | -2.5531 | 0.55275
| 225 | 5.0771 | 5.0771 | 4.947 | 0.0097872 | -2.5617 | 0.54866 | 225 | -2.5617 | 0.54866
| 240 |  | 5.078 | 4.941 | 0.010208 | -2.6981 | 0.55024 | 240 | -2.6981 | 0.55024
| 255 |  | 5.0815 | 4.935 | 0.011002 | -2.8832 | 0.55093 | 255 | -2.8832 | 0.55093
| 270 |  | 5.0823 | 4.93 | 0.01026 | -2.9969 | 0.54983 | 270 | -2.9969 | 0.54983
| 285 |  | 5.0838 | 4.921 | 0.010712 | -3.2025 | 0.55044 | 285 | -3.2025 | 0.55044
| 300 |  | 5.0857 | 4.9095 | 0.014681 | -3.4648 | 0.55745 | 300 | -3.4648 | 0.55745
| 315 |  | 5.0853 | 4.902 | 0.015079 | -3.6047 | 0.55848 | 315 | -3.6047 | 0.55848
| 330 |  | 5.0826 | 4.892 | 0.010563 | -3.7493 | 0.55115 | 330 | -3.7493 | 0.55115
| 345 |  | 5.0866 | 4.876 | 0.0094032 | -4.1395 | 0.55008 | 345 | -4.1395 | 0.55008
| 360 |  | 5.0812 | 4.868 | 0.01105 | -4.1951 | 0.55454 | 360 | -4.1951 | 0.55454
| 375 |  | 5.0858 | 4.86 | 0.01026 | -4.44 | 0.55718 | 375 | -4.44 | 0.55718
| 390 |  | 5.0907 | 4.839 | 0.013727 | -4.9445 | 0.57362 | 390 | -4.9445 | 0.57362
| 405 |  | 5.072 | 4.801 | 0.010208 | -5.3432 | 0.61471 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 4.994 | 4.78 | 0.02 | -4.2859 | 0.6535 | 120 | -4.2859 | 0.6535
| 135 |  | 4.9876 | 4.82 | 0.02 | -3.3608 | 0.66046 | 135 | -3.3608 | 0.66046
| 150 |  | 4.9924 | 4.84 | 0.02 | -3.052 | 0.66188 | 150 | -3.052 | 0.66188
| 165 |  | 4.9886 | 4.86 | 0.02 | -2.5774 | 0.66553 | 165 | -2.5774 | 0.66553
| 180 |  | 4.9913 | 4.87 | 0.04 | -2.4296 | 0.69562 | 180 | -2.4296 | 0.69562
| 195 |  | 4.9983 | 4.88 | 0.02 | -2.3667 | 0.66563 | 195 | -2.3667 | 0.66563
| 210 |  | 4.9985 | 4.86 | 0.02 | -2.7701 | 0.66294 | 210 | -2.7701 | 0.66294
| 225 |  | 5.0033 | 4.86 | 0.02 | -2.8633 | 0.66169 | 225 | -2.8633 | 0.66169
| 240 |  | 4.9991 | 4.86 | 0.02 | -2.7834 | 0.66276 | 240 | -2.7834 | 0.66276
| 255 |  | 4.9937 | 4.86 | 0.02 | -2.6773 | 0.66418 | 255 | -2.6773 | 0.66418
| 270 | 4.998 | 4.9952 | 4.86 | 0.02 | -2.7062 | 0.66379 | 270 | -2.7062 | 0.66379
| 285 |  | 4.9971 | 4.85 | 0.04 | -2.9427 | 0.69157 | 285 | -2.9427 | 0.69157
| 300 |  | 4.9934 | 4.84 | 0.02 | -3.0726 | 0.6616 | 300 | -3.0726 | 0.6616
| 315 |  | 4.9921 | 4.84 | 0.02 | -3.0472 | 0.66194 | 315 | -3.0472 | 0.66194
| 330 |  | 4.9866 | 4.83 | 0 | -3.141 | 0.65185 | 330 | -3.141 | 0.65185
| 345 |  | 4.9983 | 4.82 | 0.02 | -3.5671 | 0.65769 | 345 | -3.5671 | 0.65769
| 360 |  | 4.9913 | 4.82 | 0.02 | -3.4313 | 0.65952 | 360 | -3.4313 | 0.65952
| 375 |  | 4.9886 | 4.82 | 0.02 | -3.3793 | 0.66022 | 375 | -3.3793 | 0.66022
| 390 |  | 4.9924 | 4.83 | 0.04 | -3.2523 | 0.69026 | 390 | -3.2523 | 0.69026
| 405 |  | 4.9876 | 4.82 | 0.02 | -3.3608 | 0.66046 | 405 | -3.3608 | 0.66046
| 420 |  | 4.994 | 4.83 | 0 | -3.2847 | 0.64992 | 420 | -3.2847 | 0.64992
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
| 105 |  | 5.0675 | 4.8775 | 0.022913 | -3.7496 | 1.1276 | 105 | -3.7496 | 1.1276
| 120 |  | 5.0788 | 4.928 | 0.021176 | -2.9698 | 1.1019 | 120 | -2.9698 | 1.1019
| 135 |  | 5.0799 | 4.9455 | 0.021392 | -2.6464 | 1.0997 | 135 | -2.6464 | 1.0997
| 150 |  | 5.0698 | 4.95 | 0.01777 | -2.3622 | 1.072 | 150 | -2.3622 | 1.072
| 165 |  | 5.0721 | 4.961 | 0.015183 | -2.1907 | 1.0533 | 165 | -2.1907 | 1.0533
| 180 |  | 5.0675 | 4.969 | 0.016512 | -1.9436 | 1.0597 | 180 | -1.9436 | 1.0597
| 195 |  | 5.0731 | 4.986 | 0.014654 | -1.7168 | 1.045 | 195 | -1.7168 | 1.045
| 210 |  | 5.0836 | 4.989 | 0.016512 | -1.8606 | 1.0554 | 210 | -1.8606 | 1.0554
| 225 |  | 5.0715 | 4.999 | 0.015183 | -1.4292 | 1.0453 | 225 | -1.4292 | 1.0453
| 240 |  | 5.0675 | 4.987 | 0.016255 | -1.5887 | 1.0543 | 240 | -1.5887 | 1.0543
| 255 | 5.0635 | 5.0694 | 4.973 | 0.013416 | -1.9024 | 1.041 | 255 | -1.9024 | 1.041
| 270 |  | 5.0766 | 4.971 | 0.016512 | -2.0804 | 1.0593 | 270 | -2.0804 | 1.0593
| 285 |  | 5.0716 | 4.959 | 0.013727 | -2.2194 | 1.0456 | 285 | -2.2194 | 1.0456
| 300 |  | 5.0766 | 4.961 | 0.016512 | -2.2774 | 1.0614 | 300 | -2.2774 | 1.0614
| 315 |  | 5.0694 | 4.97 | 0.013765 | -1.9616 | 1.0435 | 315 | -1.9616 | 1.0435
| 330 |  | 5.0675 | 4.969 | 0.015183 | -1.944 | 1.0516 | 330 | -1.944 | 1.0516
| 345 |  | 5.0715 | 4.954 | 0.01729 | -2.3165 | 1.0679 | 345 | -2.3165 | 1.0679
| 360 |  | 5.0836 | 4.9375 | 0.020487 | -2.8737 | 1.0944 | 360 | -2.8737 | 1.0944
| 375 |  | 5.0731 | 4.9395 | 0.017313 | -2.6334 | 1.0712 | 375 | -2.6334 | 1.0712
| 390 |  | 5.0675 | 4.926 | 0.016983 | -2.7921 | 1.072 | 390 | -2.7921 | 1.072
| 405 |  | 5.0721 | 4.919 | 0.019974 | -3.0187 | 1.0946 | 405 | -3.0187 | 1.0946
| 420 |  | 5.0698 | 4.925 | 0.019331 | -2.8554 | 1.0885 | 420 | -2.8554 | 1.0885
| 435 |  | 5.0799 | 4.924 | 0.020622 | -3.0697 | 1.0984 | 435 | -3.0697 | 1.0984
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
| 80 |  | 5.0623 | 4.784 | 0.0091538 | -5.4981 | 0.98918 | 90 | -5.016 | 0.98972
| 95 |  | 5.0563 | 4.8149 | 0.0086089 | -4.7749 | 0.98999 | 105 | -4.3493 | 0.99019
| 110 |  | 5.0533 | 4.8443 | 0.0090358 | -4.1365 | 0.99029 | 120 | -3.7039 | 0.99052
| 125 |  | 5.0503 | 4.8742 | 0.010082 | -3.4876 | 0.99064 | 135 | -3.2473 | 0.99056
| 140 |  | 5.0503 | 4.8924 | 0.011357 | -3.1272 | 0.99052 | 150 | -2.7738 | 0.99055
| 155 |  | 5.0493 | 4.9182 | 0.011317 | -2.597 | 0.99056 | 165 | -2.4442 | 0.99066
| 170 |  | 5.0483 | 4.9288 | 0.0099391 | -2.3678 | 0.9907 | 180 | -2.3044 | 0.99069
| 185 |  | 5.0483 | 4.9336 | 0.0093414 | -2.2727 | 0.99068 | 195 | -2.2783 | 0.99081
| 200 |  | 5.0473 | 4.9322 | 0.0097733 | -2.2811 | 0.99088 | 210 | -2.3947 | 0.99091
| 215 |  | 5.0473 | 4.9236 | 0.012082 | -2.4515 | 0.99092 | 225 | -2.5182 | 0.9912
| 230 |  | 5.0453 | 4.9166 | 0.013177 | -2.5515 | 0.99134 | 240 | -2.6691 | 0.99137
| 245 |  | 5.0453 | 4.9077 | 0.014774 | -2.7279 | 0.99138 | 255 | -2.8676 | 0.99129
| 260 |  | 5.0463 | 4.8981 | 0.01418 | -2.9374 | 0.99125 | 270 | -3.033 | 0.99141
| 275 |  | 5.0453 | 4.8899 | 0.0097794 | -3.0807 | 0.99148 | 285 | -3.1651 | 0.99145
| 290 |  | 5.0458 | 4.884 | 0.0091929 | -3.2073 | 0.99143 | 300 | -3.2445 | 0.9915
| 305 |  | 5.0453 | 4.8807 | 0.0099834 | -3.2631 | 0.99154 | 315 | -3.3054 | 0.99156
| 320 |  | 5.0453 | 4.8775 | 0.0096959 | -3.3265 | 0.99156 | 330 | -3.3842 | 0.99145
| 335 |  | 5.0463 | 4.8741 | 0.0095931 | -3.413 | 0.9914 | 345 | -3.4696 | 0.99135
| 350 |  | 5.0468 | 4.8703 | 0.0063158 | -3.4979 | 0.99133 | 360 | -3.4924 | 0.99126
| 365 |  | 5.0473 | 4.8712 | 0.011714 | -3.4896 | 0.99123 | 375 | -3.5662 | 0.99125
| 380 |  | 5.0473 | 4.8654 | 0.0084187 | -3.6045 | 0.99127 | 390 | -3.6576 | 0.99135
| 395 |  | 5.0468 | 4.8609 | 0.0099707 | -3.6842 | 0.99139 | 405 | -3.7399 | 0.99148
| 410 |  | 5.0463 | 4.8562 | 0.0092755 | -3.7678 | 0.99152 | 420 | -3.7501 | 0.99138
| 425 |  | 5.0473 | 4.8585 | 0.0098929 | -3.7412 | 0.99132 | 435 | -3.8186 | 0.99154
| 440 |  | 5.0458 | 4.8512 | 0.0094297 | -3.8573 | 0.99165 | 450 | -3.8668 | 0.99172
| 455 |  | 5.0453 | 4.85 | 0.0080067 | -3.8716 | 0.99176 | 465 | -3.8962 | 0.99164
| 470 |  | 5.0463 | 4.8491 | 0.0069179 | -3.9084 | 0.99157 | 480 | -3.9301 | 0.99171
| 485 |  | 5.0453 | 4.8465 | 0.0076391 | -3.9409 | 0.99178 | 495 | -3.97 | 0.9918
| 500 | 5.045 | 5.0453 | 4.8443 | 0.0090221 | -3.9845 | 0.9918 | 510 | -3.9888 | 0.99187
| 515 |  | 5.0448 | 4.8435 | 0.010213 | -3.9909 | 0.9919 | 525 | -4.0279 | 0.99192
| 530 |  | 5.0448 | 4.8407 | 0.009985 | -4.0464 | 0.99192
| 545 |  | 5.0423 | 4.8397 | 0.010003 | -4.0186 | 0.9924
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 42 |  | 5.199 | 4.651 | 0.01 | -10.54 | 0.83 | 45
| 57 |  | 5.201 | 4.807 | 0.014 | -7.575 | 0.83 | 60 | -7.3484 | 0.83
LEI | 72 |  | 5.2 | 4.865 | 0.011 | -6.442 | 0.83 | 75 | -6.288 | 0.83
| 87 |  | 5.201 | 4.906 | 0.015 | -5.672 | 0.83 | 90 | -5.5456 | 0.83
| 102 |  | 5.198 | 4.936 | 0.013 | -5.04 | 0.83 | 105 | -4.9054 | 0.83
| 117 |  | 5.198 | 4.971 | 0.01 | -4.367 | 0.83 | 120 | -4.225 | 0.83
| 132 |  | 5.195 | 5.005 | 0.011 | -3.657 | 0.83 | 135 | -3.561 | 0.83
| 147 |  | 5.193 | 5.028 | 0.01 | -3.177 | 0.83 | 150 | -3.0848 | 0.83
| 162 |  | 5.192 | 5.051 | 0.01 | -2.716 | 0.83 | 165 | -2.7006 | 0.83
| 177 |  | 5.191 | 5.054 | 0.013 | -2.639 | 0.83 | 180 | -2.6428 | 0.83
| 192 |  | 5.191 | 5.053 | 0.01 | -2.658 | 0.83 | 195 | -2.6622 | 0.83
| 207 |  | 5.188 | 5.049 | 0.01 | -2.679 | 0.83 | 210 | -2.7368 | 0.83
| 222 |  | 5.188 | 5.034 | 0.012 | -2.968 | 0.83 | 225 | -2.991 | 0.83
| 237 |  | 5.189 | 5.029 | 0.012 | -3.083 | 0.83 | 240 | -3.1338 | 0.83
| 252 |  | 5.185 | 5.012 | 0.012 | -3.337 | 0.83 | 255 | -3.341 | 0.83
| 267 |  | 5.183 | 5.009 | 0.01 | -3.357 | 0.83 | 270 | -3.3954 | 0.83
| 282 |  | 5.184 | 5 | 0.009 | -3.549 | 0.83 | 285 | -3.5532 | 0.83
| 297 |  | 5.182 | 4.997 | 0.012 | -3.57 | 0.83 | 300 | -3.6238 | 0.83
| 312 |  | 5.184 | 4.985 | 0.013 | -3.839 | 0.83 | 315 | -3.9088 | 0.83
| 327 |  | 5.181 | 4.964 | 0.013 | -4.188 | 0.83 | 330 | -4.504 | 0.83
| 342 |  | 5.184 | 4.885 | 0.166 | -5.768 | 0.83 | 345
| 357 |  | 5.184 | 4.797 | 0.026 | -7.465 | 0.83 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 8
| 8 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.14697
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 7.9046 | 7.558 | 0.010328 | -4.3846 | 0.6452 | 45
| 52.5 |  | 7.9084 | 7.623 | 0.0094868 | -3.6092 | 0.64488 | 60 | -3.4439 | 0.64501 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 7.9052 | 7.646 | 0.0096609 | -3.2786 | 0.64515 | 75 | -3.0342 | 0.65124 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 7.9107 | 7.69 | 0.010541 | -2.7898 | 0.65734 | 90 | -2.4039 | 0.66379 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 7.9076 | 7.748 | 0.021499 | -2.0179 | 0.67024 | 105 | -1.8435 | 0.66357 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 7.9161 | 7.784 | 0.0084327 | -1.669 | 0.65689 | 120 | -1.3053 | 0.65737 |  | 0.88349 | 0.49974 | 2.1575 |  | 0.245 | 0.38684 | 1.6701 |  | 1.1923 | 1.8511 |  | 0.77577 |  | 2.4395 |  | 0.93659 | 0.78307 | 0.46151 | 0.65787 | 0.13465 | 0.20286 | 1.557 |  | 0.92065 | 0.26909 | 1.2039 | 0.79399 | 2.4358 |  | 0.027449 | 0.68697 | 0.42892 | 0.52888 | 0.063718 |  | 0.7746
| 127.5 | 7.9039 | 7.9044 | 7.83 | 0.010541 | -0.94173 | 0.65786 | 135 | -0.99267 | 0.65744 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 7.9146 | 7.832 | 0.010328 | -1.0436 | 0.65701 | 150 | -0.78395 | 0.65773 | 45
| 157.5 |  | 7.8974 | 7.856 | 0.014298 | -0.52428 | 0.65844 | 165 | -0.72135 | 0.65798 | 60 | 1.8571 | 1.1744 |  |  | 1.0771 | 0.64036 |  |  | 0.59494 |  |  | 0.01642
| 172.5 |  | 7.9086 | 7.836 | 0.0084327 | -0.91841 | 0.65751 | 180 | -1.1014 | 0.65739 | 75 | 1.8553 | 1.1676 |  |  | 0.6868 | 0.65541 |  |  | 1.0607 |  |  | 0.43225 | [12 empty] | 3.8469 | 2.7572 | 3.0536 |  |  | 3.2354
| 187.5 |  | 7.9116 | 7.81 | 0.010541 | -1.2843 | 0.65726 | 195 | -1.5911 | 0.65741 | 90 | 1.83 | 1.319 | 1.1653 |  | 0.83894 | 0.44555 | 0.50359 |  | 0.86699 | 0.08692 |  | 0.42613 |  | 0.32058 |  |  | 1.5183 | [7 empty] | 2.334 | 0.9719 | 1.245 | 1.2541 |  | 1.5443 | 0.67129 | 1.2474 | 1.1174 | 1.0822 |  | 0.97027
| 202.5 |  | 7.9081 | 7.758 | 0.011353 | -1.8978 | 0.65756 | 210 |  |  | 105 | 1.6755 | 0.93601 | 1.015 |  | 0.74013 | 0.7545 | 0.52848 |  | 0.83093 | 0.1474 |  | 0.14209 |  | 0.26965 |  |  | 1.0263 | [7 empty] | 1.2429 | 0.18233 | 0.43306 | 0.27581 |  | 0.53089 | 0.5761 | 1.1045 | 0.89334 | 0.93394 |  | 0.83978
| 217.5 |  | 7.9207 | 7.69 | 0.010541 | -2.913 | 0.64388 | 225 |  |  | 120 | 1.6792 | 0.90351 | 0.83868 |  | 0.61201 | 0.79431 | 0.73808 |  | 0.98057 | 0.011891 |  | 0.24747 |  | 0.22128 |  | 0.67285 | 0.70081 | 0.9382 | 1.4986 | 0.041888 | 0.70751 | 0.67327 |  | 0.88631 | 0.79344 | 0.73115 | 0.03362 | 0.020838 |  | 0.19397 | 0.41694 | 0.93978 | 0.71961 | 0.71075 |  | 0.63357
[8 empty] | 240 |  |  | 135 | 1.4663 | 0.65069 | 0.62332 |  | 0.15231 | 0.85989 | 0.7751 |  | 1.3118 | 0.011243 |  | 0.49438 |  | 0.47538 |  | 0.27571 | 0.32054 | 0.41948 | 0.95236 | 0.38535 | 0.37 | 0.33661 |  | 0.80856 | 0.64498 | 0.67905 | 0.060493 | 0.046969 |  | 0.50286 | 0.087873 | 0.53922 | 0.30403 | 0.30553 |  | 0.14176
[8 empty] | 255 |  |  | 150 | 1.2665 | 0.37005 | 0.49723 |  | 0.15187 | 0.98473 | 0.71899 |  | 1.4714 | 0.16419 |  | 0.5423 |  | 0.66024 |  | 0.39918 | 0.26241 | 0.12455 | 0.70355 | 0.48197 | 0.38878 | 0.20879 |  | 0.87068 | 0.29115 | 0.91377 | 0.050172 | 0.19592 |  | 0.44612 | 0.36699 | 0.01461 | 0.24866 | 0.19475 |  | 0.4212
[8 empty] | 270 |  |  | 165 | 1.1607 | 0.30951 | 0.34279 |  | 0.36526 | 0.93151 | 0.79443 |  | 1.6184 | 0.055411 |  | 0.7129 |  | 0.72304 |  | 0.13956 | 0.40088 | 0.44991 | 0.16721 | 1.0186 | 0.13945 | 0.18356 |  | 0.54817 | 0.29784 | 0.8229 | 0.011422 | 0.040449 |  | 0.67104 | 0.50566 | 0.15906 | 0.40813 | 0.38773 |  | 0.63546
[8 empty] | 285 |  |  | 180 | 0.70692 | 0.17843 | 0.1104 |  | 0.82629 | 1.0206 | 0.86632 |  | 1.7049 | 0.063299 |  | 0.72081 |  | 0.73709 |  | 0.017547 | 0.52283 | 0.51865 | 0.29046 | 1.0807 | 0.13397 | 0.18527 |  | 0.55351 | 0.29894 | 1.0558 | 0.14786 | 0.19704 |  | 0.5193 | 0.41594 | 0.20702 | 0.47922 | 0.45581 |  | 0.70819
[8 empty] | 300 |  |  | 195 | 0.1994 | 0.663 | 0.63015 |  | 1.1591 | 1.0452 | 0.96249 |  | 1.5848 | 0.0072433 |  | 0.59518 |  | 0.55206 |  | 0.38065 | 0.65589 | 0.52637 | 0.93112 | 1.3298 | 0.33249 | 0.30432 |  | 0.25669 | 1.2598 | 1.6855 | 0.7168 | 0.66841 |  | 0.12715 | 0.20306 | 0.14575 | 0.42419 | 0.4235 |  | 0.61242
[8 empty] | 315 |  |  | 210 | [5 empty] | 0.93574 | 1.0555 |  | 1.455 | 0.18348 |  | 0.56539 |  | 0.36136 |  | 0.52029 | 0.70677 | 0.52551 |  | 1.3011 | 0.40853 | 0.21099 |  | 0.1532 |  | 1.8166 | 0.9385 | 0.72145 |  | 0.3677 |  | 0.14913 | 0.39837 | 0.45392 |  | 0.57731
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.1012 | 1.1718 |  | 1.4015 | 0.15154 |  | 0.38002 |  | 0.21681 |  | 0.44694 | 0.591 | 0.43729 |  | 1.3868 | 0.34635 | 0.18233 |  | 0.037301 |  | 1.8489 | 0.80948 | 0.62343 |  | 0.40424 |  | 0.03501 | 0.32954 | 0.37543 |  | 0.44931
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.3758 | 1.3217 |  | 1.5373 | 0.059038 |  | 0.27854 |  | 0.20702 |  | 0.53573 | 0.67111 | 0.48813 |  | 1.7758 | 0.45563 | 0.37048 |  | 0.15598 |  | 2.2782 | 0.99508 | 0.88378 |  | 0.67243 |  | 0.021737 | 0.34683 | 0.36318 |  | 0.43371
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.5849 | 1.5957 |  | 1.6259 | 0.1373 |  | 0.19029 |  | 0.051733 |  | 0.59556 | 0.63895 | 0.43606 |  | 2.1637 | 0.63597 | 0.46456 |  | 0.4057 |  | 2.7366 | 1.2363 | 1.0367 |  | 0.9727 |  | 0.18759 | 0.23766 | 0.27959 |  | 0.29691
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.09785 |  | 0.021767 |  | 0.11087 |  | 0.65063 | 0.55988 | 0.33741 |  |  | 0.34757 | 0.23308 |  | 0.34143 |  |  | 1.0243 | 0.87295 |  | 0.97 |  |  | 0.22908 | 0.25859 |  | 0.22008
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.13503 |  | 0.15171 |  | 0.017658 |  | 0.26223 | 0.41276 | 0.32279 |  |  | 0.3142 | 0.16746 |  | 0.14742 |  |  | 0.57846 | 0.42262 |  | 0.40008 |  |  | 0.22461 | 0.26603 |  | 0.27166
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.2261 |  | 0.26493 |  | 0.039795 |  | 0.31263 | 0.43937 | 0.33096 |  |  | 0.62492 | 0.37502 |  | 0.33024 |  |  | 0.94001 | 0.67911 |  | 0.63164 |  |  | 0.13235 | 0.20332 |  | 0.21666
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.23008 |  | 0.095919 |  | 0.12245 |  | 0.76085 | 0.46228 | 0.20233 |  |  | 0.34675 | 0.10529 |  | 0.22804 |  |  | 1.1285 | 0.85107 |  | 0.95912 |  |  | 0.094364 | 0.16679 |  | 0.12445
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.4059 |  | 0.21309 |  | 0.5756 |  | 0.95406 |  |  |  |  | 0.42682 | 0.012822 |  | 0.59696 |  |  | 1.4054 | 0.95248 |  | 1.5004
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.50488 | [5 empty] | 0.96795 |  |  |  |  | 0.51896 | 0.004161 |  |  |  |  | 1.5053 | 0.95706
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.64327 | [5 empty] | 0.75619 |  |  |  |  | 0.88972 | 0.22058 |  |  |  |  | 1.6262 | 0.95428
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.99442 | [5 empty] | 0.42459 |  |  |  |  | 1.2529 | 0.23613 |  |  |  |  | 1.6583 | 0.6512
[8 empty] | 495 |  |  | 390 | [9 empty] | 1.3398 | [5 empty] | 0.30091 |  |  |  |  | 1.5663 | 0.20468 |  |  |  |  | 1.7894 | 0.49306
[8 empty] | 510 |  |  | 405 | [15 empty] | 0.11621 | [5 empty] | 0.1917 | [5 empty] | 0.29464
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.09626 | [5 empty] | 0.39084 | [5 empty] | 0.4638
[11 empty] | 435 | [27 empty] | 0.35881
450
45
PTB | 55 | 8.0052 | 8.045 | 7.6424 | 0.01052 | -5.0042 | 0.42 | 60 | -4.8992 | 0.42
| 70 | 8.0052 | 8.0376 | 7.6607 | 0.01035 | -4.6891 | 0.42 | 75 | -4.4943 | 0.41667
| 85 | 8.0052 | 8.0304 | 7.7007 | 0.008738 | -4.1047 | 0.41 | 90 | -3.8599 | 0.41333
| 100 | 8.0052 | 8.0245 | 7.7541 | 0.010834 | -3.3703 | 0.42 | 105 | -3.1821 | 0.42
| 115 | 8.0052 | 8.0194 | 7.7944 | 0.010832 | -2.8057 | 0.42 | 120 | -2.6383 | 0.42
| 130 | 8.0052 | 8.0134 | 7.8288 | 0.010132 | -2.3036 | 0.42 | 135 | -2.1567 | 0.42
| 145 | 8.0052 | 8.0068 | 7.8576 | 0.014799 | -1.8629 | 0.42 | 150 | -1.7897 | 0.42
| 160 | 8.0052 | 8.0052 | 7.8736 | 0.014325 | -1.6434 | 0.42 | 165 | -1.6433 | 0.42
| 175 | 8.0052 | 8.0064 | 7.8748 | 0.012288 | -1.6432 | 0.42 | 180 | -1.6625 | 0.42
| 190 | 8.0052 | 8.0104 | 7.8742 | 0.012129 | -1.7012 | 0.42 | 195 | -1.7494 | 0.42
| 205 | 8.0052 | 8.0147 | 7.8668 | 0.016729 | -1.8457 | 0.42 | 210 | -1.9084 | 0.42
| 220 | 8.0052 | 8.0179 | 7.8548 | 0.014754 | -2.0337 | 0.42 | 225 | -2.1427 | 0.42
| 235 | 8.0052 | 8.0291 | 7.8396 | 0.0087792 | -2.3606 | 0.42 | 240 | -2.4569 | 0.42
| 250 | 8.0052 | 8.0337 | 7.8208 | 0.011001 | -2.6495 | 0.42 | 255 | -2.8592 | 0.42
| 265 | 8.0052 | 8.0463 | 7.7825 | 0.013593 | -3.2787 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 8.1037 | 7.6395 | 0.013945 | -5.728 | 0.55479 | 45
| 60 |  | 8.0973 | 7.737 | 0.0097872 | -4.4494 | 0.54341 | 60 | -4.4494 | 0.54341
| 75 |  | 8.0948 | 7.768 | 0.010052 | -4.0369 | 0.54013 | 75 | -4.0369 | 0.54013
| 90 |  | 8.1016 | 7.814 | 0.013139 | -3.5497 | 0.54071 | 90 | -3.5497 | 0.54071
| 105 |  | 8.097 | 7.882 | 0.010052 | -2.6551 | 0.53846 | 105 | -2.6551 | 0.53846
| 120 |  | 8.0988 | 7.93 | 0.0091766 | -2.084 | 0.5376 | 120 | -2.084 | 0.5376
| 135 |  | 8.0969 | 7.971 | 0.013727 | -1.5547 | 0.54052 | 135 | -1.5547 | 0.54052
| 150 |  | 8.0953 | 8.006 | 0.0082078 | -1.1029 | 0.53714 | 150 | -1.1029 | 0.53714
| 165 |  | 8.0946 | 8.0145 | 0.016694 | -0.9893 | 0.54312 | 165 | -0.9893 | 0.54312
| 180 |  | 8.0906 | 8.014 | 0.014654 | -0.94715 | 0.54159 | 180 | -0.94715 | 0.54159
| 195 |  | 8.0864 | 8.004 | 0.0094032 | -1.0194 | 0.53829 | 195 | -1.0194 | 0.53829
| 210 |  | 8.0935 | 7.992 | 0.01105 | -1.2545 | 0.53882 | 210 | -1.2545 | 0.53882
| 225 | 8.0901 | 8.0901 | 7.979 | 0.012096 | -1.373 | 0.53898 | 225 | -1.373 | 0.53898
| 240 |  | 8.091 | 7.97 | 0.011239 | -1.4952 | 0.53908 | 240 | -1.4952 | 0.53908
| 255 |  | 8.098 | 7.956 | 0.0094032 | -1.7533 | 0.53752 | 255 | -1.7533 | 0.53752
| 270 |  | 8.0931 | 7.945 | 0.011002 | -1.8297 | 0.53874 | 270 | -1.8297 | 0.53874
| 285 |  | 8.0952 | 7.932 | 0.012814 | -2.0158 | 0.53987 | 285 | -2.0158 | 0.53987
| 300 |  | 8.0958 | 7.915 | 0.011002 | -2.233 | 0.53868 | 300 | -2.233 | 0.53868
| 315 |  | 8.0967 | 7.905 | 0.0088852 | -2.3674 | 0.5374 | 315 | -2.3674 | 0.5374
| 330 |  | 8.0978 | 7.886 | 0.012312 | -2.6153 | 0.53964 | 330 | -2.6153 | 0.53964
| 345 |  | 8.0997 | 7.873 | 0.007327 | -2.7986 | 0.53696 | 345 | -2.7986 | 0.53696
| 360 |  | 8.101 | 7.852 | 0.018238 | -3.0735 | 0.54557 | 360 | -3.0735 | 0.54557
| 375 |  | 8.1084 | 7.831 | 0.010208 | -3.4209 | 0.54016 | 375 | -3.4209 | 0.54016
| 390 |  | 8.106 | 7.803 | 0.013416 | -3.7377 | 0.54807 | 390 | -3.7377 | 0.54807
| 405 |  | 8.0933 | 7.732 | 0.011965 | -4.464 | 0.56527 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 7.9729 | 7.76 | 0.04 | -2.6701 | 0.61279 | 120 | -2.6701 | 0.61279
| 135 |  | 7.9685 | 7.82 | 0.04 | -1.8633 | 0.61793 | 135 | -1.8633 | 0.61793
| 150 |  | 7.9736 | 7.86 | 0.03 | -1.4248 | 0.61269 | 150 | -1.4248 | 0.61269
| 165 |  | 7.9696 | 7.9 | 0.02 | -0.87348 | 0.61098 | 165 | -0.87348 | 0.61098
| 180 |  | 7.9766 | 7.91 | 0.04 | -0.83478 | 0.62341 | 180 | -0.83478 | 0.62341
| 195 |  | 7.9794 | 7.92 | 0.02 | -0.74424 | 0.61101 | 195 | -0.74424 | 0.61101
| 210 |  | 7.9829 | 7.91 | 0.04 | -0.91281 | 0.62246 | 210 | -0.91281 | 0.62246
| 225 |  | 7.9967 | 7.91 | 0.04 | -1.0839 | 0.62037 | 225 | -1.0839 | 0.62037
| 240 |  | 7.9894 | 7.9 | 0.02 | -1.119 | 0.60797 | 240 | -1.119 | 0.60797
| 255 |  | 7.978 | 7.88 | 0.02 | -1.2287 | 0.60818 | 255 | -1.2287 | 0.60818
| 270 | 7.982 | 7.983 | 7.86 | 0.033333 | -1.5414 | 0.61352 | 270 | -1.5414 | 0.61352
| 285 |  | 7.98 | 7.84 | 0.04 | -1.7539 | 0.61769 | 285 | -1.7539 | 0.61769
| 300 |  | 7.9762 | 7.84 | 0.05 | -1.708 | 0.62771 | 300 | -1.708 | 0.62771
| 315 |  | 7.9761 | 7.81 | 0.02 | -2.0827 | 0.60314 | 315 | -2.0827 | 0.60314
| 330 |  | 7.9705 | 7.79 | 0.02 | -2.2644 | 0.60246 | 330 | -2.2644 | 0.60246
| 345 |  | 7.9794 | 7.79 | 0.02 | -2.3734 | 0.60113 | 345 | -2.3734 | 0.60113
| 360 |  | 7.9766 | 7.79 | 0.02 | -2.3392 | 0.60154 | 360 | -2.3392 | 0.60154
| 375 |  | 7.9696 | 7.78 | 0.04 | -2.3792 | 0.61477 | 375 | -2.3792 | 0.61477
| 390 |  | 7.9736 | 7.78 | 0.04 | -2.4281 | 0.61417 | 390 | -2.4281 | 0.61417
| 405 |  | 7.9685 | 7.77 | 0.02 | -2.4907 | 0.60124 | 405 | -2.4907 | 0.60124
| 420 |  | 7.9729 | 7.78 | 0.04 | -2.4192 | 0.61428 | 420 | -2.4192 | 0.61428
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
| 75 |  | 8.0629 | 7.508 | 0.025047 | -6.8821 | 0.74484 | 75 | -6.8821 | 0.74484
| 90 |  | 8.0541 | 7.68 | 0.01451 | -4.6447 | 0.6779 | 90 | -4.6447 | 0.6779
| 105 |  | 8.0673 | 7.8225 | 0.01713 | -3.0346 | 0.67565 | 105 | -3.0346 | 0.67565
| 120 |  | 8.0545 | 7.889 | 0.015183 | -2.0549 | 0.66237 | 120 | -2.0549 | 0.66237
| 135 |  | 8.0484 | 7.919 | 0.01889 | -1.6077 | 0.67495 | 135 | -1.6077 | 0.67495
| 150 |  | 8.0402 | 7.955 | 0.017014 | -1.0594 | 0.66393 | 150 | -1.0594 | 0.66393
| 165 |  | 8.0535 | 7.973 | 0.011743 | -0.99907 | 0.64418 | 165 | -0.99907 | 0.64418
| 180 |  | 8.0549 | 7.989 | 0.017741 | -0.81859 | 0.66409 | 180 | -0.81859 | 0.66409
| 195 |  | 8.0533 | 8.0215 | 0.019541 | -0.39479 | 0.66924 | 195 | -0.39479 | 0.66924
| 210 |  | 8.0502 | 8.016 | 0.022572 | -0.42545 | 0.68436 | 210 | -0.42545 | 0.68436
| 225 |  | 8.0538 | 7.9995 | 0.016376 | -0.6745 | 0.65771 | 225 | -0.6745 | 0.65771
| 240 |  | 8.0436 | 7.993 | 0.018666 | -0.62887 | 0.66772 | 240 | -0.62887 | 0.66772
| 255 | 8.0335 | 8.0454 | 7.99 | 0.015894 | -0.68851 | 0.65664 | 255 | -0.68851 | 0.65664
| 270 |  | 8.0436 | 7.967 | 0.013416 | -0.9521 | 0.64979 | 270 | -0.9521 | 0.64979
| 285 |  | 8.0538 | 7.932 | 0.017045 | -1.5126 | 0.66598 | 285 | -1.5126 | 0.66598
| 300 |  | 8.0502 | 7.936 | 0.016026 | -1.4192 | 0.66161 | 300 | -1.4192 | 0.66161
| 315 |  | 8.0533 | 7.941 | 0.015183 | -1.3944 | 0.65803 | 315 | -1.3944 | 0.65803
| 330 |  | 8.0549 | 7.942 | 0.015079 | -1.4021 | 0.65757 | 330 | -1.4021 | 0.65757
| 345 |  | 8.0535 | 7.933 | 0.016255 | -1.4958 | 0.66275 | 345 | -1.4958 | 0.66275
| 360 |  | 8.0402 | 7.9075 | 0.01713 | -1.6502 | 0.66839 | 360 | -1.6502 | 0.66839
| 375 |  | 8.0484 | 7.888 | 0.012814 | -1.9929 | 0.65436 | 375 | -1.9929 | 0.65436
| 390 |  | 8.0545 | 7.8815 | 0.020072 | -2.148 | 0.68361 | 390 | -2.148 | 0.68361
| 405 |  | 8.0673 | 7.875 | 0.019331 | -2.3838 | 0.68072 | 405 | -2.3838 | 0.68072
| 420 |  | 8.0541 | 7.8665 | 0.021343 | -2.3291 | 0.69109 | 420 | -2.3291 | 0.69109
| 435 |  | 8.0629 | 7.8645 | 0.02305 | -2.4606 | 0.70007 | 435 | -2.4606 | 0.70007
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
| 80 |  | 8.0587 | 7.744 | 0.011297 | -3.9047 | 0.62092 | 90 | -3.477 | 0.62114
| 95 |  | 8.0527 | 7.7899 | 0.012911 | -3.2631 | 0.62124 | 105 | -2.7782 | 0.62141
| 110 |  | 8.0477 | 7.8436 | 0.018005 | -2.5357 | 0.6215 | 120 | -2.0741 | 0.62164
| 125 |  | 8.0437 | 7.8954 | 0.012628 | -1.8433 | 0.62171 | 135 | -1.5641 | 0.62174
| 140 |  | 8.0427 | 7.9281 | 0.01209 | -1.4245 | 0.62175 | 150 | -1.2399 | 0.62189
| 155 |  | 8.0397 | 7.9474 | 0.01254 | -1.1476 | 0.62196 | 165 | -1.0358 | 0.622
| 170 |  | 8.0387 | 7.9599 | 0.010007 | -0.97985 | 0.62202 | 180 | -1.0002 | 0.62177
| 185 |  | 8.0437 | 7.9624 | 0.012048 | -1.0103 | 0.62164 | 195 | -1.0133 | 0.62195
| 200 |  | 8.0377 | 7.9561 | 0.010098 | -1.0148 | 0.6221 | 210 | -1.1011 | 0.62211
| 215 |  | 8.0377 | 7.9457 | 0.012189 | -1.1442 | 0.62211 | 225 | -1.2463 | 0.62217
| 230 |  | 8.0367 | 7.9324 | 0.0094412 | -1.2974 | 0.6222 | 240 | -1.4459 | 0.62221
| 245 |  | 8.0367 | 7.9145 | 0.010168 | -1.5201 | 0.62222 | 255 | -1.6386 | 0.62218
| 260 |  | 8.0377 | 7.9012 | 0.012705 | -1.6978 | 0.62216 | 270 | -1.7479 | 0.62227
| 275 |  | 8.0357 | 7.8932 | 0.010147 | -1.7729 | 0.62232 | 285 | -1.9028 | 0.62221
| 290 |  | 8.0382 | 7.88 | 0.012467 | -1.9677 | 0.62215 | 300 | -2.044 | 0.62216
| 305 |  | 8.0382 | 7.8708 | 0.010898 | -2.0821 | 0.62217 | 315 | -2.1752 | 0.62223
| 320 |  | 8.0372 | 7.8586 | 0.016963 | -2.2218 | 0.62226 | 330 | -2.2757 | 0.62227
| 335 |  | 8.0372 | 7.8521 | 0.017618 | -2.3026 | 0.62227 | 345 | -2.3771 | 0.62223
| 350 |  | 8.0382 | 7.8441 | 0.012971 | -2.4143 | 0.62221 | 360 | -2.5328 | 0.62221
| 365 |  | 8.0387 | 7.8303 | 0.011407 | -2.5921 | 0.6222 | 375 | -2.5886 | 0.62218
| 380 |  | 8.0392 | 7.8312 | 0.01475 | -2.5869 | 0.62216 | 390 | -2.6095 | 0.62222
| 395 |  | 8.0382 | 7.8275 | 0.010561 | -2.6208 | 0.62225 | 405 | -2.659 | 0.62225
| 410 |  | 8.0382 | 7.8229 | 0.012735 | -2.6781 | 0.62226 | 420 | -2.7657 | 0.62222
| 425 |  | 8.0392 | 7.8133 | 0.010256 | -2.8096 | 0.6222 | 435 | -2.8008 | 0.62228
| 440 |  | 8.0377 | 7.8129 | 0.010386 | -2.7964 | 0.62231 | 450 | -2.8131 | 0.62234
| 455 |  | 8.0372 | 7.8104 | 0.0099989 | -2.8215 | 0.62236 | 465 | -2.8581 | 0.62239
| 470 |  | 8.0367 | 7.8055 | 0.0089499 | -2.8764 | 0.62241 | 480 | -2.8789 | 0.62241
| 485 |  | 8.0367 | 7.8052 | 0.0097078 | -2.8801 | 0.62241 | 495 | -2.9075 | 0.62241
| 500 | 8.037 | 8.0367 | 7.8019 | 0.010696 | -2.9212 | 0.62241 | 510 | -2.9223 | 0.62247
| 515 |  | 8.0357 | 7.8008 | 0.0096651 | -2.9228 | 0.62249 | 525 | -2.9363 | 0.62255
| 530 |  | 8.0347 | 7.7982 | 0.010628 | -2.9431 | 0.62257
| 545 |  | 8.0337 | 7.7947 | 0.0097093 | -2.9746 | 0.62266
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
| 89 |  | 8.1578 | 8.1045 | 0.025644 | -0.65361 | 2.55 | 90 | -0.6323 | 2.55
| 104 |  | 8.1467 | 8.1195 | 0.0312 | -0.33398 | 2.55 | 105 | -0.32311 | 2.55
| 119 |  | 8.1384 | 8.1245 | 0.022589 | -0.17097 | 2.55 | 120 | -0.20568 | 2.55
| 134 |  | 8.1277 | 8.0715 | 0.031834 | -0.69161 | 2.55 | 135 | -0.76091 | 2.55
| 149 |  | 8.117 | 7.9765 | 0.03407 | -1.731 | 2.55 | 150 | -1.7519 | 2.55
| 164 |  | 8.1088 | 7.943 | 0.027739 | -2.0444 | 2.55 | 165 | -2.0551 | 2.55
| 179 |  | 8.1088 | 7.93 | 0.031623 | -2.2044 | 2.55 | 180 | -2.1984 | 2.55
| 194 |  | 8.1135 | 7.942 | 0.024623 | -2.1143 | 2.55 | 195 | -2.1266 | 2.55
| 209 |  | 8.1218 | 7.935 | 0.022361 | -2.2996 | 2.55 | 210 | -2.2944 | 2.55
| 224 |  | 8.1326 | 7.952 | 0.020417 | -2.2212 | 2.55 | 225 | -2.2333 | 2.55
| 239 |  | 8.1437 | 7.948 | 0.018238 | -2.4028 | 2.55 | 240 | -2.4006 | 2.55
| 254 |  | 8.1507 | 7.9575 | 0.016819 | -2.3706 | 2.55 | 255 | -2.3736 | 2.55
| 269 |  | 8.1539 | 7.957 | 0.018666 | -2.4151 | 2.55 | 270 | -2.4277 | 2.55
| 284 |  | 8.1507 | 7.9385 | 0.018715 | -2.604 | 2.55 | 285 | -2.6021 | 2.55
| 299 |  | 8.1387 | 7.929 | 0.021001 | -2.5767 | 2.55 | 300 | -2.5785 | 2.55
| 314 |  | 8.1261 | 7.9145 | 0.017911 | -2.6037 | 2.55 | 315 | -2.6137 | 2.55
| 329 | 8.1083 | 8.1083 | 7.885 | 0.023508 | -2.7537 | 2.55 | 330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 42 |  | 7.98 | 7.396 | 0.014 | -7.318 | 0.64 | 45
| 57 |  | 7.992 | 7.624 | 0.016 | -4.605 | 0.64 | 60 | -4.4354 | 0.64
LEI | 72 |  | 7.986 | 7.686 | 0.013 | -3.757 | 0.64 | 75 | -3.6694 | 0.64
| 87 |  | 7.984 | 7.719 | 0.014 | -3.319 | 0.64 | 90 | -3.1872 | 0.64
| 102 |  | 7.971 | 7.759 | 0.014 | -2.66 | 0.64 | 105 | -2.5344 | 0.64
| 117 |  | 7.971 | 7.809 | 0.01 | -2.032 | 0.64 | 120 | -1.874 | 0.64
| 132 |  | 7.971 | 7.872 | 0.012 | -1.242 | 0.64 | 135 | -1.1342 | 0.64
| 147 |  | 7.968 | 7.912 | 0.013 | -0.703 | 0.64 | 150 | -0.6428 | 0.64
| 162 |  | 7.97 | 7.938 | 0.012 | -0.402 | 0.64 | 165 | -0.3818 | 0.64
| 177 |  | 7.976 | 7.952 | 0.013 | -0.301 | 0.64 | 180 | -0.3336 | 0.64
| 192 |  | 7.982 | 7.945 | 0.013 | -0.464 | 0.64 | 195 | -0.514 | 0.64
| 207 |  | 7.979 | 7.922 | 0.012 | -0.714 | 0.64 | 210 | -0.7742 | 0.64
| 222 |  | 7.982 | 7.901 | 0.01 | -1.015 | 0.64 | 225 | -1.0502 | 0.64
| 237 |  | 7.977 | 7.882 | 0.01 | -1.191 | 0.64 | 240 | -1.2586 | 0.64
| 252 |  | 7.977 | 7.855 | 0.015 | -1.529 | 0.64 | 255 | -1.5918 | 0.64
| 267 |  | 7.975 | 7.828 | 0.01 | -1.843 | 0.64 | 270 | -1.8482 | 0.64
| 282 |  | 7.972 | 7.823 | 0.015 | -1.869 | 0.64 | 285 | -1.8868 | 0.64
| 297 |  | 7.968 | 7.812 | 0.012 | -1.958 | 0.64 | 300 | -2.008 | 0.64
| 312 |  | 7.97 | 7.794 | 0.012 | -2.208 | 0.64 | 315 | -2.286 | 0.64
| 327 |  | 7.969 | 7.762 | 0.013 | -2.598 | 0.64 | 330 | -2.7964 | 0.64
| 342 |  | 7.966 | 7.68 | 0.013 | -3.59 | 0.64 | 345
| 357 |  | 7.971 | 7.446 | 0.017 | -6.586 | 0.64 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 12
| 12 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.21229
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 11.8323 | 11.438 | 0.010328 | -3.3326 | 0.63386 | 45
| 52.5 |  | 11.8263 | 11.514 | 0.0096609 | -2.6407 | 0.63418 | 60 | -2.478 | 0.63397 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 11.834 | 11.56 | 0.013333 | -2.3153 | 0.63377 | 75 | -2.0822 | 0.63775 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 11.843 | 11.624 | 0.012649 | -1.8491 | 0.64173 | 90 | -1.5355 | 0.64185 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 11.8386 | 11.694 | 0.013499 | -1.2218 | 0.64197 | 105 | -0.89662 | 0.64177 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 11.8457 | 11.778 | 0.010328 | -0.57145 | 0.64158 | 120 | -0.24598 | 0.64199 |  | 1.272 | 1.027 | 2.5956 |  | 0.072092 | 0.19852 | 1.7531 |  | 1.5356 | 1.7914 |  | 1.2281 |  | 2.9536 |  | 1.1147 |  |  | 1.4565 | 0.39719 | 0.53779 | 1.1647 |  | 1.6824 | 0.47197 | 0.87351 | 0.61977 | 2.3761 |  | 0.60721
| 127.5 | 11.8392 | 11.8306 | 11.84 | 0.010541 | 0.079482 | 0.6424 | 135 | 0.18705 | 0.64195 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 11.8471 | 11.882 | 0.010328 | 0.29461 | 0.64151 | 150 | 0.40413 | 0.64213 | 45
| 157.5 |  | 11.8243 | 11.885 | 0.015811 | 0.51364 | 0.64275 | 165 | 0.35621 | 0.64252 | 60 | 1.8972 | 1.326 |  |  | 1.9157 | 0.53992 |  |  | 0.19368 |  |  | 0.66704
| 172.5 |  | 11.8325 | 11.856 | 0.0096609 | 0.19879 | 0.6423 | 180 | -0.037367 | 0.64209 | 75 | 1.9316 | 1.1526 |  |  | 0.9795 | 0.79731 |  |  | 0.98991 |  |  | 0.18451 | [12 empty] | 4.9425 | 3.864 | 4.2625 |  |  | 4.4028
| 187.5 |  | 11.8404 | 11.808 | 0.011353 | -0.27352 | 0.64187 | 195 | -0.59205 | 0.63796 | 90 | 1.8826 | 1.0344 | 1.3935 |  | 0.83404 | 0.88981 | 0.37355 |  | 1.1176 | 0.43748 |  | 0.21719 |  | 0.64271 | [10 empty] | 3.1076 | 1.6955 | 2.3483 | 1.8366 |  | 2.5395
| 202.5 |  | 11.8287 | 11.721 | 0.011972 | -0.91059 | 0.63405 | 210 |  |  | 105 | 1.8904 | 0.94951 | 1.1648 |  | 0.91034 | 1.0058 | 0.658 |  | 1.0328 | 0.27566 |  | 0.035702 |  | 0.30739 | [10 empty] | 2.0238 | 0.3669 | 1.2299 | 0.91435 |  | 1.253
| 217.5 |  | 11.8701 | 11.602 | 0.014757 | -2.2585 | 0.63184 | 225 |  |  | 120 | 2.0123 | 0.83809 | 1.2033 |  | 0.82674 | 1.2854 | 0.74417 |  | 1.2741 | 0.44035 |  | 0.0047688 |  | 0.44056 |  | 0.20171 |  |  | 1.6104 | 0.21011 | 0.90792 | 0.45935 |  | 0.90379 | 1.8694 | 0.010824 | 1.1712 | 0.68797 |  | 1.1636
[8 empty] | 240 |  |  | 135 | 1.9555 | 0.58145 | 1.168 |  | 0.60054 | 1.5342 | 0.72462 |  | 1.487 | 0.67737 |  | 0.026633 |  | 0.64562 |  | 0.22569 |  |  | 1.4223 | 0.38719 | 0.96655 | 0.28865 |  | 0.93282 | 1.6731 | 0.14312 | 1.2377 | 0.52674 |  | 1.2
[8 empty] | 255 |  |  | 150 | 1.8379 | 0.46462 | 1.0177 |  | 0.1741 | 1.5439 | 0.778 |  | 1.8767 | 0.63601 |  | 0.32285 |  | 0.93944 |  | 0.52273 |  |  | 0.82078 | 0.9736 | 0.42275 | 0.19201 |  | 0.72178 | 1.3332 | 0.39995 | 0.98837 | 0.33711 |  | 1.2891
[8 empty] | 270 |  |  | 165 | 1.6632 | 0.32611 | 0.68752 |  | 0.33113 | 1.5124 | 0.98721 |  | 2.2987 | 0.41611 |  | 0.73696 |  | 1.1193 |  | 0.65141 |  |  | 0.3411 | 1.3549 | 0.03615 | 0.3555 |  | 0.72436 | 0.97796 | 0.69766 | 0.73602 | 0.29217 |  | 1.4613
[8 empty] | 285 |  |  | 180 | 1.3117 | 0.034054 | 0.12545 |  | 0.86038 | 1.5532 | 1.2831 |  | 2.5472 | 0.17717 |  | 0.93118 |  | 1.0675 |  | 0.72067 |  |  | 0.38083 | 1.8548 | 0.38821 | 0.54017 |  | 0.48787 | 0.28763 | 1.1265 | 0.36124 | 0.17212 |  | 1.2719
[8 empty] | 300 |  |  | 195 | 0.84163 | 0.4274 | 0.51633 |  | 1.2295 | 1.4876 | 1.54 |  | 2.4559 | 0.11464 |  | 0.90195 |  | 0.75191 |  | 0.1245 |  |  | 1.0771 | 2.1821 | 0.74401 | 0.60903 |  | 0.10824 | 1.0043 | 2.1738 | 0.65067 | 0.51141 |  | 0.24531
[8 empty] | 315 |  |  | 210 | [5 empty] | 1.6109 | 1.721 |  | 2.2658 | 0.17507 |  | 0.61794 |  | 0.41899 |  | 0.24839 |  |  |  | 2.1275 | 0.59482 | 0.40857 |  | 0.012081 |  | 2.5325 | 0.88775 | 0.68072 |  | 0.27542
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.5955 | 1.8409 |  | 2.1291 | 0.29868 |  | 0.50576 |  | 0.18898 |  | 0.16567 |  |  |  | 2.2822 | 0.73893 | 0.42933 |  | 0.25664 |  | 2.6157 | 0.95364 | 0.61886 |  | 0.4442
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.5435 | 1.9965 |  | 2.2358 | 0.48733 |  | 0.64998 |  | 0.14028 |  | 0.66062 |  |  |  | 2.3658 | 0.86084 | 0.36865 |  | 0.24201 |  | 3.285 | 1.6098 | 1.0605 |  | 0.95057
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.8254 | 2.2969 |  | 2.604 | 0.51764 |  | 0.73248 |  | 0.1893 |  | 0.70132 |  |  |  | 2.5266 | 0.78882 | 0.27608 |  | 0.10038 |  | 3.556 | 1.597 | 1.0162 |  | 0.85378
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.2069 |  | 0.55531 |  | 0.3274 |  | 0.3778 |  |  |  |  | 0.39791 | 0.18891 |  | 0.12491 |  |  | 0.81945 | 0.58398 |  | 0.26894
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.196 |  | 0.83438 |  | 0.60543 |  | 0.10107 |  |  |  |  | 0.51287 | 0.30848 |  | 0.27645 |  |  | 0.42865 | 0.21834 |  | 0.39501
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.36301 |  | 0.91368 |  | 0.51559 |  | 0.22164 |  |  |  |  | 0.85425 | 0.48268 |  | 0.0043163 |  |  | 0.65971 | 0.27727 |  | 0.23975
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.45983 |  | 0.82247 |  | 0.3324 |  | 0.46503 |  |  |  |  | 0.82528 | 0.35973 |  | 0.044223 |  |  | 1.3425 | 0.84272 |  | 0.53384
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.48318 |  | 0.40292 |  | 0.092581 |  | 0.8362 |  |  |  |  | 0.65133 | 0.17424 |  | 0.26955 |  |  | 1.5766 | 1.0399 |  | 1.1679
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.61293 | [5 empty] | 0.72251 |  |  |  |  | 0.62141 | 0.020156 |  |  |  |  | 1.4195 | 0.75817
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.88949 | [5 empty] | 0.50385 |  |  |  |  | 0.99011 | 0.1146 |  |  |  |  | 1.5737 | 0.63392
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.2052 | [5 empty] | 0.18588 |  |  |  |  | 1.4637 | 0.26519 |  |  |  |  | 1.7265 | 0.46479
[8 empty] | 495 |  |  | 390 | [9 empty] | 1.3283 | [5 empty] | 0.058196 |  |  |  |  | 1.5761 | 0.26608 |  |  |  |  | 1.7195 | 0.33831
[8 empty] | 510 |  |  | 405 | [15 empty] | 0.22059 | [5 empty] | 0.104 | [5 empty] | 0.33081
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.0089494 | [5 empty] | 0.34915 | [5 empty] | 0.3515
[11 empty] | 435 | [27 empty] | 0.26241
450
45
PTB | 55 | 12.0081 | 12.0706 | 11.5808 | 0.01222 | -4.0582 | 0.41 | 60 | -3.9659 | 0.41
| 70 | 12.0081 | 12.0575 | 11.6015 | 0.011204 | -3.7814 | 0.41 | 75 | -3.603 | 0.41
| 85 | 12.0081 | 12.0432 | 11.6522 | 0.013681 | -3.2461 | 0.41 | 90 | -3.024 | 0.41
| 100 | 12.0081 | 12.0343 | 11.7238 | 0.01329 | -2.5797 | 0.41 | 105 | -2.3912 | 0.41
| 115 | 12.0081 | 12.0315 | 11.7892 | 0.0070238 | -2.0141 | 0.41 | 120 | -1.8372 | 0.41
| 130 | 12.0081 | 12.0187 | 11.8404 | 0.010215 | -1.4834 | 0.41 | 135 | -1.3592 | 0.41
| 145 | 12.0081 | 12.0127 | 11.8792 | 0.010168 | -1.1108 | 0.41 | 150 | -1.0495 | 0.41
| 160 | 12.0081 | 12.0081 | 11.8968 | 0.015737 | -0.92683 | 0.41 | 165 | -0.95972 | 0.41
| 175 | 12.0081 | 12.0164 | 11.8932 | 0.014353 | -1.0255 | 0.41 | 180 | -1.0747 | 0.41
| 190 | 12.0081 | 12.0234 | 11.8824 | 0.012675 | -1.1731 | 0.41 | 195 | -1.2548 | 0.41
| 205 | 12.0081 | 12.027 | 11.8564 | 0.011136 | -1.4183 | 0.41 | 210 | -1.5292 | 0.41
| 220 | 12.0081 | 12.0401 | 11.8292 | 0.015472 | -1.751 | 0.41 | 225 | -1.8874 | 0.41
| 235 | 12.0081 | 12.0507 | 11.7904 | 0.016452 | -2.16 | 0.41 | 240 | -2.3055 | 0.41
| 250 | 12.0081 | 12.0649 | 11.7516 | 0.020551 | -2.5965 | 0.41 | 255 | -2.784 | 0.41
| 265 | 12.0081 | 12.0767 | 11.6952 | 0.013266 | -3.1589 | 0.41 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 12.0383 | 11.519 | 0.012096 | -4.314 | 0.5174 | 45
| 60 |  | 12.0459 | 11.613 | 0.011743 | -3.594 | 0.51123 | 60 | -3.594 | 0.51123
| 75 |  | 12.0387 | 11.671 | 0.012096 | -3.0546 | 0.50985 | 75 | -3.0546 | 0.50985
| 90 |  | 12.0429 | 11.7525 | 0.015853 | -2.4116 | 0.51032 | 90 | -2.4116 | 0.51032
| 105 |  | 12.0366 | 11.832 | 0.012814 | -1.7001 | 0.50896 | 105 | -1.7001 | 0.50896
| 120 |  | 12.0369 | 11.922 | 0.010563 | -0.95481 | 0.50803 | 120 | -0.95481 | 0.50803
| 135 |  | 12.0257 | 11.989 | 0.012096 | -0.30501 | 0.50893 | 135 | -0.30501 | 0.50893
| 150 |  | 12.0227 | 12.024 | 0.010463 | 0.010979 | 0.50855 | 150 | 0.010979 | 0.50855
| 165 |  | 12.0244 | 12.034 | 0.013139 | 0.080004 | 0.50938 | 165 | 0.080004 | 0.50938
| 180 |  | 12.026 | 12.025 | 0.0088852 | -0.0085648 | 0.50794 | 180 | -0.0085648 | 0.50794
| 195 |  | 12.0298 | 12.002 | 0.015761 | -0.23134 | 0.51012 | 195 | -0.23134 | 0.51012
| 210 |  | 12.0197 | 11.969 | 0.012096 | -0.42206 | 0.50915 | 210 | -0.42206 | 0.50915
| 225 | 12.024 | 12.024 | 11.929 | 0.015861 | -0.79033 | 0.50952 | 225 | -0.79033 | 0.50952
| 240 |  | 12.0337 | 11.884 | 0.015694 | -1.2438 | 0.50984 | 240 | -1.2438 | 0.50984
| 255 |  | 12.035 | 11.851 | 0.015183 | -1.5287 | 0.5096 | 255 | -1.5287 | 0.5096
| 270 |  | 12.0245 | 11.839 | 0.013727 | -1.5425 | 0.50955 | 270 | -1.5425 | 0.50955
| 285 |  | 12.0267 | 11.823 | 0.011743 | -1.6936 | 0.50885 | 285 | -1.6936 | 0.50885
| 300 |  | 12.0362 | 11.801 | 0.010208 | -1.9539 | 0.508 | 300 | -1.9539 | 0.508
| 315 |  | 12.036 | 11.779 | 0.010208 | -2.1351 | 0.50804 | 315 | -2.1351 | 0.50804
| 330 |  | 12.0339 | 11.7685 | 0.013089 | -2.2053 | 0.50924 | 330 | -2.2053 | 0.50924
| 345 |  | 12.0347 | 11.747 | 0.015927 | -2.3904 | 0.51076 | 345 | -2.3904 | 0.51076
| 360 |  | 12.0431 | 11.722 | 0.016416 | -2.6661 | 0.51102 | 360 | -2.6661 | 0.51102
| 375 |  | 12.0459 | 11.69 | 0.015218 | -2.9548 | 0.5116 | 375 | -2.9548 | 0.5116
| 390 |  | 12.0443 | 11.671 | 0.012096 | -3.0996 | 0.51356 | 390 | -3.0996 | 0.51356
| 405 |  | 12.0489 | 11.605 | 0.011002 | -3.6844 | 0.52165 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 11.9509 | 11.75 | 0.05 | -1.6813 | 0.58068 | 120 | -1.6813 | 0.58068
| 135 |  | 11.9483 | 11.82 | 0.02 | -1.0741 | 0.57365 | 135 | -1.0741 | 0.57365
| 150 |  | 11.9491 | 11.91 | 0.03 | -0.32737 | 0.58043 | 150 | -0.32737 | 0.58043
| 165 |  | 11.9539 | 11.96 | 0.04 | 0.050909 | 0.58586 | 165 | 0.050909 | 0.58586
| 180 |  | 11.9539 | 11.99 | 0.02 | 0.30224 | 0.58131 | 180 | 0.30224 | 0.58131
| 195 |  | 11.9563 | 12 | 0.02 | 0.36549 | 0.58155 | 195 | 0.36549 | 0.58155
| 210 |  | 11.9733 | 11.98 | 0.04 | 0.056197 | 0.58494 | 210 | 0.056197 | 0.58494
| 225 |  | 11.974 | 11.95 | 0.02 | -0.20006 | 0.57745 | 225 | -0.20006 | 0.57745
| 240 |  | 11.9766 | 11.91 | 0.03 | -0.55579 | 0.5778 | 240 | -0.55579 | 0.5778
| 255 |  | 11.9667 | 11.86 | 0.06 | -0.89145 | 0.58983 | 255 | -0.89145 | 0.58983
| 270 | 11.962 | 11.9629 | 11.8167 | 0.053333 | -1.2226 | 0.5844 | 270 | -1.2226 | 0.5844
| 285 |  | 11.9636 | 11.81 | 0.04 | -1.284 | 0.57779 | 285 | -1.284 | 0.57779
| 300 |  | 11.9512 | 11.8 | 0.06 | -1.2655 | 0.58849 | 300 | -1.2655 | 0.58849
| 315 |  | 11.9464 | 11.77 | 0.04 | -1.4766 | 0.57753 | 315 | -1.4766 | 0.57753
| 330 |  | 11.9509 | 11.75 | 0.06 | -1.681 | 0.58618 | 330 | -1.681 | 0.58618
| 345 |  | 11.9563 | 11.73 | 0.05 | -1.8927 | 0.57922 | 345 | -1.8927 | 0.57922
| 360 |  | 11.9539 | 11.73 | 0.05 | -1.8728 | 0.57946 | 360 | -1.8728 | 0.57946
| 375 |  | 11.9539 | 11.74 | 0.03 | -1.7895 | 0.57182 | 375 | -1.7895 | 0.57182
| 390 |  | 11.9491 | 11.73 | 0.05 | -1.8338 | 0.57991 | 390 | -1.8338 | 0.57991
| 405 |  | 11.9483 | 11.71 | 0.02 | -1.9947 | 0.56835 | 405 | -1.9947 | 0.56835
| 420 |  | 11.9509 | 11.73 | 0.05 | -1.8487 | 0.57973 | 420 | -1.8487 | 0.57973
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
| 75 |  | 12.0381 | 11.2645 | 0.021637 | -6.426 | 0.56622 | 75 | -6.426 | 0.56622
| 90 |  | 12.0323 | 11.5225 | 0.01916 | -4.2371 | 0.54663 | 90 | -4.2371 | 0.54663
| 105 |  | 12.0327 | 11.7135 | 0.021588 | -2.6531 | 0.54438 | 105 | -2.6531 | 0.54438
| 120 |  | 12.0411 | 11.819 | 0.015183 | -1.8448 | 0.52366 | 120 | -1.8448 | 0.52366
| 135 |  | 12.0303 | 11.879 | 0.02174 | -1.2578 | 0.53723 | 135 | -1.2578 | 0.53723
| 150 |  | 12.021 | 11.9295 | 0.027621 | -0.76113 | 0.55369 | 150 | -0.76113 | 0.55369
| 165 |  | 12.0383 | 11.981 | 0.015183 | -0.47635 | 0.51658 | 165 | -0.47635 | 0.51658
| 180 |  | 12.0402 | 12.006 | 0.020622 | -0.28423 | 0.52844 | 180 | -0.28423 | 0.52844
| 195 |  | 12.0357 | 12.0675 | 0.020229 | 0.26445 | 0.5247 | 195 | 0.26445 | 0.5247
| 210 |  | 12.0327 | 12.064 | 0.023709 | 0.26 | 0.53477 | 210 | 0.26 | 0.53477
| 225 |  | 12.043 | 12.035 | 0.018209 | -0.066498 | 0.521 | 225 | -0.066498 | 0.521
| 240 |  | 12.0369 | 12.034 | 0.01729 | -0.02394 | 0.51888 | 240 | -0.02394 | 0.51888
| 255 | 12.0087 | 12.0307 | 11.992 | 0.015079 | -0.32208 | 0.51589 | 255 | -0.32208 | 0.51589
| 270 |  | 12.0369 | 11.927 | 0.021546 | -0.91287 | 0.53451 | 270 | -0.91287 | 0.53451
| 285 |  | 12.043 | 11.8785 | 0.018715 | -1.366 | 0.52912 | 285 | -1.366 | 0.52912
| 300 |  | 12.0327 | 11.8585 | 0.020333 | -1.4478 | 0.53423 | 300 | -1.4478 | 0.53423
| 315 |  | 12.0357 | 11.904 | 0.025005 | -1.094 | 0.54605 | 315 | -1.094 | 0.54605
| 330 |  | 12.0402 | 11.9205 | 0.021392 | -0.99435 | 0.53437 | 330 | -0.99435 | 0.53437
| 345 |  | 12.0383 | 11.881 | 0.017442 | -1.307 | 0.52591 | 345 | -1.307 | 0.52591
| 360 |  | 12.021 | 11.845 | 0.017014 | -1.4641 | 0.52651 | 360 | -1.4641 | 0.52651
| 375 |  | 12.0303 | 11.833 | 0.014903 | -1.6402 | 0.52246 | 375 | -1.6402 | 0.52246
| 390 |  | 12.0411 | 11.826 | 0.015355 | -1.7866 | 0.52371 | 390 | -1.7866 | 0.52371
| 405 |  | 12.0327 | 11.814 | 0.015355 | -1.8179 | 0.52424 | 405 | -1.8179 | 0.52424
| 420 |  | 12.0323 | 11.809 | 0.021981 | -1.856 | 0.54111 | 420 | -1.856 | 0.54111
| 435 |  | 12.0381 | 11.805 | 0.017014 | -1.9361 | 0.5283 | 435 | -1.9361 | 0.5283
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
| 80 |  | 12.0873 | 11.6634 | 0.029902 | -3.5073 | 0.55463 | 90 | -2.7543 | 0.555
| 95 |  | 12.0713 | 11.7843 | 0.013607 | -2.3778 | 0.55519 | 105 | -1.9158 | 0.5556
| 110 |  | 12.0563 | 11.8532 | 0.013584 | -1.6849 | 0.5558 | 120 | -1.2994 | 0.55602
| 125 |  | 12.0483 | 11.915 | 0.016949 | -1.1067 | 0.55613 | 135 | -0.83548 | 0.55608
| 140 |  | 12.0493 | 11.965 | 0.01379 | -0.6999 | 0.55606 | 150 | -0.48703 | 0.55627
| 155 |  | 12.0423 | 11.9965 | 0.018108 | -0.3806 | 0.55637 | 165 | -0.24605 | 0.55628
| 170 |  | 12.0453 | 12.0238 | 0.0128 | -0.17877 | 0.55623 | 180 | -0.14721 | 0.5562
| 185 |  | 12.0463 | 12.0305 | 0.016064 | -0.13144 | 0.55619 | 195 | -0.14145 | 0.55637
| 200 |  | 12.0403 | 12.0227 | 0.014851 | -0.14645 | 0.55646 | 210 | -0.28488 | 0.55647
| 215 |  | 12.0403 | 11.9977 | 0.013539 | -0.35409 | 0.55647 | 225 | -0.55622 | 0.5565
| 230 |  | 12.0393 | 11.9602 | 0.013135 | -0.65729 | 0.55652 | 240 | -0.86173 | 0.55659
| 245 |  | 12.0373 | 11.9213 | 0.019916 | -0.96395 | 0.55663 | 255 | -1.1229 | 0.55664
| 260 |  | 12.0373 | 11.8926 | 0.019478 | -1.2024 | 0.55664 | 270 | -1.3803 | 0.55672
| 275 |  | 12.0353 | 11.8585 | 0.014503 | -1.4693 | 0.55675 | 285 | -1.54 | 0.55673
| 290 |  | 12.0361 | 11.8465 | 0.014095 | -1.5754 | 0.55673 | 300 | -1.6697 | 0.55668
| 305 |  | 12.0378 | 11.8311 | 0.015292 | -1.7169 | 0.55666 | 315 | -1.7751 | 0.55667
| 320 |  | 12.0378 | 11.8206 | 0.014035 | -1.8041 | 0.55667 | 330 | -1.8266 | 0.55665
| 335 |  | 12.0384 | 11.8172 | 0.013213 | -1.8378 | 0.55664 | 345 | -1.9095 | 0.55656
| 350 |  | 12.0414 | 11.8072 | 0.011374 | -1.9453 | 0.55651 | 360 | -1.968 | 0.55652
| 365 |  | 12.0414 | 11.8031 | 0.011839 | -1.9794 | 0.55652 | 375 | -2.0085 | 0.5565
| 380 |  | 12.0421 | 11.7985 | 0.010612 | -2.023 | 0.55649 | 390 | -2.055 | 0.55651
| 395 |  | 12.0418 | 11.7924 | 0.011723 | -2.0709 | 0.55651 | 405 | -2.0803 | 0.55651
| 410 |  | 12.0418 | 11.7907 | 0.011399 | -2.0851 | 0.55651 | 420 | -2.1389 | 0.55651
| 425 |  | 12.0421 | 11.7813 | 0.011455 | -2.1658 | 0.55651 | 435 | -2.145 | 0.55657
| 440 |  | 12.0401 | 11.7831 | 0.010402 | -2.1346 | 0.5566 | 450 | -2.1641 | 0.55659
| 455 |  | 12.0404 | 11.7781 | 0.010613 | -2.1789 | 0.55659 | 465 | -2.1945 | 0.55658
| 470 |  | 12.0408 | 11.7756 | 0.012099 | -2.2023 | 0.55657 | 480 | -2.205 | 0.55105
| 485 |  | 12.0404 | 11.7748 | 0.011366 | -2.2063 | 0.54829 | 495 | -2.1986 | 0.5539
| 500 | 12.038 | 12.038 | 11.7738 | 0.011924 | -2.1947 | 0.5567 | 510 | -2.1963 | 0.55663
| 515 |  | 12.0404 | 11.7759 | 0.012119 | -2.1971 | 0.55659 | 525 | -2.2284 | 0.55108
| 530 |  | 12.0398 | 11.7696 | 0.011714 | -2.244 | 0.54832
| 545 |  | 12.038 | 11.7676 | 0.010341 | -2.2462 | 0.5484
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 42 |  | 12.075 | 11.151 | 0.021 | -7.652 | 0.52 | 45
| 57 |  | 12.082 | 11.555 | 0.013 | -4.362 | 0.52 | 60 | -4.1006 | 0.52
LEI | 72 |  | 12.08 | 11.711 | 0.015 | -3.055 | 0.52 | 75 | -2.9146 | 0.52
| 87 |  | 12.072 | 11.788 | 0.013 | -2.353 | 0.52 | 90 | -2.2468 | 0.52
| 102 |  | 12.073 | 11.853 | 0.014 | -1.822 | 0.52 | 105 | -1.673 | 0.52
| 117 |  | 12.066 | 11.936 | 0.013 | -1.077 | 0.52 | 120 | -0.9512 | 0.52
| 132 |  | 12.067 | 12.013 | 0.012 | -0.448 | 0.52 | 135 | -0.3252 | 0.52
| 147 |  | 12.068 | 12.088 | 0.018 | 0.166 | 0.52 | 150 | 0.2556 | 0.52
| 162 |  | 12.059 | 12.133 | 0.014 | 0.614 | 0.52 | 165 | 0.6388 | 0.52
| 177 |  | 12.062 | 12.151 | 0.014 | 0.738 | 0.52 | 180 | 0.6966 | 0.52
| 192 |  | 12.057 | 12.121 | 0.011 | 0.531 | 0.52 | 195 | 0.453 | 0.52
| 207 |  | 12.054 | 12.071 | 0.018 | 0.141 | 0.52 | 210 | 0.0464 | 0.52
| 222 |  | 12.049 | 12.009 | 0.014 | -0.332 | 0.52 | 225 | -0.4068 | 0.52
| 237 |  | 12.044 | 11.959 | 0.016 | -0.706 | 0.52 | 240 | -0.7508 | 0.52
| 252 |  | 12.038 | 11.926 | 0.018 | -0.93 | 0.52 | 255 | -0.9732 | 0.52
| 267 |  | 12.039 | 11.901 | 0.019 | -1.146 | 0.52 | 270 | -1.1214 | 0.52
| 282 |  | 12.029 | 11.906 | 0.02 | -1.023 | 0.52 | 285 | -1.0612 | 0.52
| 297 |  | 12.028 | 11.882 | 0.014 | -1.214 | 0.52 | 300 | -1.262 | 0.52
| 312 |  | 12.039 | 11.864 | 0.016 | -1.454 | 0.52 | 315 | -1.5122 | 0.52
| 327 |  | 12.037 | 11.827 | 0.012 | -1.745 | 0.52 | 330 | -1.8998 | 0.52
| 342 |  | 12.03 | 11.727 | 0.018 | -2.519 | 0.52 | 345
| 357 |  | 12.035 | 11.422 | 0.015 | -5.093 | 0.52 | 360
| 372 |  | 12.035 | 10.709 | 0.015 | -11.018 | 0.52 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## S-L 20
| 20 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.10614
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 19.7016 | 19.122 | 0.013984 | -2.9421 | 0.62431 | 45
| 52.5 |  | 19.705 | 19.245 | 0.015811 | -2.3344 | 0.62421 | 60 | -2.1643 | 0.62437 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 19.6948 | 19.302 | 0.014757 | -1.9942 | 0.62453 | 75 | -1.8003 | 0.62422 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 19.7147 | 19.398 | 0.014757 | -1.6065 | 0.6239 | 90 | -1.2971 | 0.62396 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 19.7107 | 19.516 | 0.012649 | -0.98769 | 0.62403 | 105 | -0.61928 | 0.62646 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 19.7175 | 19.668 | 0.015492 | -0.25087 | 0.62888 | 120 | 0.0079056 | 0.62898 |  | 1.7662 | 2.0677 | 2.9419 |  | 1.1495 | 0.53556 | 1.6407 |  | 0.64032 | 1.0259 |  | 1.0732 |  | 2.0875 |  | 0.95047 | 0.91323 | 0.52322 | 1.9264 | 0.41051 | 0.094469 | 1.0856 |  | 0.93923 | 1.0924 | 0.65901 | 1.0804 | 2.0699 |  | 0.036785 | 1.3675 | 0.68798 | 0.48759 | 0.070327 |  | 0.9028
| 127.5 | 19.7075 | 19.7114 | 19.764 | 0.015776 | 0.26668 | 0.62908 | 135 | 0.35434 | 0.62886 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 19.7248 | 19.812 | 0.010328 | 0.442 | 0.62865 | 150 | 0.55641 | 0.62924 | 45
| 157.5 |  | 19.6879 | 19.82 | 0.018257 | 0.67083 | 0.62983 | 165 | 0.46635 | 0.6295 | 60 | 1.9307 | 1.4854 |  |  | 2.3384 | 0.35718 |  |  | 0.64217 |  |  | 0.90747
| 172.5 |  | 19.7084 | 19.76 | 0.013333 | 0.26187 | 0.62917 | 180 | -0.050295 | 0.62908 | 75 | 2.1564 | 1.3891 |  |  | 1.0802 | 0.73144 |  |  | 1.1762 |  |  | 0.37713 | [12 empty] | 10.8666 | 10.7136 | 10.4313 |  |  | 11.0326
| 187.5 |  | 19.7145 | 19.643 | 0.016364 | -0.36246 | 0.62898 | 195 | -0.75828 | 0.62688 | 90 | 2.1041 | 1.2348 | 1.6273 |  | 0.98535 | 0.86428 | 0.32262 |  | 1.2322 | 0.46496 |  | 0.30606 |  | 0.77283 |  |  | 4.5083 | [7 empty] | 9.0516 | 8.5956 | 8.6178 | 7.9613 |  | 9.0939 | 0.36992 | 1.2398 | 0.91426 | 1.0977 |  | 0.79985
| 202.5 |  | 19.6872 | 19.46 | 0.018856 | -1.1541 | 0.62477 | 210 |  |  | 105 | 2.1949 | 1.2634 | 1.4647 |  | 0.96353 | 0.93726 | 0.61922 |  | 1.3689 | 0.2577 |  | 0.36278 |  | 0.61463 |  |  | 3.2819 | [7 empty] | 7.1185 | 6.2226 | 6.4969 | 6.0627 |  | 6.9918 | 0.085824 | 0.98616 | 0.63873 | 0.73957 |  | 0.50144
| 217.5 |  | 19.7464 | 19.23 | 0.028674 | -2.6152 | 0.62796 | 225 |  |  | 120 | 2.1999 | 1.1874 | 1.4904 |  | 0.82589 | 1.04 | 0.59528 |  | 1.5499 | 0.36916 |  | 0.43058 |  | 0.79398 |  | 4.0468 | 2.0144 | 0.33805 | 1.5598 | 0.51411 | 0.44456 | 0.072647 |  | 0.87176 | 5.3274 | 4.059 | 4.6127 | 4.1144 |  | 5.1387 | 0.35338 | 0.53714 | 0.15952 | 0.30778 |  | 0.0056366
[8 empty] | 240 |  |  | 135 | 2.0768 | 0.82113 | 1.2902 |  | 0.38985 | 1.3467 | 0.70382 |  | 1.9471 | 0.54049 |  | 0.50076 |  | 1.0386 |  | 2.8281 | 0.6353 | 0.53162 | 1.2717 | 0.72459 | 0.5206 | 0.018935 |  | 1.0178 | 3.9174 | 2.4743 | 3.4416 | 2.811 |  | 4.0182 | 1.0859 | 0.26987 | 0.74598 | 0.5238 |  | 0.94195
[8 empty] | 255 |  |  | 150 | 2.1389 | 0.70559 | 1.0601 |  | 0.047712 | 1.5621 | 1.0513 |  | 2.5642 | 0.4119 |  | 0.85811 |  | 1.2517 |  | 2.3443 | 0.032557 | 0.99049 | 0.81811 | 1.3297 | 0.14803 | 0.2537 |  | 0.97606 | 3.0297 | 1.3043 | 2.5865 | 2.0978 |  | 3.5029 | 1.3414 | 0.50799 | 1.057 | 0.88531 |  | 1.3915
[8 empty] | 270 |  |  | 165 | 2.0895 | 0.54071 | 0.65417 |  | 0.24809 | 1.7103 | 1.4902 |  | 2.756 | 0.14099 |  | 0.89446 |  | 1.009 |  | 2.1318 | 0.32576 | 1.1919 | 0.32451 | 1.8709 | 0.21992 | 0.34929 |  | 0.63302 | 2.3275 | 0.49044 | 1.9895 | 1.786 |  | 2.9312 | 1.3219 | 0.50797 | 1.1095 | 1.0469 |  | 1.458
[8 empty] | 285 |  |  | 180 | 1.6298 | 0.0049587 | 0.0019503 |  | 0.53789 | 1.8451 | 1.7832 |  | 2.5764 | 0.0075001 |  | 0.607 |  | 0.57929 |  | 2.297 | 0.42139 | 1.3586 | 0.64982 | 2.5523 | 0.72304 | 0.69468 |  | 0.15708 | 1.4813 | 0.020206 | 1.6404 | 1.5969 |  | 2.2836 | 1.0544 | 0.4205 | 1.0714 | 1.069 |  | 1.3092
[8 empty] | 300 |  |  | 195 | 1.234 | 0.46384 | 0.6466 |  | 0.86044 | 1.9646 | 2.1283 |  | 2.5079 | 0.21407 |  | 0.43076 |  | 0.19778 |  | 1.4734 | 0.74267 | 1.3372 | 1.5833 | 3.2456 | 1.248 | 1.0065 |  | 0.86448 | 0.25402 | 1.7091 | 0.233 | 0.44024 |  | 0.67035 | 0.6207 | 0.13591 | 0.83435 | 0.91689 |  | 1.0038
[8 empty] | 315 |  |  | 210 | [5 empty] | 2.0625 | 2.6886 |  | 2.6694 | 0.63744 |  | 0.4839 |  | 0.18475 |  | 0.43602 | 0.94128 | 1.116 |  | 3.5187 | 1.4025 | 0.74623 |  | 0.97079 |  | 3.1298 | 0.98733 | 0.32642 |  | 0.53401 |  | 0.18882 | 0.5495 | 0.80483 |  | 0.73841
[8 empty] | 330 |  |  | 225 | [5 empty] | 2.1838 | 2.7627 |  | 3.2408 | 0.60333 |  | 0.89004 |  | 0.24302 |  | 0.11714 | 0.90465 | 0.94858 |  | 3.6273 | 1.3977 | 0.77239 |  | 0.57222 |  | 3.6125 | 1.3142 | 0.67456 |  | 0.46424 |  | 0.39857 | 0.38461 | 0.6265 |  | 0.72889
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.1357 | 2.6259 |  | 3.205 | 0.52153 |  | 0.90299 |  | 0.3391 |  | 0.56988 | 0.84347 | 0.60737 |  | 3.0744 | 0.93 | 0.39618 |  | 0.078835 |  | 3.7758 | 1.5315 | 0.97144 |  | 0.68216 |  | 0.53507 | 0.23276 | 0.44233 |  | 0.58138
[8 empty] | 360 |  |  | 255 | [5 empty] | 2.271 | 2.4376 |  | 3.5392 | 0.23231 |  | 1.0795 |  | 0.80545 |  | 0.88443 | 0.79829 | 0.43599 |  | 2.9211 | 0.67483 | 0.4293 |  | 0.352 |  | 4.0411 | 1.6102 | 1.3269 |  | 0.5733 |  | 0.65326 | 0.16406 | 0.25707 |  | 0.58026
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.15838 |  | 1.0139 |  | 1.1425 |  | 0.54958 | 0.71065 | 0.48204 |  |  | 0.4606 | 0.60043 |  | 0.50735 |  |  | 1.0324 | 1.1561 |  | 0.068736 |  |  | 0.29718 | 0.23164 |  | 0.68892
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.10781 |  | 1.2749 |  | 1.1213 |  | 0.16039 | 0.4554 | 0.51873 |  |  | 1 | 0.86461 |  | 0.21178 |  |  | 0.85895 | 0.72425 |  | 0.38645 |  |  | 0.11609 | 0.15905 |  | 0.60732
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.29969 |  | 1.3971 |  | 1.044 |  | 0.0087485 | 0.37886 | 0.38135 |  |  | 1.0033 | 0.68235 |  | 0.32532 |  |  | 1.0123 | 0.68511 |  | 0.34068 |  |  | 0.023989 | 0.097297 |  | 0.51414
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.33853 |  | 1.3538 |  | 0.96211 |  | 0.6969 | 0.56154 | 0.27807 |  |  | 0.74045 | 0.38891 |  | 0.5545 |  |  | 1.4811 | 1.0951 |  | 0.16344 |  |  | 0.020581 | 0.11634 |  | 0.5007
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.44183 |  | 1.0325 |  | 0.5459 |  | 0.67879 |  |  |  |  | 0.76512 | 0.3122 |  | 0.21828 |  |  | 1.4893 | 0.99846 |  | 0.49076
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.47087 | [5 empty] | 1.0075 |  |  |  |  | 0.48215 | 0.0094362 |  |  |  |  | 1.527 | 1.0139
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.66309 | [5 empty] | 0.67406 |  |  |  |  | 0.73834 | 0.071096 |  |  |  |  | 1.4448 | 0.74437
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.91758 | [5 empty] | 0.29881 |  |  |  |  | 1.1572 | 0.22932 |  |  |  |  | 1.4692 | 0.52803
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.94404 | [5 empty] | 0.042966 |  |  |  |  | 1.2862 | 0.33254 |  |  |  |  | 1.349 | 0.38014
[8 empty] | 510 |  |  | 405 | [15 empty] | 0.1948 | [5 empty] | 0.12306 | [5 empty] | 0.3196
[8 empty] | 525 |  |  | 420 | [15 empty] | 0.1536 | [5 empty] | 0.26959 | [5 empty] | 0.42474
[11 empty] | 435 | [27 empty] | 0.37779
450
45
PTB | 55 | 19.9698 | 20.047 | 19.3192 | 0.014401 | -3.6303 | 0.41 | 60 | -3.6209 | 0.41
| 70 | 19.9698 | 20.0547 | 19.3323 | 0.015049 | -3.6021 | 0.41 | 75 | -3.427 | 0.41
| 85 | 19.9698 | 20.0151 | 19.3993 | 0.016155 | -3.0767 | 0.41 | 90 | -2.8839 | 0.41
| 100 | 19.9698 | 20.002 | 19.5023 | 0.022326 | -2.4981 | 0.41 | 105 | -2.279 | 0.41
| 115 | 19.9698 | 19.9928 | 19.6248 | 0.014243 | -1.8407 | 0.41 | 120 | -1.6602 | 0.41
| 130 | 19.9698 | 19.9747 | 19.7152 | 0.016955 | -1.2991 | 0.41 | 135 | -1.2202 | 0.41
| 145 | 19.9698 | 19.9692 | 19.757 | 0.014362 | -1.0625 | 0.41 | 150 | -1.0659 | 0.41
| 160 | 19.9698 | 19.9698 | 19.7556 | 0.016946 | -1.0727 | 0.41 | 165 | -1.119 | 0.41
| 175 | 19.9698 | 19.9776 | 19.7356 | 0.015021 | -1.2115 | 0.41 | 180 | -1.2863 | 0.41
| 190 | 19.9698 | 19.9811 | 19.6942 | 0.01901 | -1.4357 | 0.41 | 195 | -1.6919 | 0.41
| 205 | 19.9698 | 19.9996 | 19.5588 | 0.035424 | -2.2042 | 0.41 | 210 | -2.3981 | 0.41
| 220 | 19.9698 | 20.0078 | 19.4504 | 0.028069 | -2.7859 | 0.41 | 225 | -2.8888 | 0.41
| 235 | 19.9698 | 20.0306 | 19.4108 | 0.017187 | -3.0946 | 0.41 | 240 | -3.1629 | 0.41
| 250 | 19.9698 | 20.0456 | 19.3842 | 0.015581 | -3.2997 | 0.41 | 255 | -3.3372 | 0.41
| 265 | 19.9698 | 20.0646 | 19.38 | 0.01633 | -3.4122 | 0.41 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 20.1158 | 19.268 | 0.018238 | -4.2147 | 0.52839 | 45
| 60 |  | 20.1048 | 19.425 | 0.014327 | -3.3814 | 0.51988 | 60 | -3.3814 | 0.51988
| 75 |  | 20.0978 | 19.5075 | 0.01916 | -2.9372 | 0.51859 | 75 | -2.9372 | 0.51859
| 90 |  | 20.0965 | 19.633 | 0.017502 | -2.3062 | 0.51703 | 90 | -2.3062 | 0.51703
| 105 |  | 20.1024 | 19.77 | 0.013765 | -1.6534 | 0.51594 | 105 | -1.6534 | 0.51594
| 120 |  | 20.0982 | 19.904 | 0.013139 | -0.96611 | 0.51576 | 120 | -0.96611 | 0.51576
| 135 |  | 20.0893 | 20.025 | 0.021398 | -0.31992 | 0.51724 | 135 | -0.31992 | 0.51724
| 150 |  | 20.0846 | 20.08 | 0.01451 | -0.022754 | 0.51629 | 150 | -0.022754 | 0.51629
| 165 |  | 20.0825 | 20.087 | 0.011743 | 0.022557 | 0.51586 | 165 | 0.022557 | 0.51586
| 180 |  | 20.0869 | 20.076 | 0.015355 | -0.054364 | 0.51611 | 180 | -0.054364 | 0.51611
| 195 |  | 20.0844 | 20.0085 | 0.022542 | -0.378 | 0.51758 | 195 | -0.378 | 0.51758
| 210 |  | 20.0895 | 19.885 | 0.025236 | -1.018 | 0.518 | 210 | -1.018 | 0.518
| 225 | 20.0731 | 20.0731 | 19.786 | 0.020622 | -1.4304 | 0.51635 | 225 | -1.4304 | 0.51635
| 240 |  | 20.0879 | 19.739 | 0.015183 | -1.737 | 0.51615 | 240 | -1.737 | 0.51615
| 255 |  | 20.0884 | 19.7225 | 0.012085 | -1.8215 | 0.51582 | 255 | -1.8215 | 0.51582
| 270 |  | 20.0732 | 19.73 | 0.015218 | -1.7098 | 0.51653 | 270 | -1.7098 | 0.51653
| 285 |  | 20.0916 | 19.695 | 0.011471 | -1.9741 | 0.51577 | 285 | -1.9741 | 0.51577
| 300 |  | 20.0896 | 19.6445 | 0.017006 | -2.2157 | 0.51652 | 300 | -2.2157 | 0.51652
| 315 |  | 20.0985 | 19.632 | 0.014364 | -2.3212 | 0.5162 | 315 | -2.3212 | 0.5162
| 330 |  | 20.0922 | 19.603 | 0.013416 | -2.4349 | 0.51635 | 330 | -2.4349 | 0.51635
| 345 |  | 20.0851 | 19.584 | 0.01729 | -2.495 | 0.51801 | 345 | -2.495 | 0.51801
| 360 |  | 20.0905 | 19.554 | 0.012312 | -2.6705 | 0.51785 | 360 | -2.6705 | 0.51785
| 375 |  | 20.107 | 19.5245 | 0.012763 | -2.8971 | 0.52005 | 375 | -2.8971 | 0.52005
| 390 |  | 20.0704 | 19.482 | 0.017652 | -2.9318 | 0.52949 | 390 | -2.9318 | 0.52949
| 405 |  | 20.0742 | 19.371 | 0.016512 | -3.5031 | 0.5479 | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 19.9299 | 19.67 | 0.03 | -1.304 | 0.54822 | 120 | -1.304 | 0.54822
| 135 |  | 19.9329 | 19.79 | 0.02 | -0.71701 | 0.55043 | 135 | -0.71701 | 0.55043
| 150 |  | 19.9171 | 19.89 | 0.04 | -0.1362 | 0.55635 | 150 | -0.1362 | 0.55635
| 165 |  | 19.9319 | 19.97 | 0.04 | 0.19126 | 0.55774 | 165 | 0.19126 | 0.55774
| 180 |  | 19.9303 | 20.03 | 0.02 | 0.50014 | 0.55723 | 180 | 0.50014 | 0.55723
| 195 |  | 19.9343 | 20.05 | 0.02 | 0.58061 | 0.55756 | 195 | 0.58061 | 0.55756
| 210 |  | 19.9488 | 19.96 | 0.02 | 0.056343 | 0.55426 | 210 | 0.056343 | 0.55426
| 225 |  | 19.9524 | 19.88 | 0.03 | -0.36275 | 0.55279 | 225 | -0.36275 | 0.55279
| 240 |  | 19.9653 | 19.76 | 0.04 | -1.0285 | 0.5501 | 240 | -1.0285 | 0.5501
| 255 |  | 19.9407 | 19.68 | 0.05 | -1.3072 | 0.55096 | 255 | -1.3072 | 0.55096
| 270 | 19.943 | 19.9476 | 19.6767 | 0.056667 | -1.3583 | 0.55184 | 270 | -1.3583 | 0.55184
| 285 |  | 19.932 | 19.69 | 0.02 | -1.2143 | 0.5477 | 285 | -1.2143 | 0.5477
| 300 |  | 19.9192 | 19.63 | 0.04 | -1.4519 | 0.54904 | 300 | -1.4519 | 0.54904
| 315 |  | 19.9203 | 19.57 | 0.04 | -1.7586 | 0.54732 | 315 | -1.7586 | 0.54732
| 330 |  | 19.9193 | 19.55 | 0.04 | -1.8537 | 0.54683 | 330 | -1.8537 | 0.54683
| 345 |  | 19.9343 | 19.51 | 0.05 | -2.1283 | 0.54664 | 345 | -2.1283 | 0.54664
| 360 |  | 19.9303 | 19.51 | 0.05 | -2.1089 | 0.54685 | 360 | -2.1089 | 0.54685
| 375 |  | 19.9319 | 19.53 | 0.04 | -2.0163 | 0.54559 | 375 | -2.0163 | 0.54559
| 390 |  | 19.9171 | 19.53 | 0.04 | -1.9437 | 0.54639 | 390 | -1.9437 | 0.54639
| 405 |  | 19.9329 | 19.51 | 0.05 | -2.1217 | 0.54671 | 405 | -2.1217 | 0.54671
| 420 |  | 19.9299 | 19.52 | 0.03 | -2.0567 | 0.54407 | 420 | -2.0567 | 0.54407
435
450
465
480
495
510
525
45
DTI | 60 | [6 empty] | 60
| 75 |  | 20.0625 | 17.818 | 0.031052 | -11.1876 | 0.58766 | 75 | -11.1876 | 0.58766
| 90 |  | 20.0614 | 18.226 | 0.04109 | -9.1488 | 0.59318 | 90 | -9.1488 | 0.59318
| 105 |  | 20.0628 | 18.728 | 0.032216 | -6.6529 | 0.56099 | 105 | -6.6529 | 0.56099
| 120 |  | 20.0531 | 19.1585 | 0.029961 | -4.4614 | 0.54488 | 120 | -4.4614 | 0.54488
| 135 |  | 20.0442 | 19.46 | 0.031119 | -2.9147 | 0.53818 | 135 | -2.9147 | 0.53818
| 150 |  | 20.0233 | 19.634 | 0.023709 | -1.9441 | 0.52344 | 150 | -1.9441 | 0.52344
| 165 |  | 20.0319 | 19.742 | 0.021423 | -1.4471 | 0.51803 | 165 | -1.4471 | 0.51803
| 180 |  | 20.0527 | 19.7975 | 0.027121 | -1.2726 | 0.52336 | 180 | -1.2726 | 0.52336
| 195 |  | 20.0201 | 19.91 | 0.024061 | -0.55018 | 0.51659 | 195 | -0.55018 | 0.51659
| 210 |  | 20.0322 | 19.976 | 0.033309 | -0.28052 | 0.52764 | 210 | -0.28052 | 0.52764
| 225 |  | 20.0237 | 19.933 | 0.030625 | -0.4529 | 0.52468 | 225 | -0.4529 | 0.52468
| 240 |  | 20.0221 | 19.9045 | 0.036487 | -0.58714 | 0.5348 | 240 | -0.58714 | 0.5348
| 255 | 20.0125 | 20.0275 | 19.901 | 0.025526 | -0.63147 | 0.5186 | 255 | -0.63147 | 0.5186
| 270 |  | 20.0221 | 19.836 | 0.03939 | -0.92926 | 0.54183 | 270 | -0.92926 | 0.54183
| 285 |  | 20.0237 | 19.756 | 0.025629 | -1.3369 | 0.52253 | 285 | -1.3369 | 0.52253
| 300 |  | 20.0322 | 19.74 | 0.03095 | -1.4586 | 0.53029 | 300 | -1.4586 | 0.53029
| 315 |  | 20.0201 | 19.774 | 0.019574 | -1.2295 | 0.51531 | 315 | -1.2295 | 0.51531
| 330 |  | 20.0527 | 19.784 | 0.01667 | -1.34 | 0.51243 | 330 | -1.34 | 0.51243
| 345 |  | 20.0319 | 19.76 | 0.02753 | -1.3573 | 0.5249 | 345 | -1.3573 | 0.5249
| 360 |  | 20.0233 | 19.7045 | 0.027999 | -1.592 | 0.52702 | 360 | -1.592 | 0.52702
| 375 |  | 20.0442 | 19.6865 | 0.036023 | -1.7847 | 0.53991 | 375 | -1.7847 | 0.53991
| 390 |  | 20.0531 | 19.67 | 0.029736 | -1.9107 | 0.53039 | 390 | -1.9107 | 0.53039
| 405 |  | 20.0628 | 19.667 | 0.025976 | -1.9726 | 0.52534 | 405 | -1.9726 | 0.52534
| 420 |  | 20.0614 | 19.6725 | 0.032261 | -1.9384 | 0.53412 | 420 | -1.9384 | 0.53412
| 435 |  | 20.0625 | 19.6675 | 0.034774 | -1.9689 | 0.53832 | 435 | -1.9689 | 0.53832
| 450 | [6 empty] | 450
465
480
495
510
525
45
DWG | 50 | [6 empty] | 60
| 65 | [6 empty] | 75
| 80 |  | 20.072 | 19.3934 | 0.03523 | -3.3808 | 0.54834 | 90 | -2.6602 | 0.54874
| 95 |  | 20.044 | 19.583 | 0.015725 | -2.2999 | 0.54894 | 105 | -1.8495 | 0.54931
| 110 |  | 20.021 | 19.6958 | 0.023013 | -1.6243 | 0.5495 | 120 | -1.2471 | 0.54969
| 125 |  | 20.009 | 19.7972 | 0.022957 | -1.0585 | 0.54978 | 135 | -0.73187 | 0.54964
| 140 |  | 20.016 | 19.9022 | 0.017211 | -0.56855 | 0.54957 | 150 | -0.33643 | 0.54964
| 155 |  | 20.012 | 19.9679 | 0.017018 | -0.22037 | 0.54967 | 165 | -0.084786 | 0.54978
| 170 |  | 20.006 | 20.0026 | 0.01831 | -0.016995 | 0.54984 | 180 | -0.048652 | 0.54984
| 185 |  | 20.006 | 19.9931 | 0.018737 | -0.064481 | 0.54984 | 195 | -0.21477 | 0.54984
| 200 |  | 20.006 | 19.948 | 0.018684 | -0.28991 | 0.54984 | 210 | -0.5317 | 0.55001
| 215 |  | 19.997 | 19.8665 | 0.025114 | -0.6526 | 0.55009 | 225 | -0.97068 | 0.55013
| 230 |  | 19.996 | 19.7701 | 0.020472 | -1.1297 | 0.55015 | 240 | -1.3396 | 0.55023
| 245 |  | 19.992 | 19.7032 | 0.024274 | -1.4446 | 0.55028 | 255 | -1.6446 | 0.55028
| 260 |  | 19.993 | 19.6442 | 0.027377 | -1.7446 | 0.55028 | 270 | -1.8306 | 0.55036
| 275 |  | 19.989 | 19.6145 | 0.01484 | -1.8735 | 0.5504 | 285 | -1.8919 | 0.55053
| 290 |  | 19.9821 | 19.6022 | 0.01198 | -1.9011 | 0.55059 | 300 | -1.9872 | 0.55034
| 305 |  | 19.9961 | 19.5901 | 0.0097971 | -2.0303 | 0.55022 | 315 | -2.0632 | 0.55029
| 320 |  | 19.9926 | 19.5768 | 0.012394 | -2.0797 | 0.55032 | 330 | -2.0982 | 0.55028
| 335 |  | 19.9951 | 19.5737 | 0.01038 | -2.1074 | 0.55026 | 345 | -2.1357 | 0.55015
| 350 |  | 20.0011 | 19.5711 | 0.011997 | -2.1498 | 0.5501 | 360 | -2.1646 | 0.55014
| 365 |  | 19.9991 | 19.5647 | 0.01222 | -2.172 | 0.55015 | 375 | -2.1956 | 0.55022
| 380 |  | 19.9956 | 19.5542 | 0.016945 | -2.2074 | 0.55026 | 390 | -2.204 | 0.55023
| 395 |  | 19.9971 | 19.5567 | 0.011658 | -2.2022 | 0.55021 | 405 | -2.2181 | 0.55017
| 410 |  | 19.9996 | 19.5544 | 0.012162 | -2.226 | 0.55015 | 420 | -2.2672 | 0.55013
| 425 |  | 20.0006 | 19.543 | 0.012461 | -2.2878 | 0.55013 | 435 | -2.2624 | 0.55019
| 440 |  | 19.9971 | 19.5472 | 0.01125 | -2.2497 | 0.55022 | 450 | -2.2368 | 0.55031
| 455 |  | 19.9921 | 19.5462 | 0.012563 | -2.2303 | 0.55035 | 465 | -2.2433 | 0.55026
| 470 |  | 19.9971 | 19.5472 | 0.011282 | -2.2497 | 0.55022 | 480 | -2.2265 | 0.55023
| 485 |  | 19.9961 | 19.5532 | 0.014102 | -2.2149 | 0.55024 | 495 | -2.2132 | 0.55042
| 500 | 19.987 | 19.9867 | 19.5445 | 0.012697 | -2.2123 | 0.5505 | 510 | -2.2179 | 0.55038
| 515 |  | 19.9936 | 19.5496 | 0.013465 | -2.2206 | 0.55031 | 525 | -2.243 | 0.55031
| 530 |  | 19.9936 | 19.5429 | 0.011781 | -2.2541 | 0.55032
| 545 |  | 19.9992 | 19.5402 | 0.012173 | -2.2949 | 0.55017
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
| 89 |  | 20.365 | 20.2425 | 0.05665 | -0.60167 | 1.8 | 90 | -0.59126 | 1.8
| 104 |  | 20.3487 | 20.258 | 0.054445 | -0.44562 | 1.8 | 105 | -0.45546 | 1.8
| 119 |  | 20.3371 | 20.2165 | 0.055656 | -0.59321 | 1.8 | 120 | -0.66693 | 1.8
| 134 |  | 20.3238 | 19.9785 | 0.041202 | -1.6991 | 1.8 | 135 | -1.7193 | 1.8
| 149 |  | 20.308 | 19.9015 | 0.029069 | -2.0016 | 1.8 | 150 | -2.0053 | 1.8
| 164 |  | 20.2981 | 19.8805 | 0.040585 | -2.0572 | 1.8 | 165 | -2.0583 | 1.8
| 179 |  | 20.2947 | 19.874 | 0.030848 | -2.0729 | 1.8 | 180 | -2.0638 | 1.8
| 194 |  | 20.3011 | 19.908 | 0.02858 | -1.9362 | 1.8 | 195 | -1.9432 | 1.8
| 209 |  | 20.3086 | 19.894 | 0.035154 | -2.0417 | 1.8 | 210 | -2.0489 | 1.8
| 224 |  | 20.3224 | 19.8855 | 0.025644 | -2.1499 | 1.8 | 225 | -2.1517 | 1.8
| 239 |  | 20.3358 | 19.893 | 0.027739 | -2.1774 | 1.8 | 240 | -2.1735 | 1.8
| 254 |  | 20.3446 | 19.9135 | 0.026611 | -2.119 | 1.8 | 255 | -2.1292 | 1.8
| 269 |  | 20.3474 | 19.885 | 0.044898 | -2.2725 | 1.8 | 270 | -2.2673 | 1.8
| 284 |  | 20.3397 | 19.8935 | 0.029784 | -2.1938 | 1.8 | 285 | -2.1918 | 1.8
| 299 |  | 20.3185 | 19.879 | 0.036835 | -2.163 | 1.8 | 300 | -2.1707 | 1.8
| 314 |  | 20.2974 | 19.835 | 0.036201 | -2.2781 | 1.8 | 315 | -2.2826 | 1.8
| 329 | 20.2657 | 20.2657 | 19.7905 | 0.032359 | -2.3449 | 1.8 | 330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 42 |  | 20.1 | 18.535 | 0.025 | -7.786 | 0.49 | 45
| 57 |  | 20.082 | 19.208 | 0.022 | -4.352 | 0.49 | 60 | -4.0368 | 0.49
LEI | 72 |  | 20.068 | 19.511 | 0.02 | -2.776 | 0.49 | 75 | -2.6652 | 0.49
| 87 |  | 20.071 | 19.625 | 0.019 | -2.222 | 0.49 | 90 | -2.0858 | 0.49
| 102 |  | 20.054 | 19.745 | 0.015 | -1.541 | 0.49 | 105 | -1.3924 | 0.49
| 117 |  | 20.045 | 19.885 | 0.018 | -0.798 | 0.49 | 120 | -0.6564 | 0.49
| 132 |  | 20.049 | 20.031 | 0.019 | -0.09 | 0.49 | 135 | 0.0408 | 0.49
| 147 |  | 20.029 | 20.142 | 0.015 | 0.564 | 0.49 | 150 | 0.5948 | 0.49
| 162 |  | 20.057 | 20.201 | 0.016 | 0.718 | 0.49 | 165 | 0.666 | 0.49
| 177 |  | 20.095 | 20.187 | 0.018 | 0.458 | 0.49 | 180 | 0.3824 | 0.49
| 192 |  | 20.078 | 20.094 | 0.026 | 0.08 | 0.49 | 195 | -0.0676 | 0.49
| 207 |  | 20.071 | 19.939 | 0.037 | -0.658 | 0.49 | 210 | -0.6692 | 0.49
| 222 |  | 20.022 | 19.879 | 0.017 | -0.714 | 0.49 | 225 | -0.7898 | 0.49
| 237 |  | 20.03 | 19.811 | 0.017 | -1.093 | 0.49 | 240 | -1.0872 | 0.49
| 252 |  | 20.011 | 19.798 | 0.015 | -1.064 | 0.49 | 255 | -1.045 | 0.49
| 267 |  | 20.017 | 19.823 | 0.02 | -0.969 | 0.49 | 270 | -0.98 | 0.49
| 282 |  | 20.019 | 19.814 | 0.021 | -1.024 | 0.49 | 285 | -1.057 | 0.49
| 297 |  | 20.013 | 19.775 | 0.021 | -1.189 | 0.49 | 300 | -1.21 | 0.49
| 312 |  | 20.01 | 19.751 | 0.018 | -1.294 | 0.49 | 315 | -1.347 | 0.49
| 327 |  | 20.019 | 19.707 | 0.015 | -1.559 | 0.49 | 330 | -1.6918 | 0.49
| 342 |  | 20.018 | 19.573 | 0.015 | -2.223 | 0.49 | 345
| 357 |  | 20.026 | 19.146 | 0.026 | -4.394 | 0.49 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 0.6
| 0.6 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 9.1284
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 0.58478 | 0.652 | 0.0042164 | 11.4954 | 1.539 | 45 | 12.1595 | 1.5388
| 52.5 |  | 0.58498 | 0.66 | 0 | 12.8235 | 1.5385 | 60 | 12.9474 | 1.5402 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 0.5837 | 0.66 | 0 | 13.0713 | 1.5419 | 75 | 13.2212 | 1.5416 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 0.58392 | 0.662 | 0.0042164 | 13.371 | 1.5413 | 90 | 14.524 | 1.6275 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 0.58352 | 0.675 | 0.0052705 | 15.677 | 1.7137 | 105 | 16.1525 | 1.7119 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 0.58477 | 0.682 | 0.0042164 | 16.628 | 1.7101 | 120 | 17.6908 | 1.7106 |  | 0.015026 | 0.43444 |  | 0.96876 |  | 0.45461 |  | 0.96554 |  |  | 1.3848 | [7 empty] | 0.36391 | 0.38308 | 0.06732 |  | 1.3104
| 127.5 | 0.58464 | 0.5844 | 0.694 | 0.005164 | 18.7536 | 1.7111 | 135 | 19.9371 | 1.7122 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 0.58372 | 0.707 | 0.0048305 | 21.1207 | 1.7132 | 150 | 21.5199 | 1.7128 | 45 |  | 0.21886 |  | 0.52872 | [6 empty] | 0.73576
| 157.5 |  | 0.58399 | 0.712 | 0.0042164 | 21.9192 | 1.7123 | 165 | 22.524 | 1.6268 | 60 | 0.14569 | 0.18792 |  | 0.43597 |  | 0.33348 |  | 0.29807 |  |  | 0.6139
| 172.5 |  | 0.58394 | 0.719 | 0.0031623 | 23.1288 | 1.5413 | 180 | 23.1257 | 1.5401 | 75 | 0.61335 | 0.32923 |  | 0.42366 |  | 0.9376 |  | 0.16961 |  |  | 0.73933
| 187.5 |  | 0.58478 | 0.72 | 1.1703e-16 | 23.1225 | 1.539 | 195 | 22.8565 | 1.5368 | 90 | 0.03816 | 0.28893 |  | 0.53027 |  | 0.32944 |  | 0.49801 |  |  | 0.80676
| 202.5 |  | 0.58651 | 0.719 | 0.0031623 | 22.5905 | 1.5345 | 210 | 22.4955 | 1.5323 | 105 | 0.2666 | 0.30233 |  | 0.66037 |  | 0.041983 |  | 0.92878 |  |  | 0.9502
| 217.5 |  | 0.58823 | 0.72 | 1.1703e-16 | 22.4004 | 1.53 | 225 |  |  | 120 | 0.20606 | 0.1806 |  | 0.61966 |  | 0.021253 |  | 0.82828 |  |  | 0.79065 | [7 empty] | 0.20224 | 0.40861 | 0.37917 |  | 0.4159
[8 empty] | 240 |  |  | 135 | 0.11808 | 0.14091 |  | 0.55707 |  | 0.25956 |  | 0.44826 |  |  | 0.68987 | [7 empty] | 0.36035 | 0.24885 | 0.49538 |  | 0.19437
[8 empty] | 255 |  |  | 150 | 0.57255 | 0.0086244 |  |  |  | 0.57585 | [12 empty] | 0.41319 | 0.14791 | 0.41798
[8 empty] | 270 |  |  | 165 | 0.27802 | 0.081927 |  |  |  | 0.19212 | [12 empty] | 0.34165 | 0.07332 | 0.25788
[8 empty] | 285 |  |  | 180 | 0.14671 | 0.078036 |  |  |  | 0.066079 | [12 empty] | 0.3547 | 0.21369 | 0.27359
[8 empty] | 300 |  |  | 195 | 0.22898 | 0.06852 |  |  |  | 0.15691 | [12 empty] | 0.33098 | 0.10896 | 0.25958
[8 empty] | 315 |  |  | 210 | 0.070217 | 0.029925 |  |  |  | 0.099513 | [12 empty] | 0.14609 | 0.078373 | 0.17375
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.14557 | [13 empty] | 0.20531 | 0.34282
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.052389 | [13 empty] | 0.047689 | 0.097763
[8 empty] | 360 |  |  | 255 | [5 empty] | 0.068046 | [13 empty] | 0.081008 | 0.01299
[8 empty] | 375 |  |  | 270 | [5 empty] | 0.29756 | [13 empty] | 0.059225 | 0.23153
[8 empty] | 390 |  |  | 285 | [20 empty] | 0.57265
[8 empty] | 405 |  |  | 300 | [20 empty] | 0.28234
[8 empty] | 420 |  |  | 315 | [20 empty] | 0.045586
[8 empty] | 435 |  |  | 330 | [20 empty] | 0.012667
[8 empty] | 450 |  |  | 345 | [20 empty] | 0.24478
[8 empty] | 465 |  |  | 360 | [20 empty] | 0.1787
[8 empty] | 480 |  |  | 375 | [20 empty] | 0.0010781
[8 empty] | 495 |  |  | 390 | [20 empty] | 0.18298
[8 empty] | 510 |  |  | 405 | [20 empty] | 0.40756
[8 empty] | 525 |  |  | 420
435
450
45
PTB | 55 | 0.59757 | 0.59951 | 0.688 | 0.0070711 | 14.7613 | 0.85 | 60 | 11.5929 | 0.86333
| 70 | 0.59757 | 0.59925 | 0.63074 | 0.010715 | 5.2559 | 0.89 | 75 | 7.5142 | 0.93333
| 85 | 0.59757 | 0.59876 | 0.6708 | 0.019983 | 12.0308 | 1.02 | 90 | 14.1683 | 0.96667
| 100 | 0.59757 | 0.59788 | 0.70815 | 0.0087868 | 18.4431 | 0.86 | 105 | 18.6391 | 0.85667
| 115 | 0.59757 | 0.59804 | 0.71185 | 0.0078628 | 19.031 | 0.85 | 120 | 19.6125 | 0.84667
| 130 | 0.59757 | 0.59814 | 0.7224 | 0.0052281 | 20.7754 | 0.84 | 135 | 18.8355 | 0.88667
| 145 | 0.59757 | 0.59822 | 0.68769 | 0.018613 | 14.9557 | 0.98 | 150 | 16.1749 | 0.94333
| 160 | 0.59757 | 0.59757 | 0.7088 | 0.0097125 | 18.6134 | 0.87 | 165 | 19.9352 | 0.85667
| 175 | 0.59757 | 0.59651 | 0.7312 | 0.004397 | 22.5787 | 0.83 | 180 | 21.762 | 0.83333
| 190 | 0.59757 | 0.59703 | 0.7172 | 0.0061373 | 20.1286 | 0.84 | 195 | 20.728 | 0.84667
| 205 | 0.59757 | 0.59777 | 0.72885 | 0.0090893 | 21.927 | 0.86 | 210 | 21.8428 | 0.85
| 220 | 0.59757 | 0.59859 | 0.72833 | 0.0038069 | 21.6744 | 0.83 | 225 | 21.5244 | 0.83333
| 235 | 0.59757 | 0.59889 | 0.726 | 0.0057735 | 21.2243 | 0.84 | 240 | 21.7101 | 0.84
| 250 | 0.59757 | 0.5996 | 0.7356 | 0.0050662 | 22.6816 | 0.84 | 255 | 22.7045 | 0.84
| 265 | 0.59757 | 0.60024 | 0.7368 | 0.004761 | 22.7501 | 0.84 | 270 | 22.8044 | 0.83667
| 280 | 0.59757 | 0.60238 | 0.7404 | 0.0035119 | 22.9129 | 0.83 | 285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 0.54097 | 0.618 | 0.0041039 | 14.2392 | 2.1467 | 45 | 14.2392 | 2.1467
| 60 |  | 0.54387 | 0.624 | 0.0050262 | 14.7333 | 2.1467 | 60 | 14.7333 | 2.1467
| 75 |  | 0.54147 | 0.63 | 2.2781e-16 | 16.3499 | 2.1467 | 75 | 16.3499 | 2.1467
| 90 |  | 0.54232 | 0.636 | 0.0050262 | 17.2739 | 2.1467 | 90 | 17.2739 | 2.1467
| 105 |  | 0.54102 | 0.644 | 0.0050262 | 19.0344 | 2.1467 | 105 | 19.0344 | 2.1467
| 120 |  | 0.54182 | 0.647 | 0.0047016 | 19.4124 | 2.1467 | 120 | 19.4124 | 2.1467
| 135 |  | 0.54172 | 0.657 | 0.0047016 | 21.2804 | 2.1467 | 135 | 21.2804 | 2.1467
| 150 |  | 0.54152 | 0.6585 | 0.0036635 | 21.6022 | 2.1467 | 150 | 21.6022 | 2.1467
| 165 |  | 0.54212 | 0.66 | 0 | 21.7443 | 2.1467 | 165 | 21.7443 | 2.1467
| 180 |  | 0.54092 | 0.662 | 0.0041039 | 22.3841 | 2.1467 | 180 | 22.3841 | 2.1467
| 195 |  | 0.54212 | 0.6625 | 0.0044426 | 22.2054 | 2.1467 | 195 | 22.2054 | 2.1467
| 210 |  | 0.54162 | 0.665 | 0.0051299 | 22.7798 | 2.1467 | 210 | 22.7798 | 2.1467
| 225 | 0.54152 | 0.54152 | 0.6655 | 0.0051042 | 22.8948 | 2.1467 | 225 | 22.8948 | 2.1467
| 240 |  | 0.54172 | 0.662 | 0.0041039 | 22.2034 | 2.1467 | 240 | 22.2034 | 2.1467
| 255 |  | 0.54152 | 0.661 | 0.0030779 | 22.0638 | 2.1467 | 255 | 22.0638 | 2.1467
| 270 |  | 0.54082 | 0.649 | 0.010712 | 20.003 | 2.1467 | 270 | 20.003 | 2.1467
| 285 |  | 0.54082 | 0.6295 | 0.011459 | 16.3973 | 2.1467 | 285 | 16.3973 | 2.1467
| 300 |  | 0.54092 | 0.647 | 0.0047016 | 19.611 | 2.1467 | 300 | 19.611 | 2.1467
| 315 |  | 0.54062 | 0.6595 | 0.0022361 | 21.9896 | 2.1467 | 315 | 21.9896 | 2.1467
| 330 |  | 0.54192 | 0.66 | 0 | 21.7892 | 2.1467 | 330 | 21.7892 | 2.1467
| 345 |  | 0.54072 | 0.66 | 0 | 22.0595 | 2.1467 | 345 | 22.0595 | 2.1467
| 360 |  | 0.54212 | 0.6585 | 0.0036635 | 21.4676 | 2.1467 | 360 | 21.4676 | 2.1467
| 375 |  | 0.54137 | 0.659 | 0.0030779 | 21.7282 | 2.1467 | 375 | 21.7282 | 2.1467
| 390 |  | 0.54317 | 0.656 | 0.0050262 | 20.7725 | 2.1467 | 390 | 20.7725 | 2.1467
| 405 |  | 0.54987 | 0.6415 | 0.0074516 | 16.6639 | 2.1467 | 405 | 16.6639 | 2.1467
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 0.60469 | 0.7 | 0 | 15.761 | 2.1912 | 120 | 15.761 | 2.1912
| 135 |  | 0.60532 | 0.705 | 0.005 | 16.4669 | 2.5457 | 135 | 16.4669 | 2.5457
| 150 |  | 0.60388 | 0.71 | 0.005 | 17.5722 | 2.241 | 150 | 17.5722 | 2.241
| 165 |  | 0.59965 | 0.715 | 0.01 | 19.2371 | 2.5651 | 165 | 19.2371 | 2.5651
| 180 |  | 0.60129 | 0.72 | 0.005 | 19.7431 | 2.2899 | 180 | 19.7431 | 2.2899
| 195 |  | 0.60152 | 0.72 | 0.01 | 19.6975 | 2.3257 | 195 | 19.6975 | 2.3257
| 210 |  | 0.6028 | 0.73 | 0.005 | 21.1019 | 2.3087 | 210 | 21.1019 | 2.3087
| 225 |  | 0.60209 | 0.72 | 0.01 | 19.583 | 2.3214 | 225 | 19.583 | 2.3214
| 240 |  | 0.60202 | 0.73 | 0 | 21.2594 | 2.3021 | 240 | 21.2594 | 2.3021
| 255 |  | 0.59866 | 0.73 | 0 | 21.9383 | 2.3274 | 255 | 21.9383 | 2.3274
| 270 | 0.599 | 0.59717 | 0.73 | 0.01 | 22.2434 | 2.3882 | 270 | 22.2434 | 2.3882
| 285 |  | 0.59867 | 0.73 | 0.01 | 21.9371 | 2.3767 | 285 | 21.9371 | 2.3767
| 300 |  | 0.59669 | 0.73 | 0.005 | 22.3409 | 2.3549 | 300 | 22.3409 | 2.3549
| 315 |  | 0.59626 | 0.73 | 0.005 | 22.4304 | 2.3582 | 315 | 22.4304 | 2.3582
| 330 |  | 0.59589 | 0.725 | 0.01 | 21.6659 | 2.62 | 330 | 21.6659 | 2.62
| 345 |  | 0.60152 | 0.72 | 0 | 19.6975 | 2.2756 | 345 | 19.6975 | 2.2756
| 360 |  | 0.60129 | 0.72 | 0 | 19.7431 | 2.2773 | 360 | 19.7431 | 2.2773
| 375 |  | 0.59965 | 0.73 | 0 | 21.7386 | 2.3199 | 375 | 21.7386 | 2.3199
| 390 |  | 0.60388 | 0.74 | 0 | 22.54 | 2.3182 | 390 | 22.54 | 2.3182
| 405 |  | 0.60532 | 0.73 | 0 | 20.5969 | 2.2776 | 405 | 20.5969 | 2.2776
| 420 |  | 0.60469 | 0.73 | 0.005 | 20.7221 | 2.2947 | 420 | 20.7221 | 2.2947
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
UL-LMPS | 42.5 | 0.60138 | 0.60238 | 0.643 | 0.0053498 | 6.7435 | 2.6657 | 45 | 7.0657 | 2.6687
| 57.5 | 0.5978 | 0.59473 | 0.64633 | 0.0049013 | 8.6768 | 2.6838 | 60 | 8.745 | 2.6861
| 72.5 | 0.59584 | 0.59036 | 0.644 | 0.0049827 | 9.0858 | 2.6974 | 75 | 9.1363 | 2.6947
| 87.5 | 0.59116 | 0.58873 | 0.644 | 0.0049827 | 9.3889 | 2.6811 | 90 | 9.4065 | 2.6759
| 102.5 | 0.59816 | 0.59577 | 0.65233 | 0.0043018 | 9.4946 | 2.65 | 105 | 9.7734 | 2.6562
| 117.5 | 0.60032 | 0.5961 | 0.66267 | 0.0044978 | 11.1672 | 2.6871 | 120 | 11.7 | 2.6856
| 132.5 | 0.59008 | 0.59022 | 0.675 | 0.0062972 | 14.364 | 2.6782 | 135 | 14.5528 | 2.6762
| 147.5 | 0.58838 | 0.58732 | 0.67833 | 0.0037905 | 15.4969 | 2.6661 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 2
| 2 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 3.1109
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 1.9679 | 1.919 | 0.0073786 | -2.487 | 0.86385 | 45 | -2.2454 | 0.86418
| 52.5 |  | 1.9664 | 1.927 | 0.0082327 | -2.0038 | 0.86452 | 60 | -1.3771 | 0.86356 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 1.9708 | 1.956 | 0.0084327 | -0.7504 | 0.8626 | 75 | -0.22182 | 0.86234 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 1.972 | 1.978 | 0.0063246 | 0.30676 | 0.86209 | 90 | 0.57473 | 0.862 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 1.9724 | 1.989 | 0.0056765 | 0.8427 | 0.8619 | 105 | 1.159 | 0.86158 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 1.9739 | 2.003 | 0.0048305 | 1.4753 | 0.86125 | 120 | 1.6711 | 0.8614 |  | 0.16615 | 1.0599 | 1.1611 | 0.040675 | 0.60824 | 1.2507 | 1.317 | 0.12633 | 0.77787 | 0.2797 | 1.1118 | 0.39536 | 1.2033 | 0.60299 | 0.65245 |  |  | 0.16428 | 0.65839 | 0.83725 | 0.38827 | 0.60109 | 0.70528 | 0.022879 | [6 empty] | 0.47843 | 0.55783 | 0.019638 | 0.1691 | 0.49856 | 0.17401
| 127.5 | 1.9738 | 1.9732 | 2.01 | 0.008165 | 1.8669 | 0.86156 | 135 | 1.8355 | 0.8613 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 1.9744 | 2.01 | 0.004714 | 1.8041 | 0.86103 | 150 | 2.2397 | 0.8615 | 45 |  | 0.3523 |  | 0.03356 | 0.70101 | [5 empty] | 0.38929 | 0.36949 |  |  | 0.73867
| 157.5 |  | 1.9722 | 2.025 | 0.0070711 | 2.6753 | 0.86197 | 165 | 2.5859 | 0.86185 | 60 | 0.033683 | 0.23831 | 0.31497 | 0.13783 | 0.66437 | 0.27728 | 0.34761 | 0.10702 | 0.70902 | 0.11694 | 0.37733 | 0.44088 | 0.43068 | 0.26039 | 0.79662 | [15 empty] | 0.04021 | 0.056039 | 0.071856 | 0.13143 | 0.10518 | 0.28659
| 172.5 |  | 1.9728 | 2.022 | 0.0063246 | 2.4965 | 0.86174 | 180 | 2.5683 | 0.86171 | 75 | 0.16322 | 0.31483 | 0.1394 | 0.20512 | 0.70098 | 0.48582 | 0.27736 | 0.37477 | 0.56022 | 0.12354 | 0.11215 | 1.0026 | 0.03085 | 0.73037 | 0.90226 | [15 empty] | 0.18213 | 0.25874 | 0.034219 | 0.098999 | 0.086527 | 0.52424
| 187.5 |  | 1.9729 | 2.025 | 0.0052705 | 2.6402 | 0.86167 | 195 | 2.7215 | 0.86193 | 90 | 0.17658 | 0.45725 | 0.49527 | 0.39063 | 0.40634 | 0.64497 | 0.64987 | 0.57805 | 0.2472 | 0.11433 | 0.069649 | 0.84198 | 0.17269 | 0.8228 | 0.78069 | [15 empty] | 0.36635 | 0.45004 | 0.15163 | 0.083394 | 0.18429 | 0.56154
| 202.5 |  | 1.9717 | 2.027 | 0.0048305 | 2.8027 | 0.86218 | 210 | 2.8216 | 0.86192 | 105 | 0.26082 | 0.54044 | 0.66493 | 0.45295 | 0.27208 | 0.81465 | 0.89218 | 0.72708 | 0.030426 | 0.21504 | 0.091363 | 0.78612 | 0.29192 | 0.87213 | 0.70438 | [15 empty] | 0.381 | 0.50384 | 0.12718 | 0.0059441 | 0.16992 | 0.51008
| 217.5 |  | 1.973 | 2.029 | 0.0056765 | 2.8404 | 0.86165 | 225 |  |  | 120 | 0.35713 | 0.42625 | 0.65722 | 0.47302 | 0.21897 | 0.79462 | 0.96451 | 0.84367 | 0.11499 | 0.30272 | 0.045664 | 0.6244 | 0.26539 | 0.81957 | 0.66927 |  |  | 0.5015 | 0.26266 | 0.083845 | 0.68436 | 0.87195 | 0.73127 | 0.033366 | [6 empty] | 0.37738 | 0.54487 | 0.17726 | 0.0067137 | 0.15605 | 0.48041
[8 empty] | 240 |  |  | 135 | 0.20971 | 0.6418 | 0.77318 | 0.45828 | 0.071296 | 0.86693 | 0.9595 | 0.67606 | 0.12596 | 0.23907 | 0.17572 | 0.68072 | 0.38367 | 0.80579 | 0.50625 |  |  | 0.51837 | 0.12233 | 0.081716 | 0.75573 | 0.86848 | 0.57334 | 0.045902 | [6 empty] | 0.46145 | 0.56097 | 0.16004 | 0.024363 | 0.24271 | 0.49096
[8 empty] | 255 |  |  | 150 | 0.28651 | 0.59898 | 0.83863 |  | 0.021967 | 0.9003 | 1.0896 |  | 0.29386 | 0.34101 |  | 0.54645 |  | 0.78972 |  |  |  | 0.66001 | 0.32132 | 0.045872 | 0.91341 | 1.1004 |  | 0.32721 | [6 empty] | 0.50831 | 0.64368 | 0.22708 | 0.033414 |  | 0.49146
[8 empty] | 270 |  |  | 165 | 0.36934 | 0.50934 | 0.90885 |  | 0.081979 | 0.89188 | 1.2296 |  | 0.43356 | 0.48643 |  | 0.40114 |  | 0.80651 |  |  |  | 0.62834 | 0.24155 | 0.11976 | 0.74835 | 1.1067 |  | 0.31099 | [6 empty] | 0.51452 | 0.68836 | 0.27545 | 0.001339 |  | 0.46812
[8 empty] | 285 |  |  | 180 | 0.30784 | 0.64589 | 1.0107 |  | 0.25825 | 0.96972 | 1.2819 |  | 0.5551 | 0.47456 |  | 0.35376 |  | 0.75506 |  |  |  | 0.46119 | 0.1702 | 0.47834 | 0.46801 | 0.86036 |  | 0.093982 | [6 empty] | 0.54349 | 0.68894 | 0.24023 | 0.025904 |  | 0.41018
[8 empty] | 300 |  |  | 195 | 0.30713 | 0.66895 | 1.0612 |  | 0.33012 | 0.99259 | 1.3327 |  | 0.62774 | 0.50604 |  | 0.3035 |  | 0.74282 |  |  |  | 0.34312 | 0.43232 | 0.74522 | 0.22702 | 0.69004 |  | 0.084281 | [6 empty] | 0.55043 | 0.69558 | 0.23632 | 0.046547 |  | 0.38173
[8 empty] | 315 |  |  | 210 | 0.32725 | 0.66757 | 1.116 |  | 0.37262 | 1.0114 | 1.4049 |  | 0.69017 | 0.56227 |  | 0.25951 |  | 0.75964 |  |  |  | 0.275 | 0.37363 | 0.70518 | 0.28471 | 0.79382 |  | 0.013643 | [6 empty] | 0.45406 | 0.60808 | 0.14051 | 0.16827 |  | 0.26564
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.1033 | 1.3886 |  | 0.73496 | 0.47094 |  | 0.30072 |  | 0.70674 |  |  |  | 0.16678 |  | 0.82697 | 0.25652 | 0.68064 |  | 0.054373 | [7 empty] | 0.55516 | 0.045532 | 0.20975 |  | 0.19204
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.0949 | 1.4159 |  | 0.67168 | 0.50477 |  | 0.35524 |  | 0.78531 |  |  |  | 0.17614 |  | 0.69351 | 0.37985 | 0.81712 |  | 0.0083737 | [7 empty] | 0.502 | 0.0035861 | 0.27483 |  | 0.17023
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.1017 | 1.3953 |  | 0.62217 | 0.47907 |  | 0.41024 |  | 0.80706 |  |  |  | 0.1229 |  | 0.68659 | 0.39388 | 0.80346 |  | 0.032956 | [7 empty] | 0.44493 | 0.063537 | 0.31823 |  | 0.13796
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.47384 |  | 0.39152 |  | 0.7862 |  |  |  | 0.18453 |  |  | 0.39896 | 0.8021 |  | 0.0089809 | [8 empty] | 0.0043983 | 0.259 |  | 0.18718
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.38777 |  | 0.54132 |  | 0.83016 |  |  |  | 0.19844 |  |  | 0.44955 | 0.75919 |  | 0.10901 | [8 empty] | 0.014473 | 0.2223 |  | 0.25051
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.49203 |  | 0.37305 |  | 0.78817 |  |  |  | 0.052563 |  |  | 0.069066 | 0.54408 |  | 0.30217 | [8 empty] | 0.019875 | 0.24554 |  | 0.2021
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.57207 |  | 0.24484 |  | 0.75672 |  |  |  | 0.12536 |  |  | 0.082778 | 0.63541 |  | 0.16331 | [8 empty] | 0.086432 | 0.22503 |  | 0.20508
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.91623 |  | 0.13425 |  | 0.76696 |  |  |  | 0.16548 |  |  | 0.33878 | 0.62091 |  | 0.18878 | [8 empty] | 0.32787 | 0.17964 |  | 0.25797
[8 empty] | 450 |  |  | 345 | [9 empty] | 1.0082 |  | 0.1177 |  | 0.86927 |  |  |  | 0.22797 |  |  | 0.40136 | 0.6566 |  | 0.26478 | [8 empty] | 0.42132 | 0.13984 |  | 0.3583
[8 empty] | 465 |  |  | 360 | [9 empty] | 1.3077 | [7 empty] | 0.23632 |  |  | 0.73596 | 0.6755 | [10 empty] | 0.5884 | 0.14134
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.4133 | [10 empty] | 0.94212 | 0.60853
[8 empty] | 495 |  |  | 390 | [9 empty] | 1.3937 | [10 empty] | 1.0289 | 0.51312
[8 empty] | 510 |  |  | 405 | [9 empty] | 1.0181 | [10 empty] | 1.3005 | 0.078981
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.18105
435
450
45
PTB | 55 | 1.9887 | 1.9953 | 1.9678 | 0.020255 | -1.3786 | 0.54 | 60 | -1.2669 | 0.53333
| 70 | 1.9887 | 1.9943 | 1.9735 | 0.012631 | -1.0435 | 0.52 | 75 | -0.75541 | 0.51667
| 85 | 1.9887 | 1.9932 | 1.9896 | 0.0067577 | -0.1793 | 0.51 | 90 | -0.002342 | 0.51
| 100 | 1.9887 | 1.9904 | 1.9974 | 0.0085901 | 0.35158 | 0.51 | 105 | 0.30661 | 0.51
| 115 | 1.9887 | 1.9909 | 1.9952 | 0.0084582 | 0.21668 | 0.51 | 120 | 0.50402 | 0.51
| 130 | 1.9887 | 1.9896 | 2.0111 | 0.0050637 | 1.0787 | 0.51 | 135 | 1.1502 | 0.51
| 145 | 1.9887 | 1.9885 | 2.0142 | 0.0064331 | 1.2931 | 0.51 | 150 | 1.3034 | 0.51
| 160 | 1.9887 | 1.9887 | 2.015 | 0.0076158 | 1.3239 | 0.51 | 165 | 1.3789 | 0.51
| 175 | 1.9887 | 1.9896 | 2.0193 | 0.0078082 | 1.4887 | 0.51 | 180 | 1.5623 | 0.51
| 190 | 1.9887 | 1.9886 | 2.0226 | 0.0071213 | 1.7095 | 0.51 | 195 | 1.7177 | 0.51
| 205 | 1.9887 | 1.9892 | 2.0237 | 0.0062929 | 1.7343 | 0.51 | 210 | 1.7521 | 0.51
| 220 | 1.9887 | 1.9896 | 2.0252 | 0.0057981 | 1.7876 | 0.51 | 225 | 1.8036 | 0.51
| 235 | 1.9887 | 1.9896 | 2.0262 | 0.006373 | 1.8355 | 0.51 | 240 | 1.8348 | 0.51
| 250 | 1.9887 | 1.9923 | 2.0288 | 0.0090893 | 1.8335 | 0.51 | 255 | 1.908 | 0.51
| 265 | 1.9887 | 1.9951 | 2.0362 | 0.010612 | 2.0568 | 0.51 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 1.976 | 1.9085 | 0.01268 | -3.4175 | 0.80271 | 45 | -3.4175 | 0.80271
| 60 |  | 1.983 | 1.94 | 0.027338 | -2.1699 | 0.80271 | 60 | -2.1699 | 0.80271
| 75 |  | 1.9831 | 1.9995 | 0.013563 | 0.82546 | 0.80271 | 75 | 0.82546 | 0.80271
| 90 |  | 1.9883 | 2.03 | 0.0056195 | 2.0957 | 0.80271 | 90 | 2.0957 | 0.80271
| 105 |  | 1.9877 | 2.0465 | 0.0067082 | 2.9566 | 0.80271 | 105 | 2.9566 | 0.80271
| 120 |  | 1.99 | 2.0515 | 0.0067082 | 3.0889 | 0.80271 | 120 | 3.0889 | 0.80271
| 135 |  | 1.989 | 2.068 | 0.0076777 | 3.9703 | 0.80271 | 135 | 3.9703 | 0.80271
| 150 |  | 1.9912 | 2.0755 | 0.0068633 | 4.2321 | 0.80271 | 150 | 4.2321 | 0.80271
| 165 |  | 1.9922 | 2.0775 | 0.0071635 | 4.2801 | 0.80271 | 165 | 4.2801 | 0.80271
| 180 |  | 1.9901 | 2.084 | 0.0050262 | 4.7168 | 0.80271 | 180 | 4.7168 | 0.80271
| 195 |  | 1.9886 | 2.087 | 0.008645 | 4.9466 | 0.80271 | 195 | 4.9466 | 0.80271
| 210 |  | 1.9906 | 2.091 | 0.0055251 | 5.0421 | 0.80271 | 210 | 5.0421 | 0.80271
| 225 | 1.9892 | 1.9892 | 2.0965 | 0.0058714 | 5.3925 | 0.80271 | 225 | 5.3925 | 0.80271
| 240 |  | 1.9877 | 2.095 | 0.0051299 | 5.3966 | 0.80271 | 240 | 5.3966 | 0.80271
| 255 |  | 1.9897 | 2.099 | 0.0055251 | 5.4917 | 0.80271 | 255 | 5.4917 | 0.80271
| 270 |  | 1.9871 | 2.095 | 0.0076089 | 5.4284 | 0.80271 | 270 | 5.4284 | 0.80271
| 285 |  | 1.9829 | 2.0935 | 0.0058714 | 5.5788 | 0.80271 | 285 | 5.5788 | 0.80271
| 300 |  | 1.9851 | 2.0845 | 0.0060481 | 5.0084 | 0.80271 | 300 | 5.0084 | 0.80271
| 315 |  | 1.9854 | 2.076 | 0.0082078 | 4.5644 | 0.80271 | 315 | 4.5644 | 0.80271
| 330 |  | 1.9859 | 2.0455 | 0.028741 | 3.0022 | 0.80271 | 330 | 3.0022 | 0.80271
| 345 |  | 1.9859 | 2.037 | 0.028855 | 2.5742 | 0.80271 | 345 | 2.5742 | 0.80271
| 360 |  | 1.9836 | 2.0115 | 0.01387 | 1.4076 | 0.80271 | 360 | 1.4076 | 0.80271
| 375 |  | 1.9902 | 2.007 | 0.011743 | 0.84262 | 0.80271 | 375 | 0.84262 | 0.80271
| 390 |  | 1.9814 | 1.9995 | 0.0051042 | 0.91197 | 0.80271 | 390 | 0.91197 | 0.80271
| 405 |  | 1.9622 | 2.007 | 0.010809 | 2.2816 | 0.80271 | 405 | 2.2816 | 0.80271
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 2.0093 | 2.025 | 0.02 | 0.78185 | 1.021 | 120 | 0.78185 | 1.021
| 135 |  | 2.0065 | 2.035 | 0.02 | 1.4211 | 1.0276 | 135 | 1.4211 | 1.0276
| 150 |  | 2.0069 | 2.03 | 0.03 | 1.1513 | 1.0266 | 150 | 1.1513 | 1.0266
| 165 |  | 2.0045 | 2.04 | 0.01 | 1.7731 | 0.94926 | 165 | 1.7731 | 0.94926
| 180 |  | 2.0069 | 2.07 | 0.02 | 3.1431 | 0.99206 | 180 | 3.1431 | 0.99206
| 195 |  | 2.0061 | 2.09 | 0.02 | 4.1826 | 1.0014 | 195 | 4.1826 | 1.0014
| 210 |  | 2.008 | 2.09 | 0.02 | 4.0842 | 0.99963 | 210 | 4.0842 | 0.99963
| 225 |  | 2.009 | 2.1 | 0.01 | 4.5315 | 0.97163 | 225 | 4.5315 | 0.97163
| 240 |  | 2.0073 | 2.09 | 0.01 | 4.122 | 0.96881 | 240 | 4.122 | 0.96881
| 255 |  | 2.0063 | 2.09 | 0.005 | 4.1708 | 0.96166 | 255 | 4.1708 | 0.96166
| 270 | 2.007 | 2.0079 | 2.09 | 0.015 | 4.0882 | 0.98145 | 270 | 4.0882 | 0.98145
| 285 |  | 2.0083 | 2.09 | 0.015 | 4.0686 | 0.98109 | 285 | 4.0686 | 0.98109
| 300 |  | 2.0043 | 2.1 | 0.02 | 4.7758 | 1.0075 | 300 | 4.7758 | 1.0075
| 315 |  | 2.0041 | 2.09 | 0.015 | 4.2862 | 0.98506 | 315 | 4.2862 | 0.98506
| 330 |  | 2.002 | 2.085 | 0.015 | 4.1456 | 1.0345 | 330 | 4.1456 | 1.0345
| 345 |  | 2.0061 | 2.085 | 0.02 | 3.9334 | 1.071 | 345 | 3.9334 | 1.071
| 360 |  | 2.0069 | 2.085 | 0.015 | 3.8905 | 1.0299 | 360 | 3.8905 | 1.0299
| 375 |  | 2.0045 | 2.085 | 0.01 | 4.0181 | 1.0194 | 375 | 4.0181 | 1.0194
| 390 |  | 2.0069 | 2.095 | 0.02 | 4.3902 | 1.0514 | 390 | 4.3902 | 1.0514
| 405 |  | 2.0065 | 2.14 | 0.01 | 6.6542 | 0.99173 | 405 | 6.6542 | 0.99173
| 420 |  | 2.0093 | 2.15 | 0.01 | 7.003 | 0.99345 | 420 | 7.003 | 0.99345
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 |  | 2.1053 | 50 | 0.031128 | -5.0459 | 2.3779 | 60 | -2.6357 | 2.3546
| 65 |  | 2.1343 | 65 | 0.026351 | -1.4306 | 2.3429 | 75 | 0.33422 | 2.3428
| 80 |  | 2.1343 | 80 | 0.019441 | 1.2166 | 2.3428 | 90 | 2.5504 | 2.3435
| 95 |  | 2.1343 | 95 | 0.010713 | 3.2172 | 2.3439 | 105 | 3.812 | 2.3451
| 110 |  | 2.1333 | 110 | 0.0095455 | 4.1094 | 2.3457 | 120 | 4.2928 | 2.3437
| 125 |  | 2.1363 | 125 | 0.0085065 | 4.3845 | 2.3427 | 135 | 4.9216 | 2.3477
| 140 |  | 2.1303 | 140 | 0.007596 | 5.1901 | 2.3502 | 150 | 5.5903 | 2.3544
| 155 |  | 2.1253 | 155 | 0.0081374 | 5.7905 | 2.3565 | 165 | 6.2182 | 2.3564
| 170 |  | 2.1263 | 170 | 0.0076761 | 6.432 | 2.3563 | 180 | 6.6076 | 2.3566
| 185 |  | 2.1263 | 185 | 0.0078262 | 6.6954 | 2.3567 | 195 | 6.9628 | 2.3564
| 200 |  | 2.1273 | 200 | 0.0055356 | 7.0965 | 2.3563 | 210 | 7.2829 | 2.3581
| 215 |  | 2.1253 | 215 | 0.0057967 | 7.3761 | 2.359 | 225 | 7.2689 | 2.3566
| 230 |  | 2.1283 | 230 | 0.0067154 | 7.2153 | 2.3554 | 240 | 7.408 | 2.3572
| 245 |  | 2.1263 | 245 | 0.00639 | 7.5043 | 2.3581 | 255 | 7.4008 | 2.3579
| 260 |  | 2.1263 | 260 | 0.0057329 | 7.3491 | 2.3578 | 270 | 7.3169 | 2.3585
| 275 |  | 2.1253 | 275 | 0.0065666 | 7.3008 | 2.3588 | 285 | 7.1239 | 2.3571
| 290 |  | 2.1273 | 290 | 0.0066382 | 7.0354 | 2.3562 | 300 | 6.9685 | 2.3557
| 305 |  | 2.1278 | 305 | 0.0066788 | 6.9351 | 2.3555 | 315 | 6.8431 | 2.3549
| 320 |  | 2.1283 | 320 | 0.0064022 | 6.7972 | 2.3547 | 330 | 6.652 | 2.3552
| 335 |  | 2.1273 | 335 | 0.0060427 | 6.5794 | 2.3554 | 345 | 6.5899 | 2.3547
| 350 |  | 2.1283 | 350 | 0.0052661 | 6.5951 | 2.3544 | 360 | 6.616 | 2.354
| 365 |  | 2.1288 | 365 | 0.0058258 | 6.6265 | 2.3539 | 375 | 6.4719 | 2.3544
| 380 |  | 2.1278 | 380 | 0.0060519 | 6.3946 | 2.3546 | 390 | 6.4635 | 2.3547
| 395 |  | 2.1278 | 395 | 0.0048724 | 6.498 | 2.3548 | 405 | 6.3363 | 2.3538
| 410 |  | 2.1288 | 410 | 0.0053498 | 6.2554 | 2.3533 | 420 | 6.2742 | 2.3533
| 425 |  | 2.1288 | 425 | 0.0048466 | 6.2836 | 2.3533 | 435 | 6.3577 | 2.3531
| 440 |  | 2.1293 | 440 | 0.0057926 | 6.3948 | 2.3529 | 450 | 6.4511 | 2.3541
| 455 |  | 2.1278 | 455 | 0.0049508 | 6.4792 | 2.3547 | 465 | 6.5095 | 2.3544
| 470 |  | 2.1283 | 470 | 0.005685 | 6.5247 | 2.3543 | 480 | 6.5153 | 2.3542
| 485 |  | 2.1283 | 485 | 0.005814 | 6.5106 | 2.3542 | 495 | 6.5081 | 2.3552
| 500 | 2.127 | 2.127 | 500 | 0.0053736 | 6.5068 | 2.3557 | 510 | 6.3028 | 2.3533
| 515 |  | 2.1298 | 515 | 0.0052893 | 6.2008 | 2.3521 | 525 | 6.3113 | 2.3526
| 530 |  | 2.1293 | 530 | 0.0052959 | 6.3666 | 2.3529
UL-LMPS | 42.5 | 2.035 | 2.0256 | 1.9783 | 0.0071635 | -2.3328 | 0.71356 | 45 | -2.1342 | 0.73751
| 57.5 | 2.0348 | 2.016 | 1.993 | 0.0104 | -1.1413 | 0.85727 | 60 | -0.91752 | 0.83332
| 72.5 | 2.0304 | 2.0136 | 2.0177 | 0.0071635 | 0.20116 | 0.71356 | 75 | 0.45637 | 0.71443
| 87.5 | 2.0243 | 2.0062 | 2.041 | 0.008645 | 1.7324 | 0.71877 | 90 | 1.8665 | 0.7172
| 102.5 | 2.0139 | 2.0048 | 2.0557 | 0.0057124 | 2.5367 | 0.70939 | 105 | 2.656 | 0.70985
| 117.5 | 2.0138 | 1.9974 | 2.0623 | 0.0067082 | 3.2525 | 0.71216 | 120 | 3.2396 | 0.7593
| 132.5 | 2.0072 | 1.9982 | 2.0617 | 0.013563 | 3.1755 | 0.99505 | 135 | 3.3804 | 0.97215
| 147.5 | 2.0004 | 1.9859 | 2.0733 | 0.008507 | 4.4051 | 0.85765 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
| 50 |  | 2.0009 | 1.954 | 0.019304 | -2.3422 | 6.3 | 60 | -1.6618 | 6.3
SHMU | 65 |  | 2.0101 | 1.9835 | 0.022308 | -1.3215 | 6.3 | 75 | 1.0675 | 6.3
| 80 |  | 2.0159 | 2.0615 | 0.0087509 | 2.262 | 6.3 | 90 | 3.168 | 6.3
| 95 |  | 2.0131 | 2.086 | 0.0088258 | 3.6211 | 6.3 | 105 | 3.856 | 6.3
| 110 |  | 2.0096 | 2.0895 | 0.0075915 | 3.9735 | 6.3 | 120 | 4.3425 | 6.3
| 125 |  | 2.0067 | 2.0975 | 0.0063867 | 4.527 | 6.3 | 135 | 5.102 | 6.3
| 140 |  | 1.9959 | 2.1035 | 0.0074516 | 5.3896 | 6.3 | 150 | 5.838 | 6.3
| 155 |  | 1.9833 | 2.1035 | 0.0058714 | 6.0621 | 6.3 | 165 | 6.2281 | 6.3
| 170 |  | 1.9758 | 2.1005 | 0.0068633 | 6.3111 | 6.3 | 180 | 6.4156 | 6.3
| 185 |  | 1.9734 | 2.101 | 0.0071818 | 6.4679 | 6.3 | 195 | 6.6179 | 6.3
| 200 |  | 1.9715 | 2.1035 | 0.0058714 | 6.6928 | 6.3 | 210 | 6.0358 | 6.3
| 215 |  | 1.9833 | 2.0965 | 0.0048936 | 5.7073 | 6.3 | 225 | 5.7145 | 6.3
| 230 |  | 1.9916 | 2.1055 | 0.0051042 | 5.7182 | 6.3 | 240 | 5.3712 | 6.3
| 245 |  | 1.9943 | 2.098 | 0.0069585 | 5.1978 | 6.3 | 255 | 5.0424 | 6.3
| 260 |  | 1.9921 | 2.091 | 0.0078807 | 4.9647 | 6.3 | 270 | 5.3973 | 6.3
| 275 |  | 1.9964 | 2.1085 | 0.0074516 | 5.6137 | 6.3 | 285 | 5.4764 | 6.3
| 290 |  | 1.9951 | 2.103 | 0.007327 | 5.4078 | 6.3 | 300 | 5.1489 | 6.3
| 305 |  | 1.9882 | 2.088 | 0.011517 | 5.0195 | 6.3 | 315 | 5.1756 | 6.3
| 320 | 1.9833 | 1.9833 | 2.0875 | 0.008507 | 5.2537 | 6.3 | 330 | 5.3208 | 6.3
| 335 |  | 1.9752 | 2.081 | 0.0064072 | 5.3544 | 6.3 | 345 | 5.5537 | 6.3
| 350 |  | 1.9687 | 2.08 | 0.0045883 | 5.6533 | 6.3 | 360 | 5.5687 | 6.3
| 365 |  | 1.9725 | 2.0815 | 0.0058714 | 5.5264 | 6.3 | 375
390
405
420
435
450
465
480
495
510
525
| 40 |  | 2.143 | 2.033 | 0.03 | -5.133 | 1.4 | 45 | -4.7123 | 1.4
| 55 |  | 2.144 | 2.061 | 0.016 | -3.871 | 1.4 | 60 | -3.715 | 1.4
LEI | 70 |  | 2.145 | 2.072 | 0.008 | -3.403 | 1.4 | 75 | -2.6883 | 1.4
| 85 |  | 2.144 | 2.117 | 0.009 | -1.259 | 1.4 | 90 | -0.855 | 1.4
| 100 |  | 2.145 | 2.144 | 0.008 | -0.047 | 1.4 | 105 | 0.20167 | 1.4
| 115 |  | 2.145 | 2.16 | 0.006 | 0.699 | 1.4 | 120 | 0.90067 | 1.4
| 130 |  | 2.147 | 2.175 | 0.01 | 1.304 | 1.4 | 135 | 1.5847 | 1.4
| 145 |  | 2.144 | 2.19 | 0.009 | 2.146 | 1.4 | 150 | 2.317 | 1.4
| 160 |  | 2.144 | 2.201 | 0.007 | 2.659 | 1.4 | 165 | 2.8743 | 1.4
| 175 |  | 2.148 | 2.219 | 0.01 | 3.305 | 1.4 | 180 | 3.477 | 1.4
| 190 |  | 2.146 | 2.228 | 0.007 | 3.821 | 1.4 | 195 | 3.883 | 1.4
| 205 |  | 2.146 | 2.232 | 0.008 | 4.007 | 1.4 | 210 | 4.1327 | 1.4
| 220 |  | 2.144 | 2.238 | 0.009 | 4.384 | 1.4 | 225 | 4.3387 | 1.4
| 235 |  | 2.142 | 2.233 | 0.009 | 4.248 | 1.4 | 240 | 4.1517 | 1.4
| 250 |  | 2.147 | 2.232 | 0.011 | 3.959 | 1.4 | 255 | 4.054 | 1.4
| 265 |  | 2.144 | 2.235 | 0.009 | 4.244 | 1.4 | 270 | 4.0563 | 1.4
| 280 |  | 2.146 | 2.225 | 0.008 | 3.681 | 1.4 | 285 | 3.6817 | 1.4
| 295 |  | 2.145 | 2.224 | 0.006 | 3.683 | 1.4 | 300 | 3.701 | 1.4
| 310 |  | 2.141 | 2.221 | 0.008 | 3.737 | 1.4 | 315 | 3.7063 | 1.4
| 325 |  | 2.14 | 2.218 | 0.011 | 3.645 | 1.4 | 330 | 3.4727 | 1.4
| 340 |  | 2.142 | 2.209 | 0.008 | 3.128 | 1.4 | 345 | 2.9867 | 1.4
| 355 |  | 2.145 | 2.203 | 0.007 | 2.704 | 1.4 | 360
| 370 |  | 2.143 | 2.189 | 0.007 | 2.147 | 1.4 | 375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 5
| 5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.53072
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 4.979 | 4.723 | 0.040291 | -5.1409 | 0.84355 | 45 | -5.565 | 0.7629
| 52.5 |  | 4.9835 | 4.685 | 0.0070711 | -5.9891 | 0.68226 | 60 | -5.383 | 0.68236 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 4.982 | 4.744 | 0.011738 | -4.7769 | 0.68246 | 75 | -3.8587 | 0.6921 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 4.9877 | 4.841 | 0.011972 | -2.9406 | 0.70173 | 90 | -2.4732 | 0.69196 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 4.984 | 4.884 | 0.0084327 | -2.0058 | 0.68219 | 105 | -1.8876 | 0.68162 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 4.9923 | 4.904 | 0.0084327 | -1.7694 | 0.68104 | 120 | -1.3503 | 0.68152 |  | 1.3044 | 2.697 | 2.5732 | 0.66687 | 1.646 | 1.6593 | 1.7451 | 2.0782 | 0.64755 | 0.56272 | 3.4696 | 0.65207 | 3.1488 | 1.0108 | 2.2726 |  |  |  | 2.2487 | 1.2598 | 0.15085 | 0.6334 | 2.9221 | 0.46195
| 127.5 | 4.9852 | 4.9854 | 4.939 | 0.008756 | -0.93124 | 0.68199 | 135 | -0.8194 | 0.69167 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 4.9903 | 4.955 | 0.014337 | -0.70757 | 0.70136 | 150 | -0.15138 | 0.70217 | 45 |  | 0.54405 |  | 1.0265 | 1.8868 | [5 empty] | 1.769 | 2.6477 |  |  | 0.98869
| 157.5 |  | 4.9788 | 4.999 | 0.008756 | 0.40481 | 0.70297 | 165 | 0.54085 | 0.70259 | 60 | 0.26908 | 1.8486 | 0.54158 | 0.8961 | 0.29873 | 1.84 | 0.37085 | 1.2778 | 0.57195 | 0.91422 | 2.8339 | 1.995 | 1.2885 | 0.76511 | 0.52521
| 172.5 |  | 4.9843 | 5.018 | 0.012293 | 0.67689 | 0.70221 | 180 | 0.45195 | 0.70155 | 75 | 0.47948 | 2.3558 | 1.5343 | 0.4981 | 0.61275 | 2.2095 | 1.2819 | 0.08259 | 0.24568 | 0.28189 | 1.8171 | 1.4759 | 1.1192 | 0.92051 | 0.15489
| 187.5 |  | 4.9937 | 5.005 | 0.0097183 | 0.22701 | 0.70089 | 195 | 0.2454 | 0.691 | 90 | 0.16407 | 2.1192 | 1.7742 | 0.88818 | 0.24683 | 2.6637 | 2.0638 | 1.1767 | 0.4224 | 0.17041 | 1.1752 | 1.6592 | 1.0488 | 1.463 | 0.56897
| 202.5 |  | 4.9918 | 5.005 | 0.0084984 | 0.26379 | 0.68111 | 210 | -0.045238 | 0.67962 | 105 | 0.039691 | 2.1148 | 2.2755 | 1.1371 | 0.36942 | 2.4184 | 2.4388 | 1.2486 | 0.37291 | 0.71493 | 0.88867 | 1.51 | 1.3455 | 1.8238 | 0.66943
| 217.5 |  | 5.0138 | 4.996 | 0.0084327 | -0.35427 | 0.67813 | 225 |  |  | 120 | 0.67947 | 1.9897 | 2.3949 | 0.95599 | 0.59628 | 1.5453 | 2.0535 | 0.39486 | 0.048484 | 0.94467 | 0.95748 | 1.1509 | 1.6135 | 1.7418 | 0.27547 |  |  |  | 0.50406 | 0.077834 | 1.3297 | 1.8884 | 0.40263 | 0.10663
[8 empty] | 240 |  |  | 135 | 0.4546 | 2.0068 | 2.3782 | 0.63093 | 0.6965 | 1.8334 | 2.224 | 0.26885 | 0.36229 | 0.91206 | 1.2611 | 1.0667 | 1.808 | 1.6464 | 0.099166 |  |  |  | 0.76335 | 0.42227 | 1.088 | 1.6724 | 0.13755 | 0.032169
[8 empty] | 255 |  |  | 150 | 0.040063 | 1.8185 | 2.2171 |  | 0.71449 | 2.1022 | 2.3941 |  | 0.76151 | 0.89247 |  | 0.88384 |  | 1.4845 |  |  |  |  | 0.53846 | 0.57138 | 1.1892 | 1.7346 |  | 0.20339
[8 empty] | 270 |  |  | 165 | 0.14475 | 1.5809 | 1.9728 |  | 0.5406 | 2.0353 | 2.2778 |  | 0.73612 | 0.8249 |  | 0.85685 |  | 1.4036 |  |  |  |  | 0.43227 | 0.62904 | 1.0547 | 1.5674 |  | 0.1252
[8 empty] | 285 |  |  | 180 | 0.050958 | 1.6214 | 2.1977 |  | 0.80538 | 1.856 | 2.3632 |  | 0.85284 | 1.0339 |  | 0.60337 |  | 1.3873 |  |  |  |  | 1.2557 | 1.3737 | 0.20078 | 1.1043 |  | 0.37529
[8 empty] | 300 |  |  | 195 | 0.20605 | 1.7317 | 2.4821 |  | 0.85312 | 1.7945 | 2.5378 |  | 0.76129 | 1.2488 |  | 0.64539 |  | 1.6097 |  |  |  |  | 1.7456 | 1.7654 | 0.24712 | 0.95053 |  | 0.78682
[8 empty] | 315 |  |  | 210 | 0.39732 | 1.9602 | 2.6828 |  | 0.90758 | 1.8293 | 2.5913 |  | 0.64611 | 1.2771 |  | 0.78535 |  | 1.7453 |  |  |  |  | 1.6554 | 1.4714 | 0.066744 | 1.2177 |  | 0.64681
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.934 | 2.7931 |  | 0.79468 | 1.4032 |  | 0.72086 |  | 1.8052 | [5 empty] | 2.1057 | 0.46167 | 0.92789 |  | 1.036
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.3468 | 3.0729 |  | 1.0584 | 1.3915 |  | 0.78246 |  | 1.8436 | [5 empty] | 2.203 | 0.18816 | 1.1391 |  | 0.86691
[8 empty] | 360 |  |  | 255 | [5 empty] | 2.4029 | 3.2248 |  | 1.3976 | 1.5016 |  | 0.49546 |  | 1.7138 | [5 empty] | 2.4593 | 0.39151 | 1.0845 |  | 0.77757
[8 empty] | 375 |  |  | 270 | [5 empty] | 3.1789 | 3.7755 |  | 1.7735 | 1.5023 |  | 0.73243 |  | 1.9017 | [5 empty] | 3.0248 | 0.30625 | 1.1463 |  | 0.91881
[8 empty] | 390 |  |  | 285 | [9 empty] | 1.5681 |  | 0.61311 |  | 1.8653 | [6 empty] | 0.57882 | 0.94917 |  | 1.0285
[8 empty] | 405 |  |  | 300 | [9 empty] | 1.4852 |  | 0.78031 |  | 1.9245 | [6 empty] | 0.19059 | 1.1765 |  | 0.84259
[8 empty] | 420 |  |  | 315 | [9 empty] | 1.7831 |  | 0.24671 |  | 1.7652 | [6 empty] | 0.35027 | 1.3478 |  | 0.51572
[8 empty] | 435 |  |  | 330 | [9 empty] | 1.6831 |  | 0.6527 |  | 1.9979 | [6 empty] | 0.38542 | 1.2598 |  | 0.91593
[8 empty] | 450 |  |  | 345 | [9 empty] | 1.7005 |  | 1.4531 |  | 2.6458 | [6 empty] | 0.33136 | 1.206 |  | 1.5065
[8 empty] | 465 |  |  | 360 | [9 empty] | 1.8623 | [10 empty] | 0.80553 | 1.096
[8 empty] | 480 |  |  | 375 | [9 empty] | 2.2755 | [10 empty] | 1.1483 | 1.2175
[8 empty] | 495 |  |  | 390 | [9 empty] | 2.6545 | [10 empty] | 1.7805 | 1.0636
[8 empty] | 510 |  |  | 405 | [9 empty] | 3.0223 | [10 empty] | 3.1442 | 0.32639
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.3928
435
450
45
PTB | 55 | 4.9982 | 5.0241 | 4.7311 | 0.02699 | -5.833 | 0.46 | 60 | -5.1199 | 0.45667
| 70 | 4.9982 | 5.0204 | 4.835 | 0.016793 | -3.6938 | 0.45 | 75 | -3.3874 | 0.45333
| 85 | 4.9982 | 5.0139 | 4.8748 | 0.025351 | -2.7746 | 0.46 | 90 | -2.6347 | 0.45667
| 100 | 4.9982 | 5.01 | 4.892 | 0.013229 | -2.3548 | 0.45 | 105 | -1.8489 | 0.45
| 115 | 4.9982 | 5.0073 | 4.9654 | 0.01067 | -0.8371 | 0.45 | 120 | -0.68851 | 0.45
| 130 | 4.9982 | 5.0015 | 4.9819 | 0.012967 | -0.39134 | 0.45 | 135 | -0.37339 | 0.45
| 145 | 4.9982 | 4.9989 | 4.982 | 0.017559 | -0.33749 | 0.45 | 150 | -0.11178 | 0.45
| 160 | 4.9982 | 4.9982 | 5.0152 | 0.013118 | 0.33966 | 0.45 | 165 | 0.39772 | 0.45
| 175 | 4.9982 | 4.9995 | 5.0252 | 0.015285 | 0.51384 | 0.45 | 180 | 0.5023 | 0.45
| 190 | 4.9982 | 5.002 | 5.0259 | 0.015002 | 0.47922 | 0.45 | 195 | 0.44746 | 0.45
| 205 | 4.9982 | 5.0038 | 5.023 | 0.016365 | 0.38393 | 0.45 | 210 | 0.34121 | 0.45
| 220 | 4.9982 | 5.0057 | 5.0185 | 0.013503 | 0.25578 | 0.45 | 225 | 0.10429 | 0.45
| 235 | 4.9982 | 5.0148 | 5.0048 | 0.019289 | -0.19869 | 0.45 | 240 | -0.27837 | 0.45
| 250 | 4.9982 | 5.0175 | 4.9956 | 0.018879 | -0.43773 | 0.45 | 255 | -0.62769 | 0.45333
| 265 | 4.9982 | 5.0236 | 4.973 | 0.020721 | -1.0076 | 0.46 | 270 | -1.4438 | 0.45667
| 280 | 4.9982 | 5.0321 | 4.9156 | 0.019871 | -2.316 | 0.45 | 285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 5.0769 | 4.8235 | 0.018144 | -4.9909 | 0.5 | 45 | -4.9909 | 0.5
| 60 |  | 5.0745 | 4.895 | 0.022595 | -3.5369 | 0.5 | 60 | -3.5369 | 0.5
| 75 |  | 5.0711 | 4.9955 | 0.013945 | -1.4904 | 0.5 | 75 | -1.4904 | 0.5
| 90 |  | 5.0679 | 5.0505 | 0.014681 | -0.34294 | 0.5 | 90 | -0.34294 | 0.5
| 105 |  | 5.0702 | 5.0815 | 0.015652 | 0.22327 | 0.5 | 105 | 0.22327 | 0.5
| 120 |  | 5.0698 | 5.102 | 0.018525 | 0.63553 | 0.5 | 120 | 0.63553 | 0.5
| 135 |  | 5.0663 | 5.127 | 0.013018 | 1.1975 | 0.5 | 135 | 1.1975 | 0.5
| 150 |  | 5.0649 | 5.1505 | 0.017313 | 1.6895 | 0.5 | 150 | 1.6895 | 0.5
| 165 |  | 5.0601 | 5.1685 | 0.014609 | 2.1416 | 0.5 | 165 | 2.1416 | 0.5
| 180 |  | 5.064 | 5.17 | 0.013377 | 2.0926 | 0.5 | 180 | 2.0926 | 0.5
| 195 |  | 5.0664 | 5.167 | 0.01689 | 1.985 | 0.5 | 195 | 1.985 | 0.5
| 210 |  | 5.0648 | 5.1615 | 0.016631 | 1.9087 | 0.5 | 210 | 1.9087 | 0.5
| 225 | 5.0653 | 5.0653 | 5.1545 | 0.021145 | 1.7614 | 0.5 | 225 | 1.7614 | 0.5
| 240 |  | 5.0662 | 5.154 | 0.01875 | 1.7325 | 0.5 | 240 | 1.7325 | 0.5
| 255 |  | 5.0697 | 5.1425 | 0.012513 | 1.4354 | 0.5 | 255 | 1.4354 | 0.5
| 270 |  | 5.0705 | 5.136 | 0.020105 | 1.2912 | 0.5 | 270 | 1.2912 | 0.5
| 285 |  | 5.072 | 5.125 | 0.013955 | 1.0444 | 0.5 | 285 | 1.0444 | 0.5
| 300 |  | 5.0739 | 5.1265 | 0.016311 | 1.0361 | 0.5 | 300 | 1.0361 | 0.5
| 315 |  | 5.0735 | 5.104 | 0.015694 | 0.60057 | 0.5 | 315 | 0.60057 | 0.5
| 330 |  | 5.0708 | 5.1005 | 0.018202 | 0.5861 | 0.5 | 330 | 0.5861 | 0.5
| 345 |  | 5.0748 | 5.0985 | 0.015313 | 0.46741 | 0.5 | 345 | 0.46741 | 0.5
| 360 |  | 5.0694 | 5.0825 | 0.015853 | 0.25881 | 0.5 | 360 | 0.25881 | 0.5
| 375 |  | 5.074 | 5.057 | 0.015594 | -0.33563 | 0.5 | 375 | -0.33563 | 0.5
| 390 |  | 5.0789 | 5.0395 | 0.018489 | -0.77634 | 0.5 | 390 | -0.77634 | 0.5
| 405 |  | 5.0602 | 4.9985 | 0.02207 | -1.2199 | 0.5 | 405 | -1.2199 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 5.0035 | 4.965 | 0.05 | -0.76895 | 0.7642 | 120 | -0.76895 | 0.7642
| 135 |  | 4.9971 | 5 | 0.045 | 0.058845 | 0.75074 | 135 | 0.058845 | 0.75074
| 150 |  | 5.0018 | 5.025 | 0.045 | 0.4638 | 0.72841 | 150 | 0.4638 | 0.72841
| 165 |  | 4.998 | 5.05 | 0.04 | 1.0402 | 0.74762 | 165 | 1.0402 | 0.74762
| 180 |  | 5.0007 | 5.095 | 0.04 | 1.8857 | 0.72801 | 180 | 1.8857 | 0.72801
| 195 |  | 5.0077 | 5.12 | 0.045 | 2.2419 | 0.74097 | 195 | 2.2419 | 0.74097
| 210 |  | 5.0079 | 5.1 | 0.035 | 1.8392 | 0.74311 | 210 | 1.8392 | 0.74311
| 225 |  | 5.0127 | 5.125 | 0.045 | 2.2404 | 0.73818 | 225 | 2.2404 | 0.73818
| 240 |  | 5.0086 | 5.105 | 0.035 | 1.9251 | 0.71851 | 240 | 1.9251 | 0.71851
| 255 |  | 5.0031 | 5.095 | 0.035 | 1.8362 | 0.71871 | 255 | 1.8362 | 0.71871
| 270 | 5.007 | 5.0046 | 5.085 | 0.04 | 1.6062 | 0.72563 | 270 | 1.6062 | 0.72563
| 285 |  | 5.0065 | 5.09 | 0.06 | 1.6681 | 0.79358 | 285 | 1.6681 | 0.79358
| 300 |  | 5.0029 | 5.065 | 0.04 | 1.242 | 0.79761 | 300 | 1.242 | 0.79761
| 315 |  | 5.0016 | 5.05 | 0.055 | 0.96861 | 0.75658 | 315 | 0.96861 | 0.75658
| 330 |  | 4.9961 | 5.045 | 0.035 | 0.97946 | 0.71409 | 330 | 0.97946 | 0.71409
| 345 |  | 5.0077 | 5.05 | 0.045 | 0.84409 | 0.87211 | 345 | 0.84409 | 0.87211
| 360 |  | 5.0007 | 5.055 | 0.04 | 1.0858 | 0.7228 | 360 | 1.0858 | 0.7228
| 375 |  | 4.998 | 5.04 | 0.02 | 0.84008 | 0.71877 | 375 | 0.84008 | 0.71877
| 390 |  | 5.0018 | 5.055 | 0.045 | 1.0636 | 0.73223 | 390 | 1.0636 | 0.73223
| 405 |  | 4.9971 | 5.1 | 0.035 | 2.06 | 0.74599 | 405 | 2.06 | 0.74599
| 420 |  | 5.0035 | 5.1 | 0.035 | 1.9292 | 0.7215 | 420 | 1.9292 | 0.7215
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 |  | 5.0027 | 4.6534 | 0.069522 | -6.9816 | 1.0218 | 60 | -4.6676 | 0.99891
| 65 |  | 5.0667 | 4.8888 | 0.043886 | -3.5105 | 0.98745 | 75 | -1.8366 | 0.98812
| 80 |  | 5.0587 | 5.0081 | 0.027273 | -0.9996 | 0.98845 | 90 | -0.13353 | 0.9892
| 95 |  | 5.0527 | 5.0678 | 0.024653 | 0.29951 | 0.98958 | 105 | 1.1023 | 0.99004
| 110 |  | 5.0497 | 5.1256 | 0.02036 | 1.5037 | 0.99028 | 120 | 1.7976 | 0.99072
| 125 |  | 5.0467 | 5.1448 | 0.017072 | 1.9445 | 0.99094 | 135 | 2.3197 | 0.99102
| 140 |  | 5.0467 | 5.1732 | 0.0145 | 2.5073 | 0.99106 | 150 | 2.7877 | 0.99127
| 155 |  | 5.0457 | 5.1934 | 0.015203 | 2.9279 | 0.99137 | 165 | 3.1569 | 0.99158
| 170 |  | 5.0447 | 5.2097 | 0.017409 | 3.2714 | 0.99168 | 180 | 3.3653 | 0.99171
| 185 |  | 5.0447 | 5.2168 | 0.016281 | 3.4122 | 0.99172 | 195 | 3.5223 | 0.99189
| 200 |  | 5.0437 | 5.2241 | 0.015578 | 3.5774 | 0.99198 | 210 | 3.4809 | 0.99194
| 215 |  | 5.0437 | 5.2168 | 0.014065 | 3.4327 | 0.99193 | 225 | 3.4891 | 0.99221
| 230 |  | 5.0417 | 5.219 | 0.015482 | 3.5174 | 0.99235 | 240 | 3.446 | 0.99232
| 245 |  | 5.0417 | 5.2136 | 0.012285 | 3.4102 | 0.99231 | 255 | 3.2842 | 0.99214
| 260 |  | 5.0427 | 5.2051 | 0.011271 | 3.2212 | 0.99205 | 270 | 3.1409 | 0.99216
| 275 |  | 5.0417 | 5.198 | 0.014931 | 3.1008 | 0.99221 | 285 | 2.975 | 0.99211
| 290 |  | 5.0422 | 5.189 | 0.015896 | 2.9121 | 0.99206 | 300 | 2.8647 | 0.99211
| 305 |  | 5.0417 | 5.1849 | 0.014607 | 2.841 | 0.99214 | 315 | 2.796 | 0.99212
| 320 |  | 5.0417 | 5.1815 | 0.015356 | 2.7736 | 0.99212 | 330 | 2.6582 | 0.99196
| 335 |  | 5.0427 | 5.1738 | 0.015638 | 2.6005 | 0.99187 | 345 | 2.5606 | 0.9918
| 350 |  | 5.0432 | 5.1713 | 0.011038 | 2.5407 | 0.99176 | 360 | 2.5511 | 0.9917
| 365 |  | 5.0437 | 5.1726 | 0.012676 | 2.5563 | 0.99167 | 375 | 2.4651 | 0.99164
| 380 |  | 5.0437 | 5.1657 | 0.01234 | 2.4195 | 0.99163 | 390 | 2.4911 | 0.99172
| 395 |  | 5.0432 | 5.1706 | 0.015744 | 2.5269 | 0.99176 | 405 | 2.5006 | 0.99182
| 410 |  | 5.0427 | 5.1681 | 0.015219 | 2.4874 | 0.99185 | 420 | 2.4541 | 0.99171
| 425 |  | 5.0437 | 5.1666 | 0.010876 | 2.4374 | 0.99164 | 435 | 2.4551 | 0.99184
| 440 |  | 5.0422 | 5.1664 | 0.0152 | 2.4639 | 0.99194 | 450 | 2.4429 | 0.992
| 455 |  | 5.0417 | 5.1643 | 0.012869 | 2.4324 | 0.99203 | 465 | 2.4479 | 0.9919
| 470 |  | 5.0427 | 5.1665 | 0.01196 | 2.4557 | 0.99184 | 480 | 2.464 | 0.99197
| 485 |  | 5.0417 | 5.1661 | 0.015182 | 2.4681 | 0.99204 | 495 | 2.4588 | 0.99204
| 500 | 5.042 | 5.0417 | 5.1654 | 0.011657 | 2.4542 | 0.99203 | 510 | 2.4253 | 0.99209
| 515 |  | 5.0412 | 5.1627 | 0.012657 | 2.4108 | 0.99212 | 525 | 2.391 | 0.99212
| 530 |  | 5.0412 | 5.1612 | 0.014919 | 2.3811 | 0.99212
UL-LMPS | 42.5 | 5.0329 | 5.0447 | 4.7065 | 0.016631 | -6.7038 | 0.66415 | 45 | -6.7395 | 0.66742
| 57.5 | 5.0181 | 5.009 | 4.6625 | 0.027886 | -6.9179 | 0.68375 | 60 | -6.371 | 0.68444
| 72.5 | 4.9873 | 4.978 | 4.797 | 0.029753 | -3.6366 | 0.68789 | 75 | -3.3067 | 0.68371
| 87.5 | 4.9735 | 4.9687 | 4.8863 | 0.015594 | -1.6575 | 0.66282 | 90 | -1.5006 | 0.66229
| 102.5 | 4.9576 | 4.9595 | 4.924 | 0.012763 | -0.71591 | 0.65962 | 105 | -0.65101 | 0.66045
| 117.5 | 4.9422 | 4.9502 | 4.934 | 0.016944 | -0.32652 | 0.66457 | 120 | -0.30873 | 0.66401
| 132.5 | 4.916 | 4.9252 | 4.9143 | 0.014244 | -0.21979 | 0.66122 | 135 | -0.10251 | 0.72867
| 147.5 | 4.8912 | 4.8986 | 4.9223 | 0.022821 | 0.48385 | 1.0659 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 5.192 | 4.781 | 0.024 | -7.916 | 0.83 | 45 | -7.916 | 0.83
| 60 |  | 5.191 | 4.893 | 0.024 | -5.741 | 0.83 | 60 | -5.741 | 0.83
LEI | 75 |  | 5.191 | 5.029 | 0.024 | -3.121 | 0.83 | 75 | -3.121 | 0.83
| 90 |  | 5.192 | 5.079 | 0.023 | -2.176 | 0.83 | 90 | -2.176 | 0.83
| 105 |  | 5.191 | 5.116 | 0.019 | -1.445 | 0.83 | 105 | -1.445 | 0.83
| 120 |  | 5.191 | 5.158 | 0.023 | -0.636 | 0.83 | 120 | -0.636 | 0.83
| 135 |  | 5.191 | 5.192 | 0.02 | 0.019 | 0.83 | 135 | 0.019 | 0.83
| 150 |  | 5.191 | 5.228 | 0.021 | 0.713 | 0.83 | 150 | 0.713 | 0.83
| 165 |  | 5.187 | 5.249 | 0.019 | 1.195 | 0.83 | 165 | 1.195 | 0.83
| 180 |  | 5.189 | 5.263 | 0.027 | 1.426 | 0.83 | 180 | 1.426 | 0.83
| 195 |  | 5.19 | 5.256 | 0.018 | 1.272 | 0.83 | 195 | 1.272 | 0.83
| 210 |  | 5.187 | 5.241 | 0.019 | 1.041 | 0.83 | 210 | 1.041 | 0.83
| 225 |  | 5.183 | 5.233 | 0.017 | 0.965 | 0.83 | 225 | 0.965 | 0.83
| 240 |  | 5.182 | 5.227 | 0.021 | 0.868 | 0.83 | 240 | 0.868 | 0.83
| 255 |  | 5.182 | 5.228 | 0.015 | 0.888 | 0.83 | 255 | 0.888 | 0.83
| 270 |  | 5.182 | 5.207 | 0.013 | 0.482 | 0.83 | 270 | 0.482 | 0.83
| 285 |  | 5.184 | 5.203 | 0.021 | 0.367 | 0.83 | 285 | 0.367 | 0.83
| 300 |  | 5.184 | 5.193 | 0.024 | 0.174 | 0.83 | 300 | 0.174 | 0.83
| 315 |  | 5.185 | 5.202 | 0.027 | 0.328 | 0.83 | 315 | 0.328 | 0.83
| 330 |  | 5.182 | 5.175 | 0.027 | -0.135 | 0.83 | 330 | -0.135 | 0.83
| 345 |  | 5.184 | 5.125 | 0.023 | -1.138 | 0.83 | 345 | -1.138 | 0.83
| 360 |  | 5.184 | 4.954 | 0.038 | -4.437 | 0.83 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 8
| 8 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.37559
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 7.996 | 7.471 | 0.011005 | -6.5662 | 0.65032 | 45 | -6.7762 | 0.65642
| 52.5 |  | 7.9999 | 7.441 | 0.01912 | -6.9862 | 0.66251 | 60 | -5.8793 | 0.66264 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 7.9966 | 7.615 | 0.021213 | -4.7724 | 0.66278 | 75 | -3.668 | 0.66255 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 8.0021 | 7.797 | 0.011595 | -2.5636 | 0.66232 | 90 | -2.1696 | 0.66245 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 7.999 | 7.857 | 0.013375 | -1.7755 | 0.66258 | 105 | -1.5469 | 0.66847 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 8.0076 | 7.902 | 0.018135 | -1.3184 | 0.67436 | 120 | -0.82734 | 0.67485 |  | 0.055258 | 0.90582 | 0.14666 | 3.0011 | 0.098358 | 1.1713 | 0.23153 | 3.4593 | 0.058743 | 0.78243 | 4.2727 | 1.042 | 3.2613 | 0.25238 | 2.9717 |  |  | 0.034715 | 0.27181 | 0.3739 | 0.6109 | 0.13401 | 3.2995 | 0.37727 | [6 empty] | 0.069214 | 0.089229 | 0.2473 | 0.014789 | 1.2149 | 0.1066
| 127.5 | 7.9953 | 7.9959 | 7.969 | 0.01912 | -0.33631 | 0.67535 | 135 | -0.16847 | 0.67492 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 8.006 | 8.006 | 0.016465 | -0.00062437 | 0.67449 | 150 | 0.5889 | 0.67522 | 45 |  | 0.88007 |  | 1.4898 | 1.8812 | [5 empty] | 0.76538 | 1.1912 |  |  | 0.37579
| 157.5 |  | 7.9889 | 8.083 | 0.020575 | 1.1784 | 0.67594 | 165 | 1.0637 | 0.68172 | 60 | 0.11141 | 0.77033 | 0.81782 | 1.4151 | 0.072054 | 0.80309 | 1.071 | 1.527 | 0.029573 | 1.7013 | 0.80355 | 0.70428 | 2.254 | 0.90408 | 1.3643 | [15 empty] | 0.78111 | 0.83312 | 1.0591 | 0.47876 | 1.3214 | 0.80973
| 172.5 |  | 8.0001 | 8.076 | 0.02319 | 0.94891 | 0.68749 | 180 | 0.89265 | 0.68736 | 75 | 0.29839 | 0.094636 | 0.85222 | 0.70111 | 0.27741 | 0.45896 | 1.3081 | 0.51508 | 0.019287 | 0.85241 | 0.87216 | 0.40477 | 1.5745 | 1.1506 | 0.43341 | [15 empty] | 0.68242 | 0.79481 | 0.65887 | 0.36906 | 0.94806 | 0.78775
| 187.5 |  | 8.0031 | 8.07 | 0.022608 | 0.83639 | 0.68724 | 195 | 0.57113 | 0.68114 | 90 | 1.1114 | 0.29695 | 0.86883 | 1.0741 | 1.1465 | 0.92401 | 2.175 | 0.1323 | 0.20389 | 1.2769 | 0.89289 | 0.97219 | 1.9772 | 2.0635 | 0.059705 | [15 empty] | 0.76475 | 1.1495 | 0.87819 | 0.44642 | 1.1715 | 1.1958
| 202.5 |  | 7.9995 | 8.024 | 0.018379 | 0.30587 | 0.67504 | 210 | 0.070678 | 0.66827 | 105 | 0.97907 | 0.066462 | 1.113 | 0.84781 | 0.71195 | 1.0555 | 2.334 | 0.0010018 | 0.16992 | 1.3159 | 0.87689 | 0.72667 | 1.9994 | 1.8713 | 0.14549 | [15 empty] | 0.87318 | 1.2181 | 0.90873 | 0.46362 | 1.1958 | 1.1432
| 217.5 |  | 8.0122 | 7.999 | 0.015951 | -0.16451 | 0.66149 | 225 |  |  | 120 | 0.95616 | 0.2012 | 1.1304 | 0.61731 | 0.52307 | 1.3616 | 2.3416 | 0.24809 | 0.37052 | 1.0612 | 0.89422 | 0.79323 | 1.785 | 1.701 | 0.10215 |  |  | 1.1423 | 0.98624 | 0.16689 | 1.3168 | 2.1926 | 0.3616 | 0.46996 | [6 empty] | 0.76899 | 1.1064 | 0.70986 | 0.351 | 1.0056 | 0.96904
[8 empty] | 240 |  |  | 135 | 1.1377 | 0.098022 | 1.1754 | 0.86157 | 0.33849 | 1.4487 | 2.5885 | 0.09338 | 0.77425 | 1.2198 | 1.0474 | 0.48024 | 2.0262 | 1.5554 | 0.55252 |  |  | 0.71899 | 0.45979 | 0.63607 | 0.61709 | 1.6859 | 0.43813 | 0.12302 | [6 empty] | 0.5438 | 0.93734 | 0.51749 | 0.10778 | 0.88204 | 0.67341
[8 empty] | 255 |  |  | 150 | 1.4909 | 0.0011776 | 0.9866 |  | 0.45161 | 1.7414 | 2.7388 |  | 2.0714 | 1.1113 |  | 0.50758 |  | 0.54185 |  |  |  | 0.5766 | 0.96364 | 0.39752 | 1.0789 | 2.0098 |  | 1.4483 | [6 empty] | 0.21149 | 0.71633 | 0.21504 | 0.15586 |  | 0.041576
[8 empty] | 270 |  |  | 165 | 1.6223 | 0.12875 | 0.79416 |  | 0.10784 | 1.7465 | 2.6576 |  | 1.5552 | 1.0329 |  | 0.011658 |  | 0.92972 |  |  |  | 0.2113 | 0.69484 | 0.84804 | 0.64912 | 1.5355 |  | 0.60412 | [6 empty] | 0.053693 | 0.49458 | 0.0091596 | 0.35194 |  | 0.012983
[8 empty] | 285 |  |  | 180 | 1.6255 | 0.39692 | 0.87219 |  | 0.39214 | 2.4075 | 2.7667 |  | 1.2292 | 0.57038 |  | 0.85549 |  | 1.3116 |  |  |  | 0.23017 | 0.12545 | 1.8074 | 0.26323 | 0.75537 |  | 0.52778 | [6 empty] | 0.18137 | 0.36796 | 0.32476 | 0.51097 |  | 0.032823
[8 empty] | 300 |  |  | 195 | 1.7844 | 0.60201 | 1.0993 |  | 0.67819 | 2.8329 | 3.1902 |  | 1.0488 | 0.60932 |  | 1.3855 |  | 1.8378 |  |  |  | 0.29775 | 0.2221 | 2.068 | 0.36015 | 0.87928 |  | 0.91297 | [6 empty] | 0.212 | 0.38759 | 0.42728 | 0.62547 |  | 0.044404
[8 empty] | 315 |  |  | 210 | 1.4824 | 0.69081 | 1.4791 |  | 0.50638 | 2.5594 | 3.289 |  | 0.92657 | 0.93779 |  | 1.2717 |  | 2.0314 |  |  |  | 0.33136 | 0.74655 | 2.3906 | 0.12753 | 0.73453 |  | 1.276 | [6 empty] | 0.050187 | 0.445 | 0.2917 | 0.60023 |  | 0.14009
[8 empty] | 330 |  |  | 225 | [5 empty] | 2.9039 | 3.5038 |  | 0.99224 | 0.84735 |  | 1.4991 |  | 2.1596 |  |  |  | 0.33964 |  | 2.8188 | 0.23376 | 0.5571 |  | 1.5855 | [7 empty] | 0.57222 | 0.26399 | 0.54277 |  | 0.2441
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.8684 | 3.766 |  | 1.5642 | 1.1271 |  | 0.92426 |  | 1.882 |  |  |  | 0.24734 |  | 2.7164 | 0.26026 | 0.7555 |  | 1.0599 | [7 empty] | 0.66708 | 0.15948 | 0.53211 |  | 0.15374
[8 empty] | 360 |  |  | 255 | [5 empty] | 3.4515 | 3.7629 |  | 1.6634 | 0.6248 |  | 1.3206 |  | 1.7916 |  |  |  | 0.23201 |  | 3.0899 | 0.09884 | 0.46796 |  | 1.2774 | [7 empty] | 0.79365 | 0.20087 | 0.40636 |  | 0.24634
[8 empty] | 375 |  |  | 270 | [5 empty] | 3.7643 | 3.7968 |  | 1.8511 | 0.38879 |  | 1.4095 |  | 1.659 |  |  |  | 0.18938 |  | 3.1848 | 0.072081 | 0.4153 |  | 1.2011 | [7 empty] | 0.87017 | 0.217 | 0.34399 |  | 0.26037
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.14936 |  | 0.96364 |  | 1.0283 |  |  |  | 0.14993 |  |  | 0.094164 | 0.045759 |  | 0.94887 | [8 empty] | 0.11921 | 0.16765 |  | 0.20682
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.10112 |  | 0.77508 |  | 0.80978 |  |  |  | 0.083749 |  |  | 0.20911 | 0.1095 |  | 0.85642 | [8 empty] | 0.0085378 | 0.042023 |  | 0.2527
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.20233 |  | 0.91574 |  | 1.0321 |  |  |  | 0.0018905 |  |  | 0.12412 | 0.064546 |  | 0.93099 | [8 empty] | 0.041563 | 0.025994 |  | 0.34962
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.15217 |  | 1.8524 |  | 1.8538 |  |  |  | 0.11407 |  |  | 0.28759 | 0.13338 |  | 1.9046 | [8 empty] | 0.014218 | 0.064595 |  | 0.61008
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.48036 |  | 5.4355 |  | 5.4704 |  |  |  | 0.16578 |  |  | 0.76179 | 0.28646 |  | 5.5019 | [8 empty] | 0.10159 | 0.058828 |  | 1.9318
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.60895 | [7 empty] | 0.18986 |  |  | 1.0707 | 0.44953 | [10 empty] | 0.17996 | 0.023883
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.722 | [10 empty] | 1.0043 | 0.28783
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.86618 | [10 empty] | 1.3071 | 0.44788
[8 empty] | 510 |  |  | 405 | [9 empty] | 1.362 | [10 empty] | 2.8856 | 1.4453
[8 empty] | 525 |  |  | 420 | [21 empty] | 1.0011
435
450
45
PTB | 55 | 8.0175 | 8.0573 | 7.5046 | 0.029289 | -6.8597 | 0.42 | 60 | -5.9762 | 0.42
| 70 | 8.0175 | 8.05 | 7.7111 | 0.022072 | -4.2093 | 0.42 | 75 | -3.9276 | 0.42
| 85 | 8.0175 | 8.0427 | 7.7721 | 0.02025 | -3.364 | 0.42 | 90 | -3.1361 | 0.42
| 100 | 8.0175 | 8.0369 | 7.8214 | 0.018199 | -2.6805 | 0.42 | 105 | -2.4029 | 0.42
| 115 | 8.0175 | 8.0317 | 7.8833 | 0.020191 | -1.8477 | 0.42 | 120 | -1.6679 | 0.42
| 130 | 8.0175 | 8.0257 | 7.9207 | 0.022432 | -1.3085 | 0.42 | 135 | -1.1669 | 0.41667
| 145 | 8.0175 | 8.0191 | 7.9482 | 0.01492 | -0.88382 | 0.41 | 150 | -0.71751 | 0.41333
| 160 | 8.0175 | 8.0175 | 7.9867 | 0.025115 | -0.38491 | 0.42 | 165 | -0.37109 | 0.42
| 175 | 8.0175 | 8.0187 | 7.9912 | 0.021038 | -0.34344 | 0.42 | 180 | -0.55209 | 0.42
| 190 | 8.0175 | 8.0228 | 7.945 | 0.033734 | -0.96939 | 0.42 | 195 | -1.0063 | 0.42
| 205 | 8.0175 | 8.0271 | 7.9404 | 0.024573 | -1.08 | 0.42 | 210 | -1.2251 | 0.42
| 220 | 8.0175 | 8.0302 | 7.9085 | 0.019553 | -1.5152 | 0.42 | 225 | -1.6901 | 0.42
| 235 | 8.0175 | 8.0414 | 7.8774 | 0.039085 | -2.0398 | 0.42 | 240 | -2.1658 | 0.42
| 250 | 8.0175 | 8.046 | 7.8515 | 0.028917 | -2.4178 | 0.42 | 255 | -2.4994 | 0.42
| 265 | 8.0175 | 8.0586 | 7.8441 | 0.038656 | -2.6626 | 0.42 | 270 | -2.8364 | 0.42333
| 280 | 8.0175 | 8.0742 | 7.8171 | 0.048063 | -3.1841 | 0.43 | 285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 8.0551 | 7.445 | 0.023731 | -7.5741 | 0.5 | 45 | -7.5741 | 0.5
| 60 |  | 8.0487 | 7.519 | 0.027701 | -6.5812 | 0.5 | 60 | -6.5812 | 0.5
| 75 |  | 8.0462 | 7.758 | 0.031389 | -3.5818 | 0.5 | 75 | -3.5818 | 0.5
| 90 |  | 8.053 | 7.8565 | 0.022542 | -2.4401 | 0.5 | 90 | -2.4401 | 0.5
| 105 |  | 8.0484 | 7.919 | 0.014832 | -1.6078 | 0.5 | 105 | -1.6078 | 0.5
| 120 |  | 8.0502 | 7.9985 | 0.015985 | -0.64222 | 0.5 | 120 | -0.64222 | 0.5
| 135 |  | 8.0483 | 8.042 | 0.016733 | -0.078277 | 0.5 | 135 | -0.078277 | 0.5
| 150 |  | 8.0467 | 8.094 | 0.021861 | 0.58782 | 0.5 | 150 | 0.58782 | 0.5
| 165 |  | 8.046 | 8.122 | 0.020417 | 0.94457 | 0.5 | 165 | 0.94457 | 0.5
| 180 |  | 8.042 | 8.1435 | 0.028149 | 1.2615 | 0.5 | 180 | 1.2615 | 0.5
| 195 |  | 8.0378 | 8.1285 | 0.024767 | 1.1278 | 0.5 | 195 | 1.1278 | 0.5
| 210 |  | 8.0449 | 8.1015 | 0.022308 | 0.70293 | 0.5 | 210 | 0.70293 | 0.5
| 225 | 8.0415 | 8.0415 | 8.0815 | 0.028149 | 0.49742 | 0.5 | 225 | 0.49742 | 0.5
| 240 |  | 8.0424 | 8.042 | 0.019628 | -0.0049736 | 0.5 | 240 | -0.0049736 | 0.5
| 255 |  | 8.0494 | 8.0575 | 0.032261 | 0.10063 | 0.5 | 255 | 0.10063 | 0.5
| 270 |  | 8.0445 | 8.045 | 0.03395 | 0.0062154 | 0.5 | 270 | 0.0062154 | 0.5
| 285 |  | 8.0466 | 8.043 | 0.027928 | -0.044739 | 0.5 | 285 | -0.044739 | 0.5
| 300 |  | 8.0472 | 8.033 | 0.036864 | -0.17646 | 0.5 | 300 | -0.17646 | 0.5
| 315 |  | 8.0481 | 8.026 | 0.039789 | -0.2746 | 0.5 | 315 | -0.2746 | 0.5
| 330 |  | 8.0492 | 8.0255 | 0.036487 | -0.29444 | 0.5 | 330 | -0.29444 | 0.5
| 345 |  | 8.0511 | 8.008 | 0.032541 | -0.53533 | 0.5 | 345 | -0.53533 | 0.5
| 360 |  | 8.0524 | 7.9955 | 0.031867 | -0.70662 | 0.5 | 360 | -0.70662 | 0.5
| 375 |  | 8.0598 | 7.9945 | 0.041861 | -0.81019 | 0.5 | 375 | -0.81019 | 0.5
| 390 |  | 8.0574 | 7.98 | 0.035836 | -0.96061 | 0.5 | 390 | -0.96061 | 0.5
| 405 |  | 8.0447 | 7.9305 | 0.033635 | -1.4196 | 0.5 | 405 | -1.4196 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 7.9946 | 7.85 | 0.045 | -1.8087 | 0.62735 | 120 | -1.8087 | 0.62735
| 135 |  | 7.9902 | 7.94 | 0.045 | -0.62813 | 0.63472 | 135 | -0.62813 | 0.63472
| 150 |  | 7.9953 | 7.965 | 0.055 | -0.37921 | 0.64213 | 150 | -0.37921 | 0.64213
| 165 |  | 7.9913 | 8.02 | 0.055 | 0.3588 | 0.6506 | 165 | 0.3588 | 0.6506
| 180 |  | 7.9983 | 8.08 | 0.06 | 1.0214 | 0.66372 | 180 | 1.0214 | 0.66372
| 195 |  | 8.0011 | 8.065 | 0.045 | 0.79863 | 0.66651 | 195 | 0.79863 | 0.66651
| 210 |  | 8.0046 | 8.07 | 0.045 | 0.81724 | 0.64216 | 210 | 0.81724 | 0.64216
| 225 |  | 8.0184 | 8.075 | 0.045 | 0.706 | 0.63645 | 225 | 0.706 | 0.63645
| 240 |  | 8.0111 | 8.03 | 0.08 | 0.2357 | 0.68124 | 240 | 0.2357 | 0.68124
| 255 |  | 7.9997 | 8.015 | 0.07 | 0.19076 | 0.66373 | 255 | 0.19076 | 0.66373
| 270 | 8.0038 | 8.0048 | 8 | 0.065 | -0.059495 | 0.66332 | 270 | -0.059495 | 0.66332
| 285 |  | 8.0017 | 8.005 | 0.075 | 0.041548 | 0.66981 | 285 | 0.041548 | 0.66981
| 300 |  | 7.9979 | 8 | 0.09 | 0.025639 | 0.7369 | 300 | 0.025639 | 0.7369
| 315 |  | 7.9978 | 7.985 | 0.065 | -0.16044 | 0.67445 | 315 | -0.16044 | 0.67445
| 330 |  | 7.9922 | 7.99 | 0.05 | -0.02752 | 0.6858 | 330 | -0.02752 | 0.6858
| 345 |  | 8.0011 | 8.015 | 0.075 | 0.17372 | 0.68939 | 345 | 0.17372 | 0.68939
| 360 |  | 7.9983 | 8.02 | 0.065 | 0.27129 | 0.66568 | 360 | 0.27129 | 0.66568
| 375 |  | 7.9913 | 8 | 0.07 | 0.10852 | 0.66769 | 375 | 0.10852 | 0.66769
| 390 |  | 7.9953 | 8.015 | 0.06 | 0.24615 | 0.67923 | 390 | 0.24615 | 0.67923
| 405 |  | 7.9902 | 8.085 | 0.055 | 1.1866 | 0.65162 | 405 | 1.1866 | 0.65162
| 420 |  | 7.9946 | 8.055 | 0.065 | 0.75557 | 0.66071 | 420 | 0.75557 | 0.66071
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 |  | 8.0763 | 7.5358 | 0.070588 | -6.6928 | 0.63283 | 60 | -5.0709 | 0.63011
| 65 |  | 8.1183 | 7.7725 | 0.06705 | -4.2599 | 0.62876 | 75 | -2.8304 | 0.62128
| 80 |  | 8.0983 | 7.927 | 0.038038 | -2.1157 | 0.61755 | 90 | -1.3176 | 0.61778
| 95 |  | 8.0923 | 8.018 | 0.0313 | -0.91856 | 0.61789 | 105 | -0.45072 | 0.61813
| 110 |  | 8.0873 | 8.0698 | 0.027052 | -0.2168 | 0.61825 | 120 | 0.29117 | 0.61846
| 125 |  | 8.0833 | 8.1274 | 0.025687 | 0.54515 | 0.61857 | 135 | 0.99474 | 0.61864
| 140 |  | 8.0823 | 8.1809 | 0.025817 | 1.2195 | 0.61868 | 150 | 1.5656 | 0.61886
| 155 |  | 8.0793 | 8.2198 | 0.023064 | 1.7386 | 0.61896 | 165 | 1.8534 | 0.61902
| 170 |  | 8.0783 | 8.2327 | 0.02385 | 1.9109 | 0.61905 | 180 | 1.7633 | 0.61878
| 185 |  | 8.0833 | 8.2199 | 0.028968 | 1.6895 | 0.61864 | 195 | 1.6639 | 0.61895
| 200 |  | 8.0773 | 8.2107 | 0.024682 | 1.6511 | 0.6191 | 210 | 1.5281 | 0.61909
| 215 |  | 8.0773 | 8.1958 | 0.025296 | 1.4667 | 0.61908 | 225 | 1.2431 | 0.61912
| 230 |  | 8.0763 | 8.1677 | 0.033562 | 1.1313 | 0.61913 | 240 | 0.98683 | 0.61912
| 245 |  | 8.0763 | 8.1502 | 0.038753 | 0.91461 | 0.61912 | 255 | 0.65042 | 0.61906
| 260 |  | 8.0773 | 8.1192 | 0.042194 | 0.51832 | 0.61902 | 270 | 0.34834 | 0.61912
| 275 |  | 8.0753 | 8.0966 | 0.040396 | 0.26335 | 0.61917 | 285 | 0.086684 | 0.61904
| 290 |  | 8.0778 | 8.0777 | 0.026167 | -0.0016506 | 0.61898 | 300 | -0.087482 | 0.61898
| 305 |  | 8.0778 | 8.0673 | 0.020322 | -0.1304 | 0.61898 | 315 | -0.096567 | 0.61903
| 320 |  | 8.0768 | 8.0704 | 0.029995 | -0.079652 | 0.61905 | 330 | -0.16054 | 0.61906
| 335 |  | 8.0768 | 8.0606 | 0.014787 | -0.20099 | 0.61906 | 345 | -0.11266 | 0.619
| 350 |  | 8.0778 | 8.0723 | 0.022704 | -0.0685 | 0.61898 | 360 | -0.17083 | 0.61895
| 365 |  | 8.0783 | 8.0604 | 0.020713 | -0.22199 | 0.61894 | 375 | -0.17495 | 0.61891
| 380 |  | 8.0788 | 8.0666 | 0.025463 | -0.15142 | 0.6189 | 390 | -0.19848 | 0.61895
| 395 |  | 8.0778 | 8.0599 | 0.012234 | -0.22201 | 0.61898 | 405 | -0.22118 | 0.61898
| 410 |  | 8.0778 | 8.06 | 0.0218 | -0.22077 | 0.61898 | 420 | -0.2257 | 0.61893
| 425 |  | 8.0788 | 8.0604 | 0.021974 | -0.22817 | 0.6189 | 435 | -0.20179 | 0.61898
| 440 |  | 8.0773 | 8.0621 | 0.026105 | -0.18859 | 0.61902 | 450 | -0.20759 | 0.61904
| 455 |  | 8.0768 | 8.0593 | 0.017047 | -0.21708 | 0.61906 | 465 | -0.21214 | 0.61908
| 470 |  | 8.0763 | 8.0594 | 0.015927 | -0.20967 | 0.61909 | 480 | -0.2724 | 0.6191
| 485 |  | 8.0763 | 8.0518 | 0.01833 | -0.30377 | 0.6191 | 495 | -0.22865 | 0.61909
| 500 | 8.076 | 8.0763 | 8.0609 | 0.018158 | -0.19109 | 0.61909 | 510 | -0.23734 | 0.61915
| 515 |  | 8.0753 | 8.0543 | 0.019624 | -0.26046 | 0.61917 | 525 | -0.19113 | 0.61922
| 530 |  | 8.0743 | 8.0617 | 0.016728 | -0.15646 | 0.61925
UL-LMPS | 42.5 | 8.1007 | 8.1096 | 7.4327 | 0.015861 | -8.3469 | 0.6509 | 45 | -8.2678 | 0.65605
| 57.5 | 8.0569 | 8.0327 | 7.4003 | 0.022775 | -7.8724 | 0.68183 | 60 | -7.3224 | 0.67801
| 72.5 | 8.0237 | 8.0074 | 7.6413 | 0.027606 | -4.5721 | 0.65892 | 75 | -4.3741 | 0.65892
| 87.5 | 7.9949 | 7.9756 | 7.7057 | 0.027582 | -3.384 | 0.6589 | 90 | -3.2508 | 0.65825
| 102.5 | 7.9629 | 7.969 | 7.763 | 0.022589 | -2.5846 | 0.65498 | 105 | -2.402 | 0.65524
| 117.5 | 7.9277 | 7.9318 | 7.8137 | 0.024709 | -1.4892 | 0.65655 | 120 | -1.4532 | 0.65689
| 132.5 | 7.8881 | 7.8986 | 7.798 | 0.027222 | -1.2733 | 0.65859 | 135 | -1.0813 | 0.72524
| 147.5 | 7.8369 | 7.8492 | 7.8397 | 0.028741 | -0.1212 | 1.0585 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
| 50 |  | 8.139 | 7.83 | 0.066332 | -3.7965 | 2.55 | 60 | -3.8005 | 2.55
SHMU | 65 |  | 8.1499 | 7.84 | 0.050158 | -3.8025 | 2.55 | 75 | -1.8519 | 2.55
| 80 |  | 8.1605 | 8.089 | 0.0493 | -0.87658 | 2.55 | 90 | -0.13435 | 2.55
| 95 |  | 8.1657 | 8.185 | 0.052566 | 0.23677 | 2.55 | 105 | 0.77814 | 2.55
| 110 |  | 8.1357 | 8.221 | 0.048656 | 1.0488 | 2.55 | 120 | 1.2216 | 2.55
| 125 |  | 8.1193 | 8.2255 | 0.029465 | 1.3079 | 2.55 | 135 | 1.2804 | 2.55
| 140 |  | 8.1113 | 8.214 | 0.031018 | 1.2667 | 2.55 | 150 | 1.1524 | 2.55
| 155 |  | 8.0938 | 8.1825 | 0.050354 | 1.0953 | 2.55 | 165 | 0.92052 | 2.55
| 170 |  | 8.0827 | 8.15 | 0.040131 | 0.83314 | 2.55 | 180 | 0.40883 | 2.55
| 185 |  | 8.0856 | 8.1015 | 0.031502 | 0.19667 | 2.55 | 195 | 0.0059272 | 2.55
| 200 |  | 8.0907 | 8.0835 | 0.029607 | -0.089446 | 2.55 | 210 | -0.062958 | 2.55
| 215 |  | 8.104 | 8.1 | 0.031623 | -0.049713 | 2.55 | 225 | -0.19571 | 2.55
| 230 |  | 8.1339 | 8.112 | 0.027261 | -0.26871 | 2.55 | 240 | -0.4237 | 2.55
| 245 |  | 8.1549 | 8.114 | 0.034702 | -0.50119 | 2.55 | 255 | -0.42677 | 2.55
| 260 |  | 8.1598 | 8.128 | 0.041877 | -0.38955 | 2.55 | 270 | -0.56353 | 2.55
| 275 |  | 8.1636 | 8.1105 | 0.059866 | -0.65051 | 2.55 | 285 | -0.35773 | 2.55
| 290 |  | 8.1512 | 8.134 | 0.035452 | -0.21134 | 2.55 | 300 | -0.19888 | 2.55
| 305 |  | 8.1302 | 8.1145 | 0.043465 | -0.19264 | 2.55 | 315 | -0.16547 | 2.55
| 320 | 8.1133 | 8.1133 | 8.101 | 0.028819 | -0.15189 | 2.55 | 330 | -0.33177 | 2.55
| 335 |  | 8.0806 | 8.0465 | 0.035433 | -0.42171 | 2.55 | 345 | -0.2686 | 2.55
| 350 |  | 8.063 | 8.0475 | 0.031602 | -0.19205 | 2.55 | 360 | -0.23414 | 2.55
| 365 |  | 8.0861 | 8.0655 | 0.033791 | -0.25518 | 2.55 | 375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 8.067 | 7.37 | 0.029 | -8.64 | 0.64 | 45 | -8.64 | 0.64
| 60 |  | 8.066 | 7.586 | 0.029 | -5.951 | 0.64 | 60 | -5.951 | 0.64
LEI | 75 |  | 8.063 | 7.745 | 0.021 | -3.944 | 0.64 | 75 | -3.944 | 0.64
| 90 |  | 8.066 | 7.799 | 0.018 | -3.31 | 0.64 | 90 | -3.31 | 0.64
| 105 |  | 8.061 | 7.879 | 0.021 | -2.258 | 0.64 | 105 | -2.258 | 0.64
| 120 |  | 8.062 | 7.953 | 0.02 | -1.352 | 0.64 | 120 | -1.352 | 0.64
| 135 |  | 8.066 | 8.025 | 0.024 | -0.508 | 0.64 | 135 | -0.508 | 0.64
| 150 |  | 8.062 | 8.146 | 0.086 | 1.042 | 0.64 | 150 | 1.042 | 0.64
| 165 |  | 8.059 | 8.136 | 0.085 | 0.955 | 0.64 | 165 | 0.955 | 0.64
| 180 |  | 8.057 | 8.097 | 0.03 | 0.496 | 0.64 | 180 | 0.496 | 0.64
| 195 |  | 8.052 | 8.043 | 0.035 | -0.112 | 0.64 | 195 | -0.112 | 0.64
| 210 |  | 8.053 | 8.018 | 0.035 | -0.435 | 0.64 | 210 | -0.435 | 0.64
| 225 |  | 8.054 | 7.986 | 0.042 | -0.844 | 0.64 | 225 | -0.844 | 0.64
| 240 |  | 8.049 | 7.982 | 0.035 | -0.832 | 0.64 | 240 | -0.832 | 0.64
| 255 |  | 8.049 | 7.962 | 0.039 | -1.081 | 0.64 | 255 | -1.081 | 0.64
| 270 |  | 8.046 | 7.945 | 0.155 | -1.255 | 0.64 | 270 | -1.255 | 0.64
| 285 |  | 8.045 | 7.972 | 0.037 | -0.907 | 0.64 | 285 | -0.907 | 0.64
| 300 |  | 8.046 | 7.976 | 0.032 | -0.87 | 0.64 | 300 | -0.87 | 0.64
| 315 |  | 8.043 | 7.955 | 0.04 | -1.094 | 0.64 | 315 | -1.094 | 0.64
| 330 |  | 8.045 | 7.888 | 0.037 | -1.952 | 0.64 | 330 | -1.952 | 0.64
| 345 |  | 8.039 | 7.605 | 0.076 | -5.399 | 0.64 | 345 | -5.399 | 0.64
| 360 |  | 8.041 | 7.014 | 0.077 | -12.772 | 0.64 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 12
| 12 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.17963
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 11.9591 | 11.171 | 0.026013 | -6.5901 | 0.64386 | 45 | -6.4034 | 0.64402
| 52.5 |  | 11.9531 | 11.21 | 0.026247 | -6.2167 | 0.64418 | 60 | -5.3314 | 0.64398 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 11.9608 | 11.429 | 0.024698 | -4.4461 | 0.64377 | 75 | -3.5839 | 0.64771 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 11.9698 | 11.644 | 0.021187 | -2.7218 | 0.65164 | 90 | -2.3865 | 0.64758 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 11.9654 | 11.72 | 0.01633 | -2.0513 | 0.64352 | 105 | -1.6834 | 0.64751 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 11.9725 | 11.815 | 0.020683 | -1.3154 | 0.65149 | 120 | -0.76809 | 0.6519 |  | 0.24063 | 0.34273 | 0.72791 | 3.9193 | 0.055955 | 0.7087 | 1.1525 | 4.8571 | 0.20787 | 0.45386 | 4.0329 | 0.45209 | 3.4809 | 0.87615 | 4.3684 |  |  |  | 0.9128 | 1.3295 | 0.67434 | 0.24166 | 3.0694 | 1.0656
| 127.5 | 11.966 | 11.9574 | 11.931 | 0.019692 | -0.22075 | 0.65232 | 135 | -0.010649 | 10.3521 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 11.9739 | 11.9978 | 3.7941 | 0.19945 | 20.052 | 150 | 0.79395 | 10.3565 | 45 |  | 2.3747 |  | 3.3441 | 2.6948 | [5 empty] | 1.3583 | 0.40348 |  |  | 0.98805
| 157.5 |  | 11.9511 | 12.117 | 0.028304 | 1.3885 | 0.66103 | 165 | 1.2408 | 0.66498 | 60 | 0.55591 | 2.3464 | 1.1236 | 3.1886 | 1.6253 | 2.2444 | 0.75013 | 3.1851 | 1.3547 | 1.2521 | 1.224 | 0.78337 | 2.2617 | 0.49477 | 1.8846
| 172.5 |  | 11.9593 | 12.09 | 0.03559 | 1.0931 | 0.66894 | 180 | 0.88836 | 0.66872 | 75 | 0.045478 | 0.53717 | 0.40043 | 1.5477 | 0.60609 | 0.61123 | 0.43528 | 1.7724 | 0.69207 | 0.13076 | 1.1893 | 0.087513 | 1.2569 | 0.21171 | 1.0968
| 187.5 |  | 11.9672 | 12.049 | 0.03414 | 0.68366 | 0.66849 | 195 | 0.46935 | 0.673 | 90 | 0.48591 | 0.58007 | 0.21842 | 1.4956 | 1.1292 | 0.14886 | 0.80139 | 1.2754 | 0.82991 | 0.88101 | 1.0886 | 0.63706 | 1.818 | 1.4729 | 0.5207
| 202.5 |  | 11.9555 | 11.986 | 0.03893 | 0.25503 | 0.67751 | 210 | -0.25961 | 0.68468 | 105 | 0.38987 | 0.30721 | 0.54733 | 1.3458 | 0.7434 | 0.076237 | 1.0939 | 1.1947 | 0.46623 | 0.95607 | 1.1923 | 0.50353 | 1.9857 | 1.4192 | 0.73697
| 217.5 |  | 11.9969 | 11.904 | 0.052324 | -0.77426 | 0.69185 | 225 |  |  | 120 | 0.55165 | 0.4282 | 0.19065 | 1.1668 | 0.43364 | 0.11667 | 0.84308 | 0.82568 | 0.1 | 0.68589 | 0.87548 | 0.013179 | 1.4427 | 0.68687 | 0.85174 |  |  |  | 1.6534 | 1.3998 | 1.4155 | 1.9871 | 0.43404 | 1.3814
[8 empty] | 240 |  |  | 135 | 0.057102 | 0.028626 | 0.014815 | 0.14913 | 0.038715 | 0.43554 | 0.60955 | 1.1518 | 0.27502 | 0.18631 | 1.4326 | 0.14072 | 1.5392 | 0.31706 | 1.2956 |  |  |  | 0.15041 | 1.2514 | 1.5406 | 1.6458 | 0.013275 | 1.3921
[8 empty] | 255 |  |  | 150 | 0.091902 | 0.029533 | 0.06075 |  | 0.030465 | 0.95452 | 0.4487 |  | 0.91997 | 0.42155 |  | 0.013041 |  | 0.40204 |  |  |  |  | 0.17631 | 1.1781 | 1.9232 | 1.4485 |  | 1.8806
[8 empty] | 270 |  |  | 165 | 1.7395 | 0.60461 | 1.1116 |  | 0.90232 | 1.3127 | 0.58685 |  | 0.90414 | 0.61007 |  | 0.3553 |  | 0.26182 |  |  |  |  | 1.4841 | 0.065047 | 1.0499 | 0.44607 |  | 0.7096
[8 empty] | 285 |  |  | 180 | 1.8741 | 0.32724 | 0.75589 |  | 0.91157 | 1.8268 | 1.1809 |  | 1.0556 | 0.50928 |  | 0.68585 |  | 0.15175 |  |  |  |  | 0.45579 | 1.4387 | 0.17471 | 0.2956 |  | 0.4479
[8 empty] | 300 |  |  | 195 | 1.949 | 0.19761 | 0.29621 |  | 0.71135 | 2.0906 | 1.8381 |  | 1.3971 | 0.12286 |  | 0.60385 |  | 0.45356 |  |  |  |  | 0.37781 | 1.5751 | 0.22528 | 0.1066 |  | 0.31323
[8 empty] | 315 |  |  | 210 | 1.5414 | 0.082844 | 0.34415 |  | 0.098612 | 1.9793 | 2.1951 |  | 1.9581 | 0.30946 |  | 0.019933 |  | 0.28529 |  |  |  |  | 0.024744 | 1.6098 | 0.056445 | 0.32414 |  | 0.072901
[8 empty] | 330 |  |  | 225 | [5 empty] | 2.3795 | 2.711 |  | 2.0741 | 0.43955 |  | 0.23662 |  | 0.65719 | [5 empty] | 2.0719 | 0.028844 | 0.42591 |  | 0.18468
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.6323 | 2.9787 |  | 1.7689 | 0.46745 |  | 0.75103 |  | 1.1738 | [5 empty] | 1.7686 | 0.3963 | 0.78394 |  | 0.24998
[8 empty] | 360 |  |  | 255 | [5 empty] | 3.0756 | 3.1893 |  | 2.0858 | 0.274 |  | 0.85994 |  | 1.0872 | [5 empty] | 1.8985 | 0.70837 | 0.92325 |  | 0.059031
[8 empty] | 375 |  |  | 270 | [5 empty] | 3.3456 | 3.2536 |  | 2.0164 | 0.096402 |  | 1.1704 |  | 1.2077 | [5 empty] | 2.3506 | 0.52439 | 0.58936 |  | 0.52181
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.069223 |  | 1.2523 |  | 1.1227 | [6 empty] | 0.50967 | 0.42799 |  | 0.60042
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.02038 |  | 0.75899 |  | 0.74179 | [6 empty] | 0.26581 | 0.2733 |  | 0.41085
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.015022 |  | 0.81115 |  | 0.75661 | [6 empty] | 0.37657 | 0.34792 |  | 0.34954
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.10125 |  | 0.85574 |  | 0.71425 | [6 empty] | 0.19757 | 0.10027 |  | 0.55613
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.070471 |  | 0.81462 |  | 0.70548 | [6 empty] | 0.19535 | 0.1246 |  | 0.52963
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.066338 | [10 empty] | 0.098655 | 0.034342
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.22608 | [10 empty] | 0.19853 | 0.0086712
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.44978 | [10 empty] | 0.37646 | 0.049387
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.73386 | [10 empty] | 1.3034 | 0.58691
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.56372
435
450
45
PTB | 55 | 11.9658 | 12.0283 | 11.2231 | 0.037285 | -6.6948 | 0.42 | 60 | -5.7703 | 0.42
| 70 | 11.9658 | 12.0152 | 11.544 | 0.036968 | -3.9215 | 0.42 | 75 | -3.62 | 0.42
| 85 | 11.9658 | 12.0009 | 11.6388 | 0.028914 | -3.0169 | 0.42 | 90 | -2.7716 | 0.42
| 100 | 11.9658 | 11.992 | 11.7185 | 0.028662 | -2.281 | 0.42 | 105 | -1.9923 | 0.42
| 115 | 11.9658 | 11.9892 | 11.8196 | 0.027581 | -1.415 | 0.42 | 120 | -1.2072 | 0.42
| 130 | 11.9658 | 11.9764 | 11.8816 | 0.037381 | -0.79174 | 0.42 | 135 | -0.60235 | 0.42
| 145 | 11.9658 | 11.9704 | 11.9436 | 0.029422 | -0.22357 | 0.42 | 150 | -0.15875 | 0.42
| 160 | 11.9658 | 11.9658 | 11.9623 | 0.043387 | -0.029121 | 0.42 | 165 | -0.16563 | 0.42333
| 175 | 11.9658 | 11.9741 | 11.9216 | 0.055353 | -0.43866 | 0.43 | 180 | -0.6359 | 0.42667
| 190 | 11.9658 | 11.9811 | 11.8577 | 0.046674 | -1.0304 | 0.42 | 195 | -1.116 | 0.42
| 205 | 11.9658 | 11.9847 | 11.8304 | 0.039506 | -1.2872 | 0.42 | 210 | -1.5283 | 0.42
| 220 | 11.9658 | 11.9978 | 11.7565 | 0.052531 | -2.0105 | 0.42 | 225 | -2.1206 | 0.42
| 235 | 11.9658 | 12.0084 | 11.7273 | 0.051966 | -2.3408 | 0.42 | 240 | -2.4954 | 0.42
| 250 | 11.9658 | 12.0226 | 11.6854 | 0.04101 | -2.8046 | 0.42 | 255 | -2.7567 | 0.42
| 265 | 11.9658 | 12.0344 | 11.7142 | 0.039663 | -2.6608 | 0.42 | 270 | -2.8987 | 0.42
| 280 | 11.9658 | 12.0434 | 11.637 | 0.036864 | -3.3743 | 0.42 | 285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 12.0888 | 11.075 | 0.033166 | -8.386 | 0.5 | 45 | -8.386 | 0.5
| 60 |  | 12.0964 | 11.2145 | 0.030345 | -7.2904 | 0.5 | 60 | -7.2904 | 0.5
| 75 |  | 12.0892 | 11.6015 | 0.03265 | -4.0339 | 0.5 | 75 | -4.0339 | 0.5
| 90 |  | 12.0934 | 11.746 | 0.025833 | -2.8724 | 0.5 | 90 | -2.8724 | 0.5
| 105 |  | 12.0871 | 11.8525 | 0.030758 | -1.9407 | 0.5 | 105 | -1.9407 | 0.5
| 120 |  | 12.0874 | 11.951 | 0.03323 | -1.1282 | 0.5 | 120 | -1.1282 | 0.5
| 135 |  | 12.0761 | 12.039 | 0.029001 | -0.30738 | 0.5 | 135 | -0.30738 | 0.5
| 150 |  | 12.0731 | 12.132 | 0.031221 | 0.48769 | 0.5 | 150 | 0.48769 | 0.5
| 165 |  | 12.0748 | 12.1625 | 0.03447 | 0.72614 | 0.5 | 165 | 0.72614 | 0.5
| 180 |  | 12.0765 | 12.15 | 0.036563 | 0.60887 | 0.5 | 180 | 0.60887 | 0.5
| 195 |  | 12.0803 | 12.1165 | 0.030826 | 0.29991 | 0.5 | 195 | 0.29991 | 0.5
| 210 |  | 12.0702 | 12.0475 | 0.031098 | -0.18782 | 0.5 | 210 | -0.18782 | 0.5
| 225 | 12.0745 | 12.0745 | 12.013 | 0.04578 | -0.50909 | 0.5 | 225 | -0.50909 | 0.5
| 240 |  | 12.0841 | 11.998 | 0.039815 | -0.71267 | 0.5 | 240 | -0.71267 | 0.5
| 255 |  | 12.0854 | 12.004 | 0.029629 | -0.6737 | 0.5 | 255 | -0.6737 | 0.5
| 270 |  | 12.0749 | 11.9985 | 0.031166 | -0.63288 | 0.5 | 270 | -0.63288 | 0.5
| 285 |  | 12.0771 | 12.009 | 0.033388 | -0.56404 | 0.5 | 285 | -0.56404 | 0.5
| 300 |  | 12.0866 | 11.996 | 0.034397 | -0.74975 | 0.5 | 300 | -0.74975 | 0.5
| 315 |  | 12.0864 | 11.9965 | 0.032971 | -0.74398 | 0.5 | 315 | -0.74398 | 0.5
| 330 |  | 12.0843 | 11.9925 | 0.043995 | -0.75983 | 0.5 | 330 | -0.75983 | 0.5
| 345 |  | 12.0851 | 11.9925 | 0.02593 | -0.7664 | 0.5 | 345 | -0.7664 | 0.5
| 360 |  | 12.0935 | 12.0035 | 0.029961 | -0.74437 | 0.5 | 360 | -0.74437 | 0.5
| 375 |  | 12.0964 | 11.9765 | 0.039239 | -0.99096 | 0.5 | 375 | -0.99096 | 0.5
| 390 |  | 12.0948 | 11.958 | 0.040079 | -1.1308 | 0.5 | 390 | -1.1308 | 0.5
| 405 |  | 12.0994 | 11.932 | 0.028023 | -1.3833 | 0.5 | 405 | -1.3833 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 11.996 | 11.725 | 0.075 | -2.259 | 0.59654 | 120 | -2.259 | 0.59654
| 135 |  | 11.9934 | 11.805 | 0.08 | -1.5708 | 0.62468 | 135 | -1.5708 | 0.62468
| 150 |  | 11.9942 | 11.87 | 0.055 | -1.0352 | 0.58721 | 150 | -1.0352 | 0.58721
| 165 |  | 11.999 | 11.985 | 0.065 | -0.11639 | 0.60143 | 165 | -0.11639 | 0.60143
| 180 |  | 11.9989 | 12.055 | 0.075 | 0.46736 | 0.61133 | 180 | 0.46736 | 0.61133
| 195 |  | 12.0014 | 12.015 | 0.08 | 0.11372 | 0.6331 | 195 | 0.11372 | 0.6331
| 210 |  | 12.0183 | 11.99 | 0.095 | -0.23566 | 0.66038 | 210 | -0.23566 | 0.66038
| 225 |  | 12.019 | 11.955 | 0.09 | -0.53254 | 0.61546 | 225 | -0.53254 | 0.61546
| 240 |  | 12.0216 | 11.895 | 0.13 | -1.0532 | 0.67544 | 240 | -1.0532 | 0.67544
| 255 |  | 12.0117 | 11.86 | 0.075 | -1.2632 | 0.64045 | 255 | -1.2632 | 0.64045
| 270 | 12.007 | 12.008 | 11.88 | 0.1 | -1.0658 | 0.63191 | 270 | -1.0658 | 0.63191
| 285 |  | 12.0087 | 11.89 | 0.12 | -0.98811 | 0.64034 | 285 | -0.98811 | 0.64034
| 300 |  | 11.9963 | 11.88 | 0.1 | -0.96943 | 0.63303 | 300 | -0.96943 | 0.63303
| 315 |  | 11.9914 | 11.865 | 0.08 | -1.0545 | 0.63061 | 315 | -1.0545 | 0.63061
| 330 |  | 11.9959 | 11.885 | 0.1 | -0.92482 | 0.64432 | 330 | -0.92482 | 0.64432
| 345 |  | 12.0014 | 11.89 | 0.1 | -0.92783 | 0.63297 | 345 | -0.92783 | 0.63297
| 360 |  | 11.9989 | 11.9 | 0.075 | -0.82442 | 0.61333 | 360 | -0.82442 | 0.61333
| 375 |  | 11.999 | 11.9 | 0.08 | -0.82478 | 0.6468 | 375 | -0.82478 | 0.6468
| 390 |  | 11.9942 | 11.895 | 0.08 | -0.82681 | 0.60817 | 390 | -0.82681 | 0.60817
| 405 |  | 11.9934 | 11.955 | 0.09 | -0.32008 | 0.61898 | 405 | -0.32008 | 0.61898
| 420 |  | 11.996 | 11.955 | 0.12 | -0.34168 | 0.64754 | 420 | -0.34168 | 0.64754
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 |  | 11.9753 | 10.9941 | 0.19777 | -8.1938 | 0.57795 | 60 | -6.3169 | 0.56762
| 65 |  | 12.1063 | 11.4552 | 0.12879 | -5.3785 | 0.56245 | 75 | -3.9333 | 0.55641
| 80 |  | 12.1133 | 11.7244 | 0.080383 | -3.2108 | 0.55339 | 90 | -2.1963 | 0.55374
| 95 |  | 12.0973 | 11.893 | 0.047383 | -1.6891 | 0.55392 | 105 | -1.2066 | 0.55434
| 110 |  | 12.0823 | 11.9657 | 0.040462 | -0.96532 | 0.55455 | 120 | -0.60134 | 0.55479
| 125 |  | 12.0743 | 12.0237 | 0.037146 | -0.41935 | 0.5549 | 135 | -0.16426 | 0.55487
| 140 |  | 12.0753 | 12.0709 | 0.043718 | -0.036714 | 0.55485 | 150 | 0.1638 | 0.55507
| 155 |  | 12.0683 | 12.1002 | 0.045772 | 0.26405 | 0.55517 | 165 | 0.25738 | 0.55508
| 170 |  | 12.0713 | 12.102 | 0.036856 | 0.25405 | 0.55504 | 180 | 0.21758 | 0.555
| 185 |  | 12.0723 | 12.0964 | 0.034348 | 0.19935 | 0.55499 | 195 | 0.2055 | 0.55517
| 200 |  | 12.0663 | 12.0915 | 0.032437 | 0.20857 | 0.55527 | 210 | 0.050001 | 0.55526
| 215 |  | 12.0663 | 12.0628 | 0.038603 | -0.029283 | 0.55526 | 225 | -0.17129 | 0.5553
| 230 |  | 12.0653 | 12.0361 | 0.036361 | -0.24229 | 0.55531 | 240 | -0.3534 | 0.55537
| 245 |  | 12.0633 | 12.014 | 0.03373 | -0.40895 | 0.55541 | 255 | -0.46311 | 0.55541
| 260 |  | 12.0633 | 12.0042 | 0.036618 | -0.49019 | 0.55541 | 270 | -0.55878 | 0.55547
| 275 |  | 12.0613 | 11.9898 | 0.029431 | -0.59308 | 0.5555 | 285 | -0.61725 | 0.55548
| 290 |  | 12.0621 | 11.9862 | 0.027526 | -0.62934 | 0.55547 | 300 | -0.73409 | 0.55542
| 305 |  | 12.0638 | 11.9689 | 0.0263 | -0.78647 | 0.5554 | 315 | -0.75552 | 0.5554
| 320 |  | 12.0638 | 11.9745 | 0.027008 | -0.74005 | 0.5554 | 330 | -0.83764 | 0.55538
| 335 |  | 12.0644 | 11.9575 | 0.022401 | -0.88644 | 0.55537 | 345 | -0.82055 | 0.55528
| 350 |  | 12.0674 | 11.9724 | 0.017517 | -0.78761 | 0.55523 | 360 | -0.79534 | 0.55523
| 365 |  | 12.0674 | 11.971 | 0.024154 | -0.79921 | 0.55523 | 375 | -0.81723 | 0.55521
| 380 |  | 12.0681 | 11.9684 | 0.017079 | -0.82624 | 0.5552 | 390 | -0.78519 | 0.55521
| 395 |  | 12.0678 | 11.9755 | 0.01713 | -0.76466 | 0.55521 | 405 | -0.81935 | 0.55522
| 410 |  | 12.0678 | 11.9656 | 0.02324 | -0.8467 | 0.55522 | 420 | -0.83306 | 0.55521
| 425 |  | 12.0681 | 11.9684 | 0.018239 | -0.82624 | 0.5552 | 435 | -0.865 | 0.55526
| 440 |  | 12.0661 | 11.9594 | 0.021541 | -0.88439 | 0.5553 | 450 | -0.87903 | 0.55528
| 455 |  | 12.0664 | 11.9607 | 0.022713 | -0.87635 | 0.55528 | 465 | -0.84282 | 0.55527
| 470 |  | 12.0668 | 11.9671 | 0.02476 | -0.82605 | 0.55526 | 480 | -0.86732 | 0.55527
| 485 |  | 12.0664 | 11.9593 | 0.024329 | -0.88795 | 0.55528 | 495 | -0.86793 | 0.55535
| 500 | 12.064 | 12.064 | 11.9605 | 0.019223 | -0.85792 | 0.55539 | 510 | -0.86579 | 0.55532
| 515 |  | 12.0664 | 11.9615 | 0.025966 | -0.86972 | 0.55528 | 525 | -0.87933 | 0.5553
| 530 |  | 12.0658 | 11.9591 | 0.020234 | -0.88414 | 0.55531
UL-LMPS | 42.5 | 12.1997 | 12.2347 | 11.0543 | 0.032326 | -9.648 | 0.65213 | 45 | -9.5301 | 0.65358
| 57.5 | 12.1372 | 12.1313 | 11.0467 | 0.047736 | -8.9406 | 0.6608 | 60 | -8.3278 | 0.66034
| 72.5 | 12.0766 | 12.0742 | 11.4387 | 0.043392 | -5.2638 | 0.65803 | 75 | -5.0397 | 0.65802
| 87.5 | 12.0115 | 12.0004 | 11.53 | 0.043319 | -3.9197 | 0.65799 | 90 | -3.7922 | 0.65706
| 102.5 | 11.9597 | 11.9676 | 11.59 | 0.032991 | -3.1549 | 0.65244 | 105 | -2.9441 | 0.65276
| 117.5 | 11.9047 | 11.9261 | 11.7007 | 0.036921 | -1.8905 | 0.65438 | 120 | -1.8666 | 0.65506
| 132.5 | 11.8242 | 11.865 | 11.6577 | 0.044066 | -1.7474 | 0.65844 | 135 | -1.5582 | 0.6928
| 147.5 | 11.7483 | 11.7835 | 11.7113 | 0.037571 | -0.61204 | 0.86461 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 12.134 | 11.08 | 0.128 | -8.686 | 0.52 | 45 | -8.686 | 0.52
| 60 |  | 12.135 | 11.321 | 0.039 | -6.708 | 0.52 | 60 | -6.708 | 0.52
LEI | 75 |  | 12.126 | 11.629 | 0.03 | -4.099 | 0.52 | 75 | -4.099 | 0.52
| 90 |  | 12.133 | 11.727 | 0.026 | -3.346 | 0.52 | 90 | -3.346 | 0.52
| 105 |  | 12.14 | 11.859 | 0.029 | -2.315 | 0.52 | 105 | -2.315 | 0.52
| 120 |  | 12.123 | 11.985 | 0.022 | -1.138 | 0.52 | 120 | -1.138 | 0.52
| 135 |  | 12.13 | 12.08 | 0.029 | -0.412 | 0.52 | 135 | -0.412 | 0.52
| 150 |  | 12.123 | 12.181 | 0.027 | 0.478 | 0.52 | 150 | 0.478 | 0.52
| 165 |  | 12.12 | 12.176 | 0.062 | 0.462 | 0.52 | 165 | 0.462 | 0.52
| 180 |  | 12.119 | 12.131 | 0.056 | 0.099 | 0.52 | 180 | 0.099 | 0.52
| 195 |  | 12.12 | 12.102 | 0.032 | -0.149 | 0.52 | 195 | -0.149 | 0.52
| 210 |  | 12.119 | 12.098 | 0.032 | -0.173 | 0.52 | 210 | -0.173 | 0.52
| 225 |  | 12.11 | 12.027 | 0.056 | -0.685 | 0.52 | 225 | -0.685 | 0.52
| 240 |  | 12.113 | 11.959 | 0.054 | -1.271 | 0.52 | 240 | -1.271 | 0.52
| 255 |  | 12.108 | 11.949 | 0.028 | -1.313 | 0.52 | 255 | -1.313 | 0.52
| 270 |  | 12.107 | 11.925 | 0.032 | -1.503 | 0.52 | 270 | -1.503 | 0.52
| 285 |  | 12.107 | 11.926 | 0.033 | -1.495 | 0.52 | 285 | -1.495 | 0.52
| 300 |  | 12.098 | 11.939 | 0.029 | -1.314 | 0.52 | 300 | -1.314 | 0.52
| 315 |  | 12.097 | 11.934 | 0.028 | -1.347 | 0.52 | 315 | -1.347 | 0.52
| 330 |  | 12.104 | 11.935 | 0.034 | -1.396 | 0.52 | 330 | -1.396 | 0.52
| 345 |  | 12.1 | 11.934 | 0.043 | -1.372 | 0.52 | 345 | -1.372 | 0.52
| 360 |  | 12.093 | 11.856 | 0.043 | -1.96 | 0.52 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-S 20
| 20 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.40825
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 19.8991 | 18.701 | 0.072026 | -6.0207 | 0.66335 | 45 | -6.2346 | 0.65073
| 52.5 |  | 19.9024 | 18.619 | 0.052377 | -6.4485 | 0.63811 | 60 | -5.5724 | 0.66341 | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 19.8922 | 18.958 | 0.091141 | -4.6962 | 0.68871 | 75 | -3.5965 | 0.66577 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 19.9121 | 19.415 | 0.045522 | -2.4967 | 0.64282 | 90 | -2.1025 | 0.64289 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 19.9081 | 19.568 | 0.04077 | -1.7084 | 0.64295 | 105 | -1.2305 | 0.64284 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 19.9149 | 19.765 | 0.037786 | -0.75268 | 0.64273 | 120 | -0.32074 | 0.64283 |  | 0.78077 | 1.814 | 1.3707 | 4.6984 | 1.4953 | 1.2504 | 0.74785 | 4.5708 | 0.87277 | 0.43039 | 3.3313 | 0.36833 | 3.6181 | 0.078959 | 3.6789 |  |  | 0.71839 | 2.0694 | 1.5622 | 0.3607 | 0.75961 | 2.8777 | 0.71356 | [6 empty] | 0.29441 | 0.056687 | 0.56335 | 0.36983 | 2.11 | 0.40781
| 127.5 | 19.9049 | 19.9089 | 19.931 | 0.039001 | 0.1112 | 0.64293 | 135 | 0.35615 | 0.64522 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 19.9222 | 20.042 | 0.044171 | 0.60111 | 0.64752 | 150 | 0.92826 | 0.65566 | 45 |  | 2.6346 |  | 3.4244 | 2.3392 | [5 empty] | 1.1259 | 0.34971 |  |  | 1.4438
| 157.5 |  | 19.8854 | 20.135 | 0.058547 | 1.2554 | 0.66381 | 165 | 0.99991 | 0.671 | 60 | 0.35706 | 2.0552 | 1.4489 | 3.2067 | 0.19493 | 2.8623 | 2.0991 | 4.0576 | 0.17982 | 0.60994 | 1.4761 | 2.5687 | 1.9772 | 1.8523 | 3.7688 | [15 empty] | 0.77815 | 0.63768 | 1.7929 | 1.5104 | 2.4433 | 0.70527
| 172.5 |  | 19.9058 | 20.054 | 0.074416 | 0.74441 | 0.67819 | 180 | 0.488 | 0.67056 | 75 | 0.43784 | 0.71424 | 0.87665 | 1.8943 | 0.31188 | 1.3572 | 1.5191 | 2.6296 | 0.8813 | 0.20574 | 1.3787 | 0.46246 | 1.1545 | 0.65273 | 1.7977 | [15 empty] | 0.67107 | 0.48881 | 1.0349 | 1.1183 | 1.6573 | 0.84008
| 187.5 |  | 19.9119 | 19.958 | 0.064429 | 0.23159 | 0.66292 | 195 | 0.036239 | 0.66086 | 90 | 0.10998 | 1.1957 | 0.64894 | 2.3156 | 0.88944 | 1.2805 | 0.63494 | 2.5256 | 0.92209 | 0.56712 | 1.3422 | 0.3503 | 1.8124 | 0.23357 | 1.6618 | [15 empty] | 0.72468 | 0.79813 | 1.3104 | 1.0519 | 1.909 | 1.1635
| 202.5 |  | 19.8846 | 19.853 | 0.060194 | -0.15911 | 0.6588 | 210 | -0.59058 | 0.65782 | 105 | 0.22438 | 1.2144 | 0.44609 | 2.1227 | 0.16616 | 1.1734 | 0.27541 | 2.1896 | 0.059499 | 0.81143 | 1.1192 | 1.1795 | 1.8099 | 0.31885 | 2.1666 | [15 empty] | 0.65768 | 0.78181 | 1.2507 | 0.885 | 1.7473 | 0.7519
| 217.5 |  | 19.9438 | 19.74 | 0.061644 | -1.022 | 0.65684 | 225 |  |  | 120 | 0.32861 | 1.0404 | 0.80705 | 1.9489 | 0.35035 | 0.85145 | 0.58199 | 1.885 | 0.038842 | 0.22415 | 1.1181 | 0.77802 | 1.2823 | 0.52336 | 1.7937 |  |  | 1.514 | 2.4145 | 2.4437 | 1.5654 | 1.7269 | 0.29108 | 2.3143 | [6 empty] | 0.32242 | 0.48415 | 0.82517 | 0.72073 | 1.3375 | 0.49612
[8 empty] | 240 |  |  | 135 | 0.69207 | 0.92725 | 1.0471 | 2.0103 | 0.2614 | 0.30922 | 0.46816 | 1.6222 | 0.48011 | 0.16348 | 1.3358 | 0.75163 | 1.1714 | 0.8883 | 1.9417 |  |  | 0.98159 | 2.172 | 1.7783 | 1.4174 | 1.2135 | 0.1427 | 2.1433 | [6 empty] | 0.08483 | 0.23248 | 0.35582 | 0.42528 | 1.018 | 0.03733
[8 empty] | 255 |  |  | 150 | 1.4316 | 0.91527 | 1.2952 |  | 0.52724 | 0.54593 | 0.047063 |  | 1.0183 | 0.4547 |  | 0.44406 |  | 0.88541 |  |  |  | 0.56736 | 1.9136 | 0.70623 | 1.1615 | 0.6956 |  | 1.5851 | [6 empty] | 0.38553 | 0.26933 | 0.04549 | 0.2453 |  | 0.1429
[8 empty] | 270 |  |  | 165 | 1.9025 | 1.0638 | 1.1858 |  | 0.98397 | 0.91938 | 0.70141 |  | 1.0305 | 0.17235 |  | 0.098202 |  | 0.26809 |  |  |  | 0.13257 | 1.3992 | 0.43646 | 0.41913 | 0.23955 |  | 0.51591 | [6 empty] | 0.55631 | 0.32242 | 0.053359 | 0.02288 |  | 0.095106
[8 empty] | 285 |  |  | 180 | 1.8594 | 1.0972 | 0.66235 |  | 0.91309 | 0.82912 | 1.2737 |  | 1.0652 | 0.45604 |  | 0.21751 |  | 0.24982 |  |  |  | 0.32697 | 0.20739 | 1.7515 | 0.9377 | 0.4785 |  | 0.74047 | [6 empty] | 0.42598 | 0.43681 | 0.096615 | 0.1048 |  | 0.0044398
[8 empty] | 300 |  |  | 195 | 1.9588 | 0.82418 | 0.383 |  | 1.064 | 1.2607 | 1.7006 |  | 0.99056 | 0.46827 |  | 0.26637 |  | 0.72795 |  |  |  | 0.060847 | 0.6131 | 1.3714 | 0.18932 | 0.25751 |  | 0.4399 | [6 empty] | 0.24411 | 0.66518 | 0.1479 | 0.059269 |  | 0.26114
[8 empty] | 315 |  |  | 210 | 1.6424 | 0.7813 | 0.16854 |  | 0.71207 | 0.94688 | 1.9969 |  | 1.0427 | 1.0381 |  | 0.08347 |  | 0.96497 |  |  |  | 0.2267 | 0.30348 | 1.416 | 0.49832 | 0.5074 |  | 0.42251 | [6 empty] | 0.07468 | 0.84469 | 0.45308 | 0.0071121 |  | 0.41827
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.1487 | 2.2646 |  | 1.2044 | 1.1088 |  | 0.043807 |  | 1.0744 |  |  |  | 0.43044 |  | 1.2897 | 0.19429 | 0.86992 |  | 0.15399 | [7 empty] | 0.9967 | 0.52188 | 0.030031 |  | 0.504
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.3826 | 2.4457 |  | 1.579 | 1.0677 |  | 0.1755 |  | 0.90621 |  |  |  | 0.60378 |  | 1.2836 | 0.0079114 | 1.0158 |  | 0.17282 | [7 empty] | 1.1787 | 0.60737 | 0.13299 |  | 0.53371
[8 empty] | 360 |  |  | 255 | [5 empty] | 1.786 | 2.5118 |  | 1.9141 | 0.76173 |  | 0.1075 |  | 0.66357 |  |  |  | 0.57969 |  | 1.4989 | 0.1709 | 0.88774 |  | 0.27353 | [7 empty] | 1.2443 | 0.50878 | 0.1698 |  | 0.46385
[8 empty] | 375 |  |  | 270 | [5 empty] | 2.2106 | 2.8606 |  | 2.3119 | 0.70611 |  | 0.078675 |  | 0.63531 |  |  |  | 0.61086 |  | 1.808 | 0.27295 | 0.93824 |  | 0.34965 | [7 empty] | 1.4033 | 0.49417 | 0.17976 |  | 0.46146
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.76655 |  | 0.14932 |  | 0.62821 |  |  |  | 0.58156 |  |  | 0.078969 | 0.81251 |  | 0.2224 | [8 empty] | 0.55113 | 0.20966 |  | 0.48851
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.57277 |  | 0.036388 |  | 0.54173 |  |  |  | 0.75871 |  |  | 0.53918 | 1.0715 |  | 0.57777 | [8 empty] | 0.52464 | 0.26874 |  | 0.50991
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.72928 |  | 0.15247 |  | 0.58762 |  |  |  | 0.7166 |  |  | 0.14695 | 0.83691 |  | 0.29259 | [8 empty] | 0.65762 | 0.33187 |  | 0.59381
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.69608 |  | 0.24889 |  | 0.46139 |  |  |  | 0.59269 |  |  | 0.079075 | 0.74941 |  | 0.31955 | [8 empty] | 0.56172 | 0.25127 |  | 0.45686
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.5979 |  | 1.7771 |  | 2.3125 |  |  |  | 0.57682 |  |  | 0.26455 | 0.8254 |  | 1.4197 | [8 empty] | 0.46357 | 0.19708 |  | 1.2184
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.6897 | [7 empty] | 0.4218 |  |  | 0.015178 | 0.6471 | [10 empty] | 0.43213 | 0.12545
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.73568 | [10 empty] | 0.013317 | 0.70804
[8 empty] | 495 |  |  | 390 | [9 empty] | 0.69888 | [10 empty] | 0.14509 | 0.81144
[8 empty] | 510 |  |  | 405 | [9 empty] | 0.7136 | [10 empty] | 0.21619 | 0.45764
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.55694
435
450
45
PTB | 55 | 19.9577 | 20.0349 | 18.8162 | 0.076267 | -6.0832 | 0.43 | 60 | -5.2547 | 0.43
| 70 | 19.9577 | 20.0426 | 19.3215 | 0.052816 | -3.5977 | 0.43 | 75 | -3.2061 | 0.43
| 85 | 19.9577 | 20.003 | 19.5183 | 0.036316 | -2.4229 | 0.43 | 90 | -2.1987 | 0.43
| 100 | 19.9577 | 19.9899 | 19.64 | 0.056285 | -1.7504 | 0.43 | 105 | -1.4268 | 0.43
| 115 | 19.9577 | 19.9807 | 19.825 | 0.059883 | -0.77949 | 0.43 | 120 | -0.60812 | 0.43
| 130 | 19.9577 | 19.9626 | 19.9096 | 0.054953 | -0.26539 | 0.43 | 135 | -0.2503 | 0.43
| 145 | 19.9577 | 19.9571 | 19.9132 | 0.066063 | -0.22011 | 0.43 | 150 | -0.33729 | 0.43
| 160 | 19.9577 | 19.9577 | 19.8436 | 0.061704 | -0.57166 | 0.43 | 165 | -0.70366 | 0.43
| 175 | 19.9577 | 19.9655 | 19.7723 | 0.075805 | -0.96765 | 0.43 | 180 | -1.1763 | 0.43
| 190 | 19.9577 | 19.969 | 19.6508 | 0.071885 | -1.5937 | 0.43 | 195 | -1.7029 | 0.43
| 205 | 19.9577 | 19.9875 | 19.6035 | 0.069163 | -1.9214 | 0.43 | 210 | -2.0451 | 0.43
| 220 | 19.9577 | 19.9957 | 19.5373 | 0.065456 | -2.2924 | 0.43 | 225 | -2.4324 | 0.43
| 235 | 19.9577 | 20.0185 | 19.4756 | 0.046021 | -2.7124 | 0.43 | 240 | -2.7786 | 0.43
| 250 | 19.9577 | 20.0335 | 19.4504 | 0.04372 | -2.9109 | 0.43 | 255 | -2.9793 | 0.43
| 265 | 19.9577 | 20.0526 | 19.4277 | 0.051638 | -3.1162 | 0.43 | 270 | -3.3313 | 0.43
| 280 | 19.9577 | 20.1048 | 19.3485 | 0.049746 | -3.7616 | 0.43 | 285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 |  | 20.105 | 18.366 | 0.037473 | -8.6494 | 0.5 | 45 | -8.6494 | 0.5
| 60 |  | 20.094 | 18.592 | 0.072228 | -7.4747 | 0.5 | 60 | -7.4747 | 0.5
| 75 |  | 20.087 | 19.2315 | 0.048696 | -4.2588 | 0.5 | 75 | -4.2588 | 0.5
| 90 |  | 20.0856 | 19.4445 | 0.040194 | -3.1919 | 0.5 | 90 | -3.1919 | 0.5
| 105 |  | 20.0915 | 19.622 | 0.051052 | -2.3369 | 0.5 | 105 | -2.3369 | 0.5
| 120 |  | 20.0873 | 19.8325 | 0.052403 | -1.2685 | 0.5 | 120 | -1.2685 | 0.5
| 135 |  | 20.0784 | 19.98 | 0.045306 | -0.49013 | 0.5 | 135 | -0.49013 | 0.5
| 150 |  | 20.0737 | 20.091 | 0.045872 | 0.086133 | 0.5 | 150 | 0.086133 | 0.5
| 165 |  | 20.0716 | 20.0735 | 0.084808 | 0.0094163 | 0.5 | 165 | 0.0094163 | 0.5
| 180 |  | 20.0761 | 19.969 | 0.045178 | -0.53327 | 0.5 | 180 | -0.53327 | 0.5
| 195 |  | 20.0736 | 19.928 | 0.064856 | -0.72513 | 0.5 | 195 | -0.72513 | 0.5
| 210 |  | 20.0787 | 19.8155 | 0.056052 | -1.3106 | 0.5 | 210 | -1.3106 | 0.5
| 225 | 20.0623 | 20.0623 | 19.753 | 0.048134 | -1.5415 | 0.5 | 225 | -1.5415 | 0.5
| 240 |  | 20.0771 | 19.7345 | 0.036052 | -1.7062 | 0.5 | 240 | -1.7062 | 0.5
| 255 |  | 20.0776 | 19.7575 | 0.038916 | -1.5941 | 0.5 | 255 | -1.5941 | 0.5
| 270 |  | 20.0624 | 19.738 | 0.055687 | -1.6168 | 0.5 | 270 | -1.6168 | 0.5
| 285 |  | 20.0808 | 19.743 | 0.042809 | -1.682 | 0.5 | 285 | -1.682 | 0.5
| 300 |  | 20.0788 | 19.757 | 0.05545 | -1.6025 | 0.5 | 300 | -1.6025 | 0.5
| 315 |  | 20.0877 | 19.733 | 0.028303 | -1.7656 | 0.5 | 315 | -1.7656 | 0.5
| 330 |  | 20.0814 | 19.731 | 0.050147 | -1.7447 | 0.5 | 330 | -1.7447 | 0.5
| 345 |  | 20.0743 | 19.7515 | 0.046597 | -1.6078 | 0.5 | 345 | -1.6078 | 0.5
| 360 |  | 20.0797 | 19.7325 | 0.05609 | -1.7289 | 0.5 | 360 | -1.7289 | 0.5
| 375 |  | 20.0962 | 19.732 | 0.046971 | -1.8121 | 0.5 | 375 | -1.8121 | 0.5
| 390 |  | 20.0596 | 19.708 | 0.045143 | -1.7526 | 0.5 | 390 | -1.7526 | 0.5
| 405 |  | 20.0634 | 19.6985 | 0.049019 | -1.8185 | 0.5 | 405 | -1.8185 | 0.5
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 19.9926 | 19.47 | 0.1 | -2.6142 | 0.56777 | 120 | -2.6142 | 0.56777
| 135 |  | 19.9957 | 19.655 | 0.1 | -1.7037 | 0.56255 | 135 | -1.7037 | 0.56255
| 150 |  | 19.9799 | 19.795 | 0.14 | -0.92533 | 0.58456 | 150 | -0.92533 | 0.58456
| 165 |  | 19.9946 | 19.925 | 0.035 | -0.34825 | 0.55814 | 165 | -0.34825 | 0.55814
| 180 |  | 19.9931 | 20.05 | 0.13 | 0.28474 | 0.5868 | 180 | 0.28474 | 0.5868
| 195 |  | 19.997 | 19.885 | 0.13 | -0.56014 | 0.5855 | 195 | -0.56014 | 0.5855
| 210 |  | 20.0115 | 19.835 | 0.105 | -0.88206 | 0.56838 | 210 | -0.88206 | 0.56838
| 225 |  | 20.0151 | 19.74 | 0.105 | -1.3746 | 0.56673 | 225 | -1.3746 | 0.56673
| 240 |  | 20.0281 | 19.685 | 0.135 | -1.7131 | 0.58094 | 240 | -1.7131 | 0.58094
| 255 |  | 20.0034 | 19.655 | 0.12 | -1.7418 | 0.57456 | 255 | -1.7418 | 0.57456
| 270 | 20.0054 | 20.0104 | 19.64 | 0.08 | -1.8509 | 0.56472 | 270 | -1.8509 | 0.56472
| 285 |  | 19.9948 | 19.645 | 0.085 | -1.7493 | 0.55717 | 285 | -1.7493 | 0.55717
| 300 |  | 19.982 | 19.57 | 0.08 | -2.0616 | 0.55541 | 300 | -2.0616 | 0.55541
| 315 |  | 19.9831 | 19.605 | 0.105 | -1.8919 | 0.56825 | 315 | -1.8919 | 0.56825
| 330 |  | 19.982 | 19.62 | 0.04 | -1.8116 | 0.54771 | 330 | -1.8116 | 0.54771
| 345 |  | 19.997 | 19.63 | 0.09 | -1.8353 | 0.56823 | 345 | -1.8353 | 0.56823
| 360 |  | 19.9931 | 19.65 | 0.085 | -1.716 | 0.55853 | 360 | -1.716 | 0.55853
| 375 |  | 19.9946 | 19.63 | 0.12 | -1.8236 | 0.57986 | 375 | -1.8236 | 0.57986
| 390 |  | 19.9799 | 19.605 | 0.07 | -1.8763 | 0.55705 | 390 | -1.8763 | 0.55705
| 405 |  | 19.9957 | 19.67 | 0.15 | -1.6287 | 0.59521 | 405 | -1.6287 | 0.59521
| 420 |  | 19.9926 | 19.65 | 0.1 | -1.7139 | 0.56368 | 420 | -1.7139 | 0.56368
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 |  | 19.7223 | 18.0026 | 0.29378 | -8.7197 | 0.55986 | 60 | -6.9563 | 0.55278
| 65 |  | 20.0643 | 18.8455 | 0.22832 | -6.0746 | 0.54925 | 75 | -4.433 | 0.54843
| 80 |  | 20.0853 | 19.3598 | 0.13353 | -3.6123 | 0.54802 | 90 | -2.7115 | 0.54839
| 95 |  | 20.0573 | 19.6038 | 0.092615 | -2.2612 | 0.54857 | 105 | -1.6493 | 0.54893
| 110 |  | 20.0343 | 19.7652 | 0.063677 | -1.3434 | 0.54911 | 120 | -1.0785 | 0.54931
| 125 |  | 20.0223 | 19.8329 | 0.048977 | -0.94611 | 0.54941 | 135 | -0.62869 | 0.54927
| 140 |  | 20.0293 | 19.9352 | 0.060266 | -0.46998 | 0.5492 | 150 | -0.29926 | 0.54927
| 155 |  | 20.0253 | 19.9825 | 0.069996 | -0.2139 | 0.54931 | 165 | -0.13668 | 0.54941
| 170 |  | 20.0193 | 19.9997 | 0.064656 | -0.098072 | 0.54947 | 180 | -0.14669 | 0.54947
| 185 |  | 20.0193 | 19.9851 | 0.053751 | -0.171 | 0.54947 | 195 | -0.32818 | 0.54947
| 200 |  | 20.0193 | 19.9379 | 0.062374 | -0.40677 | 0.54947 | 210 | -0.43055 | 0.54964
| 215 |  | 20.0103 | 19.9218 | 0.067167 | -0.44244 | 0.54972 | 225 | -0.60138 | 0.54974
| 230 |  | 20.0093 | 19.8731 | 0.065248 | -0.68085 | 0.54976 | 240 | -0.80091 | 0.54983
| 245 |  | 20.0053 | 19.8331 | 0.05686 | -0.86094 | 0.54987 | 255 | -0.94821 | 0.54986
| 260 |  | 20.0063 | 19.8079 | 0.051887 | -0.99185 | 0.54985 | 270 | -1.018 | 0.54993
| 275 |  | 20.0023 | 19.7961 | 0.038945 | -1.031 | 0.54997 | 285 | -1.0319 | 0.55009
| 290 |  | 19.9954 | 19.789 | 0.030165 | -1.0323 | 0.55016 | 300 | -1.1168 | 0.5499
| 305 |  | 20.0094 | 19.7775 | 0.03647 | -1.159 | 0.54978 | 315 | -1.1472 | 0.54984
| 320 |  | 20.0059 | 19.7776 | 0.030345 | -1.1412 | 0.54987 | 330 | -1.1545 | 0.54983
| 335 |  | 20.0084 | 19.7761 | 0.037541 | -1.1611 | 0.54981 | 345 | -1.1009 | 0.54969
| 350 |  | 20.0144 | 19.8001 | 0.029908 | -1.0708 | 0.54964 | 360 | -1.1442 | 0.54968
| 365 |  | 20.0124 | 19.7761 | 0.040128 | -1.1809 | 0.5497 | 375 | -1.1883 | 0.54976
| 380 |  | 20.0089 | 19.7704 | 0.041458 | -1.1921 | 0.54979 | 390 | -1.16 | 0.54976
| 395 |  | 20.0104 | 19.7815 | 0.037565 | -1.144 | 0.54975 | 405 | -1.2135 | 0.54971
| 410 |  | 20.0129 | 19.7631 | 0.02638 | -1.2483 | 0.54969 | 420 | -1.2199 | 0.54967
| 425 |  | 20.0139 | 19.7726 | 0.032917 | -1.2057 | 0.54966 | 435 | -1.2382 | 0.54972
| 440 |  | 20.0104 | 19.7594 | 0.030052 | -1.2544 | 0.54976 | 450 | -1.2013 | 0.54984
| 455 |  | 20.0054 | 19.7704 | 0.032397 | -1.1748 | 0.54989 | 465 | -1.2405 | 0.5498
| 470 |  | 20.0104 | 19.7556 | 0.039911 | -1.2734 | 0.54976 | 480 | -1.2711 | 0.54978
| 485 |  | 20.0094 | 19.7553 | 0.03038 | -1.27 | 0.54979 | 495 | -1.2077 | 0.54995
| 500 | 20 | 20 | 19.7647 | 0.02717 | -1.1765 | 0.55004 | 510 | -1.2019 | 0.54991
| 515 |  | 20.0069 | 19.7639 | 0.032669 | -1.2147 | 0.54985 | 525 | -1.2293 | 0.54985
| 530 |  | 20.0069 | 19.7595 | 0.028461 | -1.2367 | 0.54985
UL-LMPS | 42.5 | 20.3249 | 20.2699 | 18.2983 | 0.04128 | -9.7267 | 0.65 | 45 | -9.6809 | 0.65015
| 57.5 | 20.2059 | 20.1079 | 18.2073 | 0.05343 | -9.4518 | 0.65088 | 60 | -8.8282 | 0.65118
| 72.5 | 20.068 | 19.9481 | 18.809 | 0.059617 | -5.7101 | 0.65267 | 75 | -5.524 | 0.65222
| 87.5 | 19.9509 | 19.8274 | 18.9167 | 0.046708 | -4.5931 | 0.65 | 90 | -4.4221 | 0.65071
| 102.5 | 19.8615 | 19.7429 | 19.0387 | 0.06458 | -3.567 | 0.65423 | 105 | -3.3707 | 0.66076
| 117.5 | 19.7484 | 19.6785 | 19.2083 | 0.055776 | -2.3891 | 0.69338 | 120 | -2.3276 | 0.69316
| 132.5 | 19.6012 | 19.5659 | 19.1707 | 0.080298 | -2.0202 | 0.69205 | 135 | -1.854 | 0.79105
| 147.5 | 19.4696 | 19.4072 | 19.2087 | 0.063503 | -1.0232 | 1.286 | 150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
| 50 |  | 20.4428 | 19.5565 | 0.10189 | -4.3356 | 1.8 | 60 | -4.0462 | 1.8
SHMU | 65 |  | 20.4452 | 19.6475 | 0.14653 | -3.9015 | 1.8 | 75 | -2.2797 | 1.8
| 80 |  | 20.4641 | 20.1635 | 0.13406 | -1.4689 | 1.8 | 90 | -0.68617 | 1.8
| 95 |  | 20.4588 | 20.3985 | 0.11613 | -0.29481 | 1.8 | 105 | 0.054867 | 1.8
| 110 |  | 20.4336 | 20.4805 | 0.098967 | 0.22971 | 1.8 | 120 | 0.30941 | 1.8
| 125 |  | 20.4037 | 20.475 | 0.094896 | 0.34926 | 1.8 | 135 | 0.19029 | 1.8
| 140 |  | 20.3914 | 20.414 | 0.083628 | 0.1108 | 1.8 | 150 | 0.17312 | 1.8
| 155 |  | 20.3719 | 20.4135 | 0.068616 | 0.20428 | 1.8 | 165 | -0.092619 | 1.8
| 170 |  | 20.3636 | 20.3145 | 0.081594 | -0.24107 | 1.8 | 180 | -0.34852 | 1.8
| 185 |  | 20.3544 | 20.2725 | 0.076425 | -0.40225 | 1.8 | 195 | -0.44232 | 1.8
| 200 |  | 20.3667 | 20.2725 | 0.099941 | -0.46236 | 1.8 | 210 | -0.44425 | 1.8
| 215 |  | 20.3752 | 20.2865 | 0.074712 | -0.43519 | 1.8 | 225 | -0.54354 | 1.8
| 230 |  | 20.4185 | 20.2965 | 0.074993 | -0.59772 | 1.8 | 240 | -0.54479 | 1.8
| 245 |  | 20.4324 | 20.3265 | 0.094606 | -0.51833 | 1.8 | 255 | -0.62121 | 1.8
| 260 |  | 20.4475 | 20.31 | 0.088437 | -0.67265 | 1.8 | 270 | -0.67178 | 1.8
| 275 |  | 20.4629 | 20.3255 | 0.061428 | -0.67134 | 1.8 | 285 | -0.62811 | 1.8
| 290 |  | 20.4359 | 20.312 | 0.092884 | -0.60649 | 1.8 | 300 | -0.59924 | 1.8
| 305 |  | 20.4136 | 20.292 | 0.094457 | -0.59562 | 1.8 | 315 | -0.50803 | 1.8
| 320 | 20.3721 | 20.3721 | 20.2775 | 0.085954 | -0.46423 | 1.8 | 330 | -0.67056 | 1.8
| 335 |  | 20.3172 | 20.16 | 0.092963 | -0.77372 | 1.8 | 345 | -0.72137 | 1.8
| 350 |  | 20.2724 | 20.1315 | 0.086406 | -0.69519 | 1.8 | 360 | -0.90257 | 1.8
| 365 |  | 20.3326 | 20.128 | 0.084766 | -1.0063 | 1.8 | 375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 20.212 | 18.521 | 0.059 | -8.366 | 0.49 | 45 | -8.366 | 0.49
| 60 |  | 20.211 | 19.121 | 0.058 | -5.393 | 0.49 | 60 | -5.393 | 0.49
LEI | 75 |  | 20.209 | 19.424 | 0.049 | -3.884 | 0.49 | 75 | -3.884 | 0.49
| 90 |  | 20.218 | 19.63 | 0.044 | -2.908 | 0.49 | 90 | -2.908 | 0.49
| 105 |  | 20.208 | 19.929 | 0.049 | -1.381 | 0.49 | 105 | -1.381 | 0.49
| 120 |  | 20.204 | 20.075 | 0.053 | -0.638 | 0.49 | 120 | -0.638 | 0.49
| 135 |  | 20.218 | 20.242 | 0.046 | 0.119 | 0.49 | 135 | 0.119 | 0.49
| 150 |  | 20.192 | 20.282 | 0.085 | 0.446 | 0.49 | 150 | 0.446 | 0.49
| 165 |  | 20.189 | 20.207 | 0.065 | 0.089 | 0.49 | 165 | 0.089 | 0.49
| 180 |  | 20.189 | 20.117 | 0.1 | -0.357 | 0.49 | 180 | -0.357 | 0.49
| 195 |  | 20.19 | 20 | 0.078 | -0.941 | 0.49 | 195 | -0.941 | 0.49
| 210 |  | 20.188 | 19.937 | 0.062 | -1.243 | 0.49 | 210 | -1.243 | 0.49
| 225 |  | 20.189 | 19.885 | 0.051 | -1.506 | 0.49 | 225 | -1.506 | 0.49
| 240 |  | 20.21 | 19.894 | 0.063 | -1.564 | 0.49 | 240 | -1.564 | 0.49
| 255 |  | 20.169 | 19.865 | 0.07 | -1.507 | 0.49 | 255 | -1.507 | 0.49
| 270 |  | 20.159 | 19.846 | 0.048 | -1.553 | 0.49 | 270 | -1.553 | 0.49
| 285 |  | 20.175 | 19.86 | 0.041 | -1.561 | 0.49 | 285 | -1.561 | 0.49
| 300 |  | 20.155 | 19.838 | 0.057 | -1.573 | 0.49 | 300 | -1.573 | 0.49
| 315 |  | 20.16 | 19.829 | 0.04 | -1.642 | 0.49 | 315 | -1.642 | 0.49
| 330 |  | 20.156 | 19.845 | 0.082 | -1.543 | 0.49 | 330 | -1.543 | 0.49
| 345 |  | 20.142 | 19.528 | 0.102 | -3.048 | 0.49 | 345 | -3.048 | 0.49
| 360 |  | 20.146 | 18.554 | 0.162 | -7.902 | 0.49 | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-L 0.5
| 0.5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 4.0825
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 0.49994 | 0.47 | 1.1703e-16 | -5.9888 | 1.8002 | 45
| 52.5 |  | 0.49921 | 0.47 | 1.1703e-16 | -5.8507 | 1.8029 | 60 |  |  | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 0.49916 | 0.479 | 0.0031623 | -4.0381 | 1.803 | 75 | -3.8952 | 1.8038 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 0.49871 | 0.48 | 1.1703e-16 | -3.7523 | 1.8046 | 90 | -3.7844 | 1.804 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 0.49905 | 0.48 | 1.1703e-16 | -3.8166 | 1.8034 | 105 | -2.8354 | 1.8031 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 0.49926 | 0.49 | 1.1703e-16 | -1.8542 | 1.8027 | 120 | -1.8806 | 1.8022 |  | 0.23195 | 0.78908 |  |  |  | 0.61425 | [12 empty] | 0.88464 | 1.1644 | 1.6293
| 127.5 | 0.49964 | 0.49953 | 0.49 | 1.1703e-16 | -1.9069 | 1.8017 | 135 | -1.7149 | 1.8052 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 0.49758 | 0.49 | 1.1703e-16 | -1.5229 | 1.8088 | 150 | -1.6113 | 1.8071 | 45
| 157.5 |  | 0.49847 | 0.49 | 1.1703e-16 | -1.6997 | 1.8055 | 165 | -1.7705 | 1.8061 | 60
| 172.5 |  | 0.49817 | 0.489 | 0.0031623 | -1.8414 | 1.8066 | 180 | -2.7896 | 1.8039 | 75 | 0.28938 | 0.67985 |  |  |  | 0.44223
| 187.5 |  | 0.49968 | 0.481 | 0.0031623 | -3.7379 | 1.8012 | 195 |  |  | 90 | 0.27584 | 0.65765 |  |  |  | 0.4317
| 202.5 |  | 0.5002 | 0.48 | 1.1703e-16 | -4.0379 | 1.7993 | 210 |  |  | 105 | 0.28631 | 0.84205 |  |  |  | 0.61716
| 217.5 |  | 0.50125 | 0.47 | 1.1703e-16 | -6.2337 | 1.7955 | 225 |  |  | 120 | 0.33684 | 0.86007 |  |  |  | 0.58752 | [12 empty] | 0.34298 | 0.69233 | 1.1747
[8 empty] | 240 |  |  | 135 | 0.36811 | 0.92679 |  |  |  | 0.62834 | [12 empty] | 0.30707 | 0.68476 | 1.2052
[8 empty] | 255 |  |  | 150 | 0.37896 | 1.0329 |  |  |  | 0.7307 | [12 empty] | 0.27189 | 0.65812 | 1.2749
[8 empty] | 270 |  |  | 165 | 0.2812 | 0.72676 |  |  |  | 0.50031 | [12 empty] | 0.37383 | 0.67108 | 1.0742
[8 empty] | 285 |  |  | 180 | 0.076974 | 0.77401 |  |  |  | 0.74873 | [12 empty] | 0.5986 | 0.71163 | 1.3392
[8 empty] | 300 |  |  | 195 | [5 empty] | 0.42587 | [13 empty] | 0.79867 | 1.1214
[8 empty] | 315 |  |  | 210 | [5 empty] | 0.39552 | [13 empty] | 1.481 | 1.7207
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.24621 | [13 empty] | 1.4588 | 1.5645
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.22524 | [13 empty] | 1.4208 | 1.5102
[8 empty] | 360 |  |  | 255 | [20 empty] | 1.5239
[8 empty] | 375 |  |  | 270 | [20 empty] | 1.5339
[8 empty] | 390 |  |  | 285 | [20 empty] | 1.5423
[8 empty] | 405 |  |  | 300 | [20 empty] | 1.521
[8 empty] | 420 |  |  | 315 | [20 empty] | 1.5546
[8 empty] | 435 |  |  | 330 | [20 empty] | 1.5867
[8 empty] | 450 |  |  | 345 | [20 empty] | 1.5255
[8 empty] | 465 |  |  | 360 | [20 empty] | 1.536
[8 empty] | 480 |  |  | 375 | [20 empty] | 1.421
[8 empty] | 495 |  |  | 390
[8 empty] | 510 |  |  | 405
[8 empty] | 525 |  |  | 420
435
450
45
PTB | 55 | 0.49419 | 0.49698 | 0.47 | 1.1331e-16 | -5.4286 | 0.64 | 60
| 70 | 0.49419 | 0.49682 | 0.47125 | 0.0033783 | -5.1461 | 0.65 | 75 | -5.2003 | 0.64667
| 85 | 0.49419 | 0.49679 | 0.47042 | 0.0020412 | -5.3086 | 0.64 | 90 | -5.0286 | 0.65
| 100 | 0.49419 | 0.49678 | 0.47458 | 0.0050898 | -4.4687 | 0.67 | 105 | -4.1273 | 0.66667
| 115 | 0.49419 | 0.49464 | 0.4776 | 0.0043589 | -3.4446 | 0.66 | 120 | -3.3999 | 0.65667
| 130 | 0.49419 | 0.49478 | 0.4784 | 0.0037417 | -3.3105 | 0.65 | 135 | -3.3755 | 0.65333
| 145 | 0.49419 | 0.49495 | 0.4776 | 0.0043589 | -3.5055 | 0.66 | 150 | -3.3212 | 0.65333
| 160 | 0.49419 | 0.49419 | 0.4796 | 0.002 | -2.9525 | 0.64 | 165 | -3.0387 | 0.64
| 175 | 0.49419 | 0.49592 | 0.48 | 2.8328e-16 | -3.2111 | 0.64 | 180 | -3.1367 | 0.64
| 190 | 0.49419 | 0.49478 | 0.48 | 2.8379e-16 | -2.988 | 0.64 | 195 | -3.4749 | 0.65
| 205 | 0.49419 | 0.49643 | 0.47435 | 0.0050687 | -4.4487 | 0.67 | 210 | -4.5307 | 0.66667
| 220 | 0.49419 | 0.49578 | 0.4725 | 0.0044233 | -4.6949 | 0.66 | 225 | -4.6614 | 0.66333
| 235 | 0.49419 | 0.49656 | 0.47375 | 0.0049454 | -4.5944 | 0.67 | 240 | -4.2469 | 0.66
| 250 | 0.49419 | 0.49768 | 0.48 | 2.8353e-16 | -3.5519 | 0.64 | 255
| 265 | 0.49419 | 0.49791 | 0.48 | 2.8353e-16 | -3.5965 | 0.64 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 | [6 empty] | 45
| 60 |  | 0.55122 | 0.501 | 0.0030779 | -9.1102 | 2.1219 | 60
| 75 |  | 0.54882 | 0.509 | 0.0030779 | -7.255 | 2.1219 | 75 | -7.255 | 2.1219
| 90 |  | 0.54967 | 0.511 | 0.0030779 | -7.0346 | 2.1219 | 90 | -7.0346 | 2.1219
| 105 |  | 0.54837 | 0.51 | 1.1391e-16 | -6.9965 | 2.1219 | 105 | -6.9965 | 2.1219
| 120 |  | 0.54917 | 0.5155 | 0.0051042 | -6.1305 | 2.1219 | 120 | -6.1305 | 2.1219
| 135 |  | 0.54907 | 0.5145 | 0.0051042 | -6.2955 | 2.1219 | 135 | -6.2955 | 2.1219
| 150 |  | 0.54887 | 0.512 | 0.0041039 | -6.7169 | 2.1219 | 150 | -6.7169 | 2.1219
| 165 |  | 0.54947 | 0.52 | 2.2781e-16 | -5.3628 | 2.1219 | 165 | -5.3628 | 2.1219
| 180 |  | 0.54827 | 0.512 | 0.0041039 | -6.6148 | 2.1219 | 180 | -6.6148 | 2.1219
| 195 |  | 0.54947 | 0.5195 | 0.0022361 | -5.4538 | 2.1219 | 195 | -5.4538 | 2.1219
| 210 |  | 0.54897 | 0.514 | 0.0050262 | -6.3695 | 2.1219 | 210 | -6.3695 | 2.1219
| 225 | 0.54887 | 0.54887 | 0.517 | 0.0047016 | -5.8059 | 2.1219 | 225 | -5.8059 | 2.1219
| 240 |  | 0.54907 | 0.52 | 2.2781e-16 | -5.2938 | 2.1219 | 240 | -5.2938 | 2.1219
| 255 |  | 0.54887 | 0.52 | 2.2781e-16 | -5.2593 | 2.1219 | 255 | -5.2593 | 2.1219
| 270 |  | 0.54817 | 0.519 | 0.0030779 | -5.3208 | 2.1219 | 270 | -5.3208 | 2.1219
| 285 |  | 0.54817 | 0.519 | 0.0030779 | -5.3208 | 2.1219 | 285 | -5.3208 | 2.1219
| 300 |  | 0.54827 | 0.5195 | 0.0022361 | -5.2468 | 2.1219 | 300 | -5.2468 | 2.1219
| 315 |  | 0.54797 | 0.519 | 0.0030779 | -5.2862 | 2.1219 | 315 | -5.2862 | 2.1219
| 330 |  | 0.54927 | 0.519 | 0.0030779 | -5.5104 | 2.1219 | 330 | -5.5104 | 2.1219
| 345 |  | 0.54807 | 0.518 | 0.0052315 | -5.486 | 2.1219 | 345 | -5.486 | 2.1219
| 360 |  | 0.54947 | 0.5185 | 0.0036635 | -5.6358 | 2.1219 | 360 | -5.6358 | 2.1219
| 375 |  | 0.54872 | 0.5205 | 0.0022361 | -5.1423 | 2.1219 | 375 | -5.1423 | 2.1219
| 390 |  | 0.55052 | 0.521 | 0.0030779 | -5.3616 | 2.1219 | 390
| 405 | [6 empty] | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 0.50095 | 0.5 | 0 | -0.18884 | 2.101 | 120 | -0.18884 | 2.101
| 135 |  | 0.501 | 0.5 | 0 | -0.2 | 2.1005 | 135 | -0.2 | 2.1005
| 150 |  | 0.50135 | 0.5 | 0 | -0.27007 | 2.0975 | 150 | -0.27007 | 2.0975
| 165 |  | 0.49962 | 0.5 | 0 | 0.075659 | 2.1121 | 165 | 0.075659 | 2.1121
| 180 |  | 0.49917 | 0.5 | 0 | 0.16708 | 2.116 | 180 | 0.16708 | 2.116
| 195 |  | 0.49868 | 0.5 | 0.01 | 0.2643 | 2.1977 | 195 | 0.2643 | 2.1977
| 210 |  | 0.49785 | 0.51 | 0.01 | 2.4401 | 2.2458 | 210 | 2.4401 | 2.2458
| 225 |  | 0.49904 | 0.51 | 0.01 | 2.197 | 2.2355 | 225 | 2.197 | 2.2355
| 240 |  | 0.49811 | 0.51 | 0 | 2.387 | 2.1674 | 240 | 2.387 | 2.1674
| 255 |  | 0.49759 | 0.51 | 0 | 2.4942 | 2.172 | 255 | 2.4942 | 2.172
| 270 | 0.49761 | 0.49764 | 0.51 | 0 | 2.4833 | 2.1715 | 270 | 2.4833 | 2.1715
| 285 |  | 0.49743 | 0.51 | 0 | 2.527 | 2.1734 | 285 | 2.527 | 2.1734
| 300 |  | 0.4976 | 0.51 | 0 | 2.492 | 2.1719 | 300 | 2.492 | 2.1719
| 315 |  | 0.49694 | 0.51 | 0 | 2.6271 | 2.1776 | 315 | 2.6271 | 2.1776
| 330 |  | 0.49725 | 0.51 | 0 | 2.5647 | 2.175 | 330 | 2.5647 | 2.175
| 345 |  | 0.49868 | 0.51 | 0 | 2.2696 | 2.1625 | 345 | 2.2696 | 2.1625
| 360 |  | 0.49917 | 0.51 | 0 | 2.1704 | 2.1583 | 360 | 2.1704 | 2.1583
| 375 |  | 0.49962 | 0.51 | 0 | 2.0772 | 2.1543 | 375 | 2.0772 | 2.1543
| 390 |  | 0.50135 | 0.51 | 0 | 1.7245 | 2.1395 | 390 | 1.7245 | 2.1395
| 405 |  | 0.501 | 0.51 | 0 | 1.796 | 2.1425 | 405 | 1.796 | 2.1425
| 420 |  | 0.50095 | 0.51 | 0 | 1.8074 | 2.143 | 420 | 1.8074 | 2.143
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
LEI | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-L 2
| 2 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.58788
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 1.9344 | 1.819 | 0.0031623 | -5.9657 | 0.82713 | 45
| 52.5 |  | 1.9329 | 1.851 | 0.0031623 | -4.2352 | 0.82779 | 60 |  |  | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 1.9372 | 1.858 | 0.0042164 | -4.0907 | 0.82591 | 75 | -4.0163 | 0.82567 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 1.9384 | 1.862 | 0.0042164 | -3.9419 | 0.82542 | 90 | -3.7462 | 0.82533 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 1.9388 | 1.87 | 4.6811e-16 | -3.5504 | 0.82524 | 105 | -3.3043 | 0.82492 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 1.9403 | 1.881 | 0.0031623 | -3.0581 | 0.8246 | 120 | -2.7824 | 0.82475 |  | 1.1951 | 0.065369 | 0.44808 |  | 0.61674 | 1.2675 | 0.99513 |  | 1.5083 | 0.4145 |  | 0.56615 |  | 0.031597 |  |  |  |  | 0.5603 | 1.7992 | 0.49542 | 0.15306 |  | 0.17539
| 127.5 | 1.9403 | 1.9396 | 1.891 | 0.0031623 | -2.5067 | 0.8249 | 135 | -2.4343 | 0.82464 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 1.9408 | 1.895 | 0.0052705 | -2.3618 | 0.82439 | 150 | -2.3852 | 0.82484 | 45
| 157.5 |  | 1.9387 | 1.892 | 0.0042164 | -2.4086 | 0.8253 | 165 | -2.4215 | 0.82519 | 60
| 172.5 |  | 1.9392 | 1.892 | 0.0042164 | -2.4345 | 0.82508 | 180 | -2.3869 | 0.82504 | 75 | 1.1296 | 0.20617 |  |  | 0.17456 | 1.3443 |  |  | 0.59518 |  |  | 0.32988
| 187.5 |  | 1.9394 | 1.894 | 0.005164 | -2.3394 | 0.82501 | 195 |  |  | 90 | 1.1448 | 0.053264 | 0.17574 |  | 0.10135 | 1.0666 | 0.3253 |  | 0.90825 | 0.14805 |  | 0.14124 |  | 0.22322
| 202.5 |  | 1.9382 | 1.897 | 0.0048305 | -2.1255 | 0.82551 | 210 |  |  | 105 | 1.1317 | 0.23549 | 0.041275 |  | 0.24554 | 0.84093 | 0.45924 |  | 1.057 | 0.079059 |  | 0.42311 |  | 0.18883
| 217.5 |  | 1.9394 | 1.86 | 2.3406e-16 | -4.095 | 0.82499 | 225 |  |  | 120 | 1.3319 | 0.39689 | 0.054848 |  | 0.23175 | 0.85347 | 0.53438 |  | 1.1811 | 0.14748 |  | 0.53174 |  | 0.19285 |  |  |  |  | 1.1439 | 0.058941 | 0.75188 | 0.53645 |  | 1.1015
[8 empty] | 240 |  |  | 135 | 1.4203 | 0.30419 | 0.04859 |  | 0.25943 | 1.0512 | 0.68052 |  | 1.273 | 0.20344 |  | 0.48916 |  | 0.11507 |  |  |  |  | 0.92208 | 0.27523 | 0.62392 | 0.52724 |  | 0.96091
[8 empty] | 255 |  |  | 150 | 1.3284 | 0.30888 | 0.27168 |  | 0.34502 | 0.95354 | 0.87058 |  | 1.303 | 0.42843 |  | 0.57813 |  | 0.036934 |  |  |  |  | 0.60141 | 0.56154 | 0.3002 | 0.58166 |  | 0.79947
[8 empty] | 270 |  |  | 165 | 1.2664 | 0.22324 | 0.34837 |  | 0.39552 | 0.99254 | 0.92284 |  | 1.3155 | 0.46149 |  | 0.5636 |  | 0.076233 |  |  |  |  | 0.11673 | 1.0672 | 0.099216 | 0.40583 |  | 0.47696
[8 empty] | 285 |  |  | 180 | 1.22 | 0.096927 | 0.35839 |  | 0.54738 | 1.0962 | 0.91274 |  | 1.4497 | 0.40724 |  | 0.61927 |  | 0.0078532 |  |  |  |  | 0.23089 | 0.8905 | 0.1362 | 0.47518 |  | 0.71354
[8 empty] | 300 |  |  | 195 | [5 empty] | 1.1378 | 0.85334 |  | 1.3038 | 0.33419 |  | 0.46334 |  | 0.021301 | [5 empty] | 1.1985 | 0.097756 | 0.28004 |  | 0.37874
[8 empty] | 315 |  |  | 210 | [5 empty] | 0.8581 | 0.83274 |  | 1.2499 | 0.4331 |  | 0.5907 |  | 0.033614 | [5 empty] | 1.0107 | 0.15637 | 0.34864 |  | 0.46125
[8 empty] | 330 |  |  | 225 | [5 empty] | 0.75918 | 0.7744 |  | 0.93262 | 0.41954 |  | 0.3657 |  | 0.15949 | [5 empty] | 0.8298 | 0.089169 | 0.36931 |  | 0.29007
[8 empty] | 345 |  |  | 240 | [5 empty] | 0.9035 | 0.80547 |  | 0.96913 | 0.38858 |  | 0.30831 |  | 0.16647 | [5 empty] | 1.2947 | 0.37197 | 0.19032 |  | 0.015319
[8 empty] | 360 |  |  | 255 | [9 empty] | 0.44547 |  | 0.27817 |  | 0.23718 | [6 empty] | 0.51505 | 0.17324 |  | 0.12317
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.44002 |  | 0.049024 |  | 0.37336 | [6 empty] | 0.45586 | 0.19884 |  | 0.30312
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.42245 |  | 0.071952 |  | 0.34315 | [6 empty] | 0.39759 | 0.21021 |  | 0.23678
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.48831 |  | 0.10447 |  | 0.38332 | [6 empty] | 0.69263 | 0.12026 |  | 0.43258
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.5325 |  | 0.27193 |  | 0.65569 | [6 empty] | 1.1634 | 0.084555 |  | 1.1655
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.54583 | [10 empty] | 0.90769 | 0.062676
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.52476 | [10 empty] | 0.74451 | 0.12852
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.45933 | [10 empty] | 0.61224 | 0.13221
[8 empty] | 480 |  |  | 375 | [9 empty] | 0.51876 | [10 empty] | 0.80011 | 0.089218
[8 empty] | 495 |  |  | 390 | [21 empty] | 0.057759
[8 empty] | 510 |  |  | 405 | [21 empty] | 0.075048
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.055965
435
450
45
PTB | 55 | 2.0025 | 2.0091 | 1.892 | 0.0040825 | -5.8296 | 0.42 | 60
| 70 | 2.0025 | 2.0081 | 1.9018 | 0.0047559 | -5.2948 | 0.42 | 75 | -5.2557 | 0.42
| 85 | 2.0025 | 2.007 | 1.9031 | 0.0047082 | -5.1774 | 0.42 | 90 | -5.0019 | 0.42
| 100 | 2.0025 | 2.0042 | 1.911 | 0.0040925 | -4.6509 | 0.42 | 105 | -4.5453 | 0.42
| 115 | 2.0025 | 2.0047 | 1.9179 | 0.0049868 | -4.3341 | 0.42 | 120 | -4.2428 | 0.42
| 130 | 2.0025 | 2.0035 | 1.9221 | 0.0041786 | -4.0602 | 0.42 | 135 | -3.9915 | 0.42
| 145 | 2.0025 | 2.0024 | 1.9252 | 0.0057981 | -3.854 | 0.42 | 150 | -3.8418 | 0.42
| 160 | 2.0025 | 2.0025 | 1.9261 | 0.0049735 | -3.8172 | 0.42 | 165 | -3.8105 | 0.42
| 175 | 2.0025 | 2.0035 | 1.9274 | 0.0044658 | -3.797 | 0.42 | 180 | -3.7248 | 0.42
| 190 | 2.0025 | 2.0024 | 1.9307 | 0.0054954 | -3.5805 | 0.42 | 195 | -3.545 | 0.42
| 205 | 2.0025 | 2.003 | 1.9335 | 0.0048516 | -3.4739 | 0.42 | 210 | -3.4101 | 0.42
| 220 | 2.0025 | 2.0035 | 1.9377 | 0.0042967 | -3.2826 | 0.42 | 225 | -3.2573 | 0.42
| 235 | 2.0025 | 2.0035 | 1.9392 | 0.0039223 | -3.2066 | 0.42 | 240 | -3.2178 | 0.42
| 250 | 2.0025 | 2.0062 | 1.9412 | 0.0043146 | -3.2402 | 0.42 | 255
| 265 | 2.0025 | 2.009 | 1.9381 | 0.0040192 | -3.5283 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 | [6 empty] | 45
| 60 |  | 2.0101 | 1.921 | 0.0055251 | -4.4334 | 0.91025 | 60
| 75 |  | 2.0102 | 1.935 | 0.0051299 | -3.7417 | 0.86388 | 75 | -3.7417 | 0.86388
| 90 |  | 2.0154 | 1.9385 | 0.0036635 | -3.8164 | 0.84464 | 90 | -3.8164 | 0.84464
| 105 |  | 2.0148 | 1.942 | 0.0041039 | -3.6141 | 0.8392 | 105 | -3.6141 | 0.8392
| 120 |  | 2.0171 | 1.9505 | 0.0022361 | -3.3026 | 0.83162 | 120 | -3.3026 | 0.83162
| 135 |  | 2.0161 | 1.959 | 0.0030779 | -2.833 | 0.83223 | 135 | -2.833 | 0.83223
| 150 |  | 2.0183 | 1.962 | 0.0041039 | -2.7903 | 0.83307 | 150 | -2.7903 | 0.83307
| 165 |  | 2.0193 | 1.9645 | 0.0051042 | -2.7146 | 0.83491 | 165 | -2.7146 | 0.83491
| 180 |  | 2.0172 | 1.9665 | 0.0048936 | -2.5142 | 0.83527 | 180 | -2.5142 | 0.83527
| 195 |  | 2.0157 | 1.9695 | 0.0022361 | -2.2928 | 0.83009 | 195 | -2.2928 | 0.83009
| 210 |  | 2.0177 | 1.968 | 0.0041039 | -2.464 | 0.83288 | 210 | -2.464 | 0.83288
| 225 | 2.0163 | 2.0163 | 1.9675 | 0.0044426 | -2.4211 | 0.83139 | 225 | -2.4211 | 0.83139
| 240 |  | 2.0148 | 1.97 | 4.5563e-16 | -2.2244 | 0.82886 | 240 | -2.2244 | 0.82886
| 255 |  | 2.0168 | 1.969 | 0.0044721 | -2.3709 | 0.83399 | 255 | -2.3709 | 0.83399
| 270 |  | 2.0142 | 1.9665 | 0.0048936 | -2.369 | 0.83676 | 270 | -2.369 | 0.83676
| 285 |  | 2.01 | 1.9635 | 0.0048936 | -2.3118 | 0.83874 | 285 | -2.3118 | 0.83874
| 300 |  | 2.0122 | 1.9615 | 0.0036635 | -2.518 | 0.83463 | 300 | -2.518 | 0.83463
| 315 |  | 2.0125 | 1.959 | 0.0030779 | -2.6568 | 0.83511 | 315 | -2.6568 | 0.83511
| 330 |  | 2.013 | 1.9585 | 0.0048936 | -2.7058 | 0.84308 | 330 | -2.7058 | 0.84308
| 345 |  | 2.013 | 1.959 | 0.0030779 | -2.681 | 0.84289 | 345 | -2.681 | 0.84289
| 360 |  | 2.0107 | 1.9595 | 0.0051042 | -2.5448 | 0.85537 | 360 | -2.5448 | 0.85537
| 375 |  | 2.0173 | 1.963 | 0.0047016 | -2.6925 | 0.87281 | 375 | -2.6925 | 0.87281
| 390 |  | 2.0085 | 1.9605 | 0.0051042 | -2.3907 | 0.93628 | 390
| 405 | [6 empty] | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 2.0065 | 1.92 | 0 | -4.3094 | 0.86948 | 120 | -4.3094 | 0.86948
| 135 |  | 2.0037 | 1.93 | 0.01 | -3.6764 | 0.88822 | 135 | -3.6764 | 0.88822
| 150 |  | 2.0041 | 1.94 | 0.01 | -3.197 | 0.89235 | 150 | -3.197 | 0.89235
| 165 |  | 2.0016 | 1.95 | 0.01 | -2.5796 | 0.89898 | 165 | -2.5796 | 0.89898
| 180 |  | 2.0041 | 1.95 | 0.01 | -2.6993 | 0.8968 | 180 | -2.6993 | 0.8968
| 195 |  | 2.0033 | 1.96 | 0.01 | -2.1599 | 0.90202 | 195 | -2.1599 | 0.90202
| 210 |  | 2.0052 | 1.96 | 0 | -2.2525 | 0.88875 | 210 | -2.2525 | 0.88875
| 225 |  | 2.0061 | 1.96 | 0.01 | -2.2999 | 0.89947 | 225 | -2.2999 | 0.89947
| 240 |  | 2.0044 | 1.97 | 0.01 | -1.7181 | 0.90546 | 240 | -1.7181 | 0.90546
| 255 |  | 2.0035 | 1.97 | 0 | -1.6719 | 0.89477 | 255 | -1.6719 | 0.89477
| 270 | 2.0045 | 2.0051 | 1.97 | 0 | -1.75 | 0.89335 | 270 | -1.75 | 0.89335
| 285 |  | 2.0055 | 1.97 | 0.01 | -1.7685 | 0.90454 | 285 | -1.7685 | 0.90454
| 300 |  | 2.0015 | 1.97 | 0.01 | -1.5717 | 0.90813 | 300 | -1.5717 | 0.90813
| 315 |  | 2.0013 | 1.98 | 0.01 | -1.0632 | 0.91278 | 315 | -1.0632 | 0.91278
| 330 |  | 1.9992 | 1.97 | 0.01 | -1.4596 | 0.91017 | 330 | -1.4596 | 0.91017
| 345 |  | 2.0033 | 1.97 | 0.01 | -1.6607 | 0.9065 | 345 | -1.6607 | 0.9065
| 360 |  | 2.0041 | 1.97 | 0.01 | -1.7014 | 0.90576 | 360 | -1.7014 | 0.90576
| 375 |  | 2.0016 | 1.97 | 0.01 | -1.5805 | 0.90797 | 375 | -1.5805 | 0.90797
| 390 |  | 2.0041 | 1.98 | 0.01 | -1.2011 | 0.91027 | 390 | -1.2011 | 0.91027
| 405 |  | 2.0037 | 1.98 | 0.01 | -1.1809 | 0.91064 | 405 | -1.1809 | 0.91064
| 420 |  | 2.0065 | 1.98 | 0 | -1.3191 | 0.89666 | 420 | -1.3191 | 0.89666
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 | [6 empty] | 60
| 65 | [6 empty] | 75
| 80 |  | 2.1177 | 2.0135 | 0.0074374 | -4.9189 | 2.3639 | 90 | -4.198 | 2.3632
| 95 |  | 2.1177 | 2.0364 | 0.0056285 | -3.8376 | 2.3628 | 105 | -3.4104 | 2.3632
| 110 |  | 2.1167 | 2.049 | 0.0058233 | -3.1969 | 2.3634 | 120 | -2.9234 | 2.361
| 125 |  | 2.1197 | 2.0606 | 0.0047269 | -2.7866 | 2.3598 | 135 | -2.3093 | 2.364
| 140 |  | 2.1137 | 2.0699 | 0.0040306 | -2.0707 | 2.3661 | 150 | -1.6851 | 2.3696
| 155 |  | 2.1087 | 2.0772 | 0.0045042 | -1.4923 | 2.3714 | 165 | -1.5234 | 2.3707
| 170 |  | 2.1097 | 2.0772 | 0.0045099 | -1.5389 | 2.3703 | 180 | -1.4631 | 2.3703
| 185 |  | 2.1097 | 2.0796 | 0.0019612 | -1.4252 | 2.3703 | 195 | -1.4311 | 2.3695
| 200 |  | 2.1107 | 2.0804 | 0.0028389 | -1.434 | 2.3692 | 210 | -1.3464 | 2.3706
| 215 |  | 2.1087 | 2.0812 | 0.0032957 | -1.3026 | 2.3714 | 225 | -1.3392 | 2.3691
| 230 |  | 2.1117 | 2.083 | 0.0045952 | -1.3575 | 2.368 | 240 | -1.2225 | 2.3695
| 245 |  | 2.1097 | 2.0853 | 0.0049956 | -1.155 | 2.3702 | 255 | -1.2214 | 2.3702
| 260 |  | 2.1097 | 2.0832 | 0.0046853 | -1.2545 | 2.3702 | 270 | -1.2328 | 2.371
| 275 |  | 2.1087 | 2.0829 | 0.0045303 | -1.2219 | 2.3713 | 285 | -1.2212 | 2.3698
| 290 |  | 2.1107 | 2.0849 | 0.0050037 | -1.2208 | 2.3691 | 300 | -1.2585 | 2.3687
| 305 |  | 2.1112 | 2.0842 | 0.0049404 | -1.2773 | 2.3686 | 315 | -1.2834 | 2.3682
| 320 |  | 2.1117 | 2.0845 | 0.0049746 | -1.2865 | 2.368 | 330 | -1.2964 | 2.3687
| 335 |  | 2.1107 | 2.0832 | 0.0046814 | -1.3013 | 2.3691 | 345 | -1.3262 | 2.3684
| 350 |  | 2.1117 | 2.0834 | 0.0047484 | -1.3386 | 2.368 | 360 | -1.3573 | 2.3676
| 365 |  | 2.1122 | 2.0833 | 0.0046959 | -1.3667 | 2.3675 | 375 | -1.3482 | 2.3682
| 380 |  | 2.1112 | 2.0829 | 0.0045281 | -1.3389 | 2.3686 | 390 | -1.3515 | 2.3686
| 395 |  | 2.1112 | 2.0825 | 0.0043337 | -1.3579 | 2.3686 | 405 | -1.3764 | 2.3678
| 410 |  | 2.1122 | 2.0829 | 0.0045359 | -1.3856 | 2.3675 | 420 | -1.4645 | 2.3675
| 425 |  | 2.1122 | 2.0804 | 0.0028069 | -1.504 | 2.3675 | 435 | -1.5321 | 2.3671
| 440 |  | 2.1127 | 2.08 | 0 | -1.5462 | 2.367 | 450 | -1.4617 | 2.3681
| 455 |  | 2.1112 | 2.0812 | 0.0033076 | -1.4194 | 2.3686 | 465 | -1.4539 | 2.3682
| 470 |  | 2.1117 | 2.0806 | 0.0024038 | -1.4712 | 2.3681 | 480 | -1.4838 | 2.3681
| 485 |  | 2.1117 | 2.0802 | 0.0042567 | -1.4901 | 2.3681 | 495 | -1.3981 | 2.369
| 500 | 2.11 | 2.1103 | 2.0818 | 0.0038501 | -1.3521 | 2.3695 | 510 | -1.4655 | 2.3674
| 515 |  | 2.1132 | 2.081 | 0.0030047 | -1.5222 | 2.3664 | 525 | -1.5508 | 2.3668
| 530 |  | 2.1127 | 2.0796 | 0.0029473 | -1.5652 | 2.367
| 545 |  | 2.1127 | 2.081 | 0.0030224 | -1.4989 | 2.3669
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 2.013 | 1.884 | 0.005 | -6.408 | 1.4 | 45
| 60 |  | 2.016 | 1.911 | 0.006 | -5.208 | 1.4 | 60
LEI | 75 |  | 2.015 | 1.928 | 0.006 | -4.318 | 1.4 | 75 | -4.318 | 1.4
| 90 |  | 2.016 | 1.944 | 0.005 | -3.571 | 1.4 | 90 | -3.571 | 1.4
| 105 |  | 2.014 | 1.956 | 0.005 | -2.88 | 1.4 | 105 | -2.88 | 1.4
| 120 |  | 2.015 | 1.967 | 0.006 | -2.382 | 1.4 | 120 | -2.382 | 1.4
| 135 |  | 2.014 | 1.974 | 0.006 | -1.986 | 1.4 | 135 | -1.986 | 1.4
| 150 |  | 2.012 | 1.976 | 0.006 | -1.789 | 1.4 | 150 | -1.789 | 1.4
| 165 |  | 2.014 | 1.979 | 0.006 | -1.738 | 1.4 | 165 | -1.738 | 1.4
| 180 |  | 2.013 | 1.984 | 0.005 | -1.441 | 1.4 | 180 | -1.441 | 1.4
| 195 |  | 2.012 | 1.982 | 0.005 | -1.491 | 1.4 | 195 | -1.491 | 1.4
| 210 |  | 2.013 | 1.984 | 0.005 | -1.441 | 1.4 | 210 | -1.441 | 1.4
| 225 |  | 2.013 | 1.977 | 0.005 | -1.788 | 1.4 | 225 | -1.788 | 1.4
| 240 |  | 2.011 | 1.977 | 0.006 | -1.691 | 1.4 | 240 | -1.691 | 1.4
| 255 |  | 2.012 | 1.974 | 0.005 | -1.889 | 1.4 | 255 | -1.889 | 1.4
| 270 |  | 2.014 | 1.968 | 0.006 | -2.284 | 1.4 | 270 | -2.284 | 1.4
| 285 |  | 2.012 | 1.968 | 0.005 | -2.187 | 1.4 | 285 | -2.187 | 1.4
| 300 |  | 2.011 | 1.964 | 0.005 | -2.337 | 1.4 | 300 | -2.337 | 1.4
| 315 |  | 2.014 | 1.951 | 0.006 | -3.128 | 1.4 | 315 | -3.128 | 1.4
| 330 | [6 empty] | 330
| 345 | [6 empty] | 345
| 360 | [6 empty] | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-L 5
| 5 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.12247
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 4.8907 | 4.543 | 0.0067495 | -7.1102 | 0.67474 | 45
| 52.5 |  | 4.8952 | 4.663 | 0.0067495 | -4.7442 | 0.67412 | 60 |  |  | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 4.8938 | 4.696 | 0.0084327 | -4.0411 | 0.69476 | 75 | -3.6681 | 0.69436 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 4.8994 | 4.738 | 0.0078881 | -3.2951 | 0.69396 | 90 | -2.9318 | 0.69422 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 4.8957 | 4.77 | 0.008165 | -2.5684 | 0.69448 | 105 | -2.3558 | 0.69389 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 4.9041 | 4.799 | 0.0073786 | -2.1433 | 0.6933 | 120 | -1.768 | 0.69379 |  | 2.1197 | 0.25447 | 0.19016 |  | 1.0554 | 1.6748 | 1.3909 |  | 0.62904 | 0.026349 |  | 0.77954 |  | 0.70646 |  |  |  |  | 0.54063 | 1.5188 | 0.25495 | 0.24442 |  | 0.58141
| 127.5 | 4.897 | 4.8972 | 4.829 | 0.0073786 | -1.3927 | 0.69427 | 135 | -1.4214 | 0.69393 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 4.9021 | 4.831 | 0.011972 | -1.4501 | 0.69358 | 150 | -1.488 | 0.6944 | 45
| 157.5 |  | 4.8906 | 4.816 | 0.0069921 | -1.5258 | 0.69521 | 165 | -1.7028 | 0.68461 | 60
| 172.5 |  | 4.896 | 4.804 | 0.005164 | -1.8799 | 0.67401 | 180 | -2.1064 | 0.67337 | 75 | 2.743 | 0.58118 |  |  | 1.1642 | 1.8525 |  |  | 1.048 |  |  | 0.57922
| 187.5 |  | 4.9054 | 4.791 | 0.0056765 | -2.3329 | 0.67272 | 195 |  |  | 90 | 2.6867 | 0.56314 | 0.7357 |  | 1.0691 | 1.8237 | 1.2106 |  | 1.1088 | 0.24268 |  | 0.50556 |  | 0.20808
| 202.5 |  | 4.9036 | 4.78 | 0.0066667 | -2.5208 | 0.67297 | 210 |  |  | 105 | 2.3773 | 0.56716 | 0.49848 |  | 0.99027 | 1.5398 | 1.2412 |  | 0.92928 | 0.012193 |  | 0.42893 |  | 0.36373
| 217.5 |  | 4.9255 | 4.659 | 0.008756 | -5.4114 | 0.66998 | 225 |  |  | 120 | 2.3073 | 0.68266 | 0.49224 |  | 1.1179 | 1.3396 | 1.1944 |  | 0.7199 | 0.089007 |  | 0.44575 |  | 0.47632 |  |  |  |  | 2.2245 | 0.30354 | 1.4037 | 1.2862 |  | 0.86093
[8 empty] | 240 |  |  | 135 | 2.251 | 0.53809 | 0.43427 |  | 1.2588 | 1.4589 | 1.2169 |  | 0.50728 | 0.025285 |  | 0.71368 |  | 0.64877 |  |  |  |  | 2.0322 | 0.13411 | 1.3733 | 1.1934 |  | 0.54689
[8 empty] | 255 |  |  | 150 | 2.0258 | 0.26922 | 0.15477 |  | 1.0899 | 1.5678 | 1.3607 |  | 0.50687 | 0.073058 |  | 0.79843 |  | 0.76909 |  |  |  |  | 1.6303 | 0.10844 | 1.2718 | 1.1587 |  | 0.36454
[8 empty] | 270 |  |  | 165 | 1.7033 | 0.064224 | 0.14119 |  | 0.71769 | 1.644 | 1.4349 |  | 0.64663 | 0.083041 |  | 0.74255 |  | 0.73005 |  |  |  |  | 1.1373 | 0.37547 | 1.1356 | 1.053 |  | 0.29146
[8 empty] | 285 |  |  | 180 | 1.0848 | 0.40336 | 0.53003 |  | 0.3854 | 1.457 | 1.3931 |  | 0.48655 | 0.17805 |  | 0.72964 |  | 0.81087 |  |  |  |  | 0.56504 | 0.42434 | 0.93483 | 0.98238 |  | 0.11402
[8 empty] | 300 |  |  | 195 | [5 empty] | 1.406 | 1.4093 |  | 0.5083 | 0.22772 |  | 0.67215 |  | 0.80875 | [5 empty] | 0.50636 | 0.82744 | 0.94337 |  | 0.072632
[8 empty] | 315 |  |  | 210 | [5 empty] | 1.5647 | 1.4506 |  | 0.51847 | 0.15281 |  | 0.78655 |  | 0.83588 | [5 empty] | 0.81745 | 0.72177 | 0.77432 |  | 0.14952
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.5733 | 1.4851 |  | 0.55511 | 0.1784 |  | 0.76191 |  | 0.83838 | [5 empty] | 0.71245 | 0.80286 | 0.86945 |  | 0.042818
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.8058 | 1.6952 |  | 0.82448 | 0.19521 |  | 0.72077 |  | 0.81923 | [5 empty] | 1.3565 | 0.49935 | 0.6295 |  | 0.28748
[8 empty] | 360 |  |  | 255 | [9 empty] | 0.27794 |  | 0.56839 |  | 0.76589 | [6 empty] | 0.085208 | 0.36329 |  | 0.5199
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.36354 |  | 0.57914 |  | 0.85801 | [6 empty] | 0.031679 | 0.40727 |  | 0.58234
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.45893 |  | 0.56378 |  | 0.93687 | [6 empty] | 0.065903 | 0.42413 |  | 0.65923
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.62125 |  | 0.79193 |  | 1.294 | [6 empty] | 0.34287 | 0.35819 |  | 1.1649
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.64869 |  | 1.6594 |  | 2.0793 | [6 empty] | 0.40042 | 0.34065 |  | 2.1452
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.71326 | [10 empty] | 0.43781 | 0.37365
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.93838 | [10 empty] | 0.54805 | 0.51695
[8 empty] | 465 |  |  | 360 | [9 empty] | 0.97026 | [10 empty] | 0.73374 | 0.3911
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.2967 | [10 empty] | 1.3765 | 0.18257
[8 empty] | 495 |  |  | 390 | [21 empty] | 0.40961
[8 empty] | 510 |  |  | 405 | [21 empty] | 0.12773
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.22243
435
450
45
PTB | 55 | 5.118 | 5.1439 | 4.7685 | 0.007181 | -7.2979 | 0.42 | 60
| 70 | 5.118 | 5.1402 | 4.8235 | 0.008458 | -6.1627 | 0.42 | 75 | -5.9193 | 0.42
| 85 | 5.118 | 5.1337 | 4.8548 | 0.0093522 | -5.4325 | 0.42 | 90 | -5.1364 | 0.42
| 100 | 5.118 | 5.1298 | 4.8967 | 0.0087706 | -4.544 | 0.42 | 105 | -4.3059 | 0.42
| 115 | 5.118 | 5.1271 | 4.9307 | 0.0054954 | -3.8296 | 0.42 | 120 | -3.6604 | 0.42
| 130 | 5.118 | 5.1213 | 4.9512 | 0.0076561 | -3.322 | 0.42 | 135 | -3.268 | 0.42
| 145 | 5.118 | 5.1187 | 4.9569 | 0.0047068 | -3.1598 | 0.42 | 150 | -3.1506 | 0.42
| 160 | 5.118 | 5.118 | 4.9577 | 0.007646 | -3.1322 | 0.42 | 165 | -3.0867 | 0.42
| 175 | 5.118 | 5.1193 | 4.9659 | 0.0063605 | -2.9957 | 0.42 | 180 | -2.9775 | 0.42
| 190 | 5.118 | 5.1217 | 4.9711 | 0.0064051 | -2.9411 | 0.42 | 195 | -2.9789 | 0.42
| 205 | 5.118 | 5.1235 | 4.967 | 0.006688 | -3.0546 | 0.42 | 210 | -3.0284 | 0.42
| 220 | 5.118 | 5.1255 | 4.973 | 0.0072403 | -2.9761 | 0.42 | 225 | -3.0428 | 0.42
| 235 | 5.118 | 5.1346 | 4.9715 | 0.0060152 | -3.1762 | 0.42 | 240 | -3.2776 | 0.42
| 250 | 5.118 | 5.1373 | 4.9585 | 0.0086397 | -3.4802 | 0.42 | 255
| 265 | 5.118 | 5.1434 | 4.9411 | 0.011209 | -3.9324 | 0.42 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 | [6 empty] | 45
| 60 |  | 5.1097 | 4.8425 | 0.011642 | -5.2298 | 0.77647 | 60
| 75 |  | 5.1063 | 4.888 | 0.010563 | -4.2757 | 0.77186 | 75 | -4.2757 | 0.77186
| 90 |  | 5.1031 | 4.9235 | 0.010894 | -3.52 | 0.77082 | 90 | -3.52 | 0.77082
| 105 |  | 5.1054 | 4.955 | 0.0076089 | -2.9465 | 0.7668 | 105 | -2.9465 | 0.7668
| 120 |  | 5.105 | 4.9785 | 0.0074516 | -2.4785 | 0.7662 | 120 | -2.4785 | 0.7662
| 135 |  | 5.1016 | 5.0005 | 0.0060481 | -1.9813 | 0.76572 | 135 | -1.9813 | 0.76572
| 150 |  | 5.1002 | 5.01 | 0.0056195 | -1.7682 | 0.76554 | 150 | -1.7682 | 0.76554
| 165 |  | 5.0954 | 5.012 | 0.0052315 | -1.6364 | 0.76605 | 165 | -1.6364 | 0.76605
| 180 |  | 5.0993 | 5.013 | 0.0065695 | -1.692 | 0.76616 | 180 | -1.692 | 0.76616
| 195 |  | 5.1017 | 5.013 | 0.007327 | -1.7383 | 0.7663 | 195 | -1.7383 | 0.7663
| 210 |  | 5.1001 | 5.016 | 0.0059824 | -1.6486 | 0.76567 | 210 | -1.6486 | 0.76567
| 225 | 5.1005 | 5.1005 | 5.016 | 0.0050262 | -1.6573 | 0.76431 | 225 | -1.6573 | 0.76431
| 240 |  | 5.1015 | 5.0155 | 0.0060481 | -1.6854 | 0.76551 | 240 | -1.6854 | 0.76551
| 255 |  | 5.105 | 5.013 | 0.0057124 | -1.8018 | 0.76471 | 255 | -1.8018 | 0.76471
| 270 |  | 5.1058 | 5.0095 | 0.0051042 | -1.8857 | 0.76432 | 270 | -1.8857 | 0.76432
| 285 |  | 5.1073 | 5.004 | 0.0050262 | -2.0222 | 0.76414 | 285 | -2.0222 | 0.76414
| 300 |  | 5.1092 | 4.995 | 0.0068825 | -2.2348 | 0.76509 | 300 | -2.2348 | 0.76509
| 315 |  | 5.1088 | 4.993 | 0.007327 | -2.2663 | 0.76558 | 315 | -2.2663 | 0.76558
| 330 |  | 5.106 | 4.9835 | 0.0087509 | -2.3997 | 0.76746 | 330 | -2.3997 | 0.76746
| 345 |  | 5.11 | 4.97 | 0.0079472 | -2.7403 | 0.76689 | 345 | -2.7403 | 0.76689
| 360 |  | 5.1046 | 4.962 | 0.0069585 | -2.7941 | 0.76804 | 360 | -2.7941 | 0.76804
| 375 |  | 5.1093 | 4.9455 | 0.0082558 | -3.2055 | 0.77142 | 375 | -3.2055 | 0.77142
| 390 |  | 5.1142 | 4.903 | 0.0080131 | -4.1293 | 0.77865 | 390
| 405 | [6 empty] | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 4.9947 | 4.8 | 0.01 | -3.8979 | 0.64836 | 120 | -3.8979 | 0.64836
| 135 |  | 4.9883 | 4.82 | 0.01 | -3.3733 | 0.65271 | 135 | -3.3733 | 0.65271
| 150 |  | 4.993 | 4.84 | 0.02 | -3.0645 | 0.66177 | 150 | -3.0645 | 0.66177
| 165 |  | 4.9892 | 4.85 | 0.01 | -2.7905 | 0.65649 | 165 | -2.7905 | 0.65649
| 180 |  | 4.9919 | 4.86 | 0.01 | -2.6425 | 0.65712 | 180 | -2.6425 | 0.65712
| 195 |  | 4.9989 | 4.87 | 0.01 | -2.5794 | 0.65662 | 195 | -2.5794 | 0.65662
| 210 |  | 4.9991 | 4.88 | 0.01 | -2.3826 | 0.65792 | 210 | -2.3826 | 0.65792
| 225 |  | 5.0039 | 4.88 | 0.02 | -2.4762 | 0.66422 | 225 | -2.4762 | 0.66422
| 240 |  | 4.9998 | 4.89 | 0.02 | -2.1959 | 0.66662 | 240 | -2.1959 | 0.66662
| 255 |  | 4.9943 | 4.9 | 0.02 | -1.889 | 0.66938 | 255 | -1.889 | 0.66938
| 270 | 4.9982 | 4.9958 | 4.9 | 0.02 | -1.9181 | 0.66899 | 270 | -1.9181 | 0.66899
| 285 |  | 4.9977 | 4.9 | 0.02 | -1.9548 | 0.66849 | 285 | -1.9548 | 0.66849
| 300 |  | 4.9941 | 4.9 | 0.02 | -1.8837 | 0.66945 | 300 | -1.8837 | 0.66945
| 315 |  | 4.9928 | 4.9 | 0.01 | -1.858 | 0.66227 | 315 | -1.858 | 0.66227
| 330 |  | 4.9873 | 4.89 | 0.02 | -1.9505 | 0.66992 | 330 | -1.9505 | 0.66992
| 345 |  | 4.9989 | 4.89 | 0.02 | -2.1793 | 0.66684 | 345 | -2.1793 | 0.66684
| 360 |  | 4.9919 | 4.89 | 0.02 | -2.0415 | 0.66869 | 360 | -2.0415 | 0.66869
| 375 |  | 4.9892 | 4.9 | 0.02 | -1.7883 | 0.67073 | 375 | -1.7883 | 0.67073
| 390 |  | 4.993 | 4.89 | 0.02 | -2.0631 | 0.6684 | 390 | -2.0631 | 0.6684
| 405 |  | 4.9883 | 4.9 | 0.02 | -1.7696 | 0.67098 | 405 | -1.7696 | 0.67098
| 420 |  | 4.9947 | 4.9 | 0.02 | -1.8957 | 0.66929 | 420 | -1.8957 | 0.66929
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 | [6 empty] | 60
| 65 | [6 empty] | 75
| 80 |  | 5.0577 | 4.8247 | 0.011739 | -4.6062 | 0.98965 | 90 | -3.826 | 0.99012
| 95 |  | 5.0517 | 4.8781 | 0.0076485 | -3.4358 | 0.99036 | 105 | -2.9618 | 0.9906
| 110 |  | 5.0487 | 4.9111 | 0.0071098 | -2.7248 | 0.99073 | 120 | -2.3665 | 0.99103
| 125 |  | 5.0457 | 4.9353 | 0.0057761 | -2.1874 | 0.99119 | 135 | -1.9495 | 0.99114
| 140 |  | 5.0457 | 4.9533 | 0.0087057 | -1.8306 | 0.99112 | 150 | -1.6762 | 0.99122
| 155 |  | 5.0447 | 4.964 | 0.0063135 | -1.599 | 0.99127 | 165 | -1.5318 | 0.99139
| 170 |  | 5.0437 | 4.9681 | 0.0068913 | -1.4982 | 0.99145 | 180 | -1.4678 | 0.99145
| 185 |  | 5.0437 | 4.9704 | 0.0048772 | -1.4526 | 0.99145 | 195 | -1.4515 | 0.99158
| 200 |  | 5.0427 | 4.9695 | 0.0054999 | -1.451 | 0.99164 | 210 | -1.4562 | 0.99164
| 215 |  | 5.0427 | 4.9691 | 0.0059864 | -1.4589 | 0.99164 | 225 | -1.4328 | 0.9919
| 230 |  | 5.0407 | 4.9691 | 0.0047823 | -1.4198 | 0.99203 | 240 | -1.4396 | 0.99204
| 245 |  | 5.0407 | 4.9676 | 0.007119 | -1.4495 | 0.99204 | 255 | -1.452 | 0.99191
| 260 |  | 5.0417 | 4.9684 | 0.0041552 | -1.4532 | 0.99184 | 270 | -1.4283 | 0.99197
| 275 |  | 5.0407 | 4.9693 | 0.0057444 | -1.4158 | 0.99203 | 285 | -1.4448 | 0.99197
| 290 |  | 5.0412 | 4.9676 | 0.0055669 | -1.4593 | 0.99194 | 300 | -1.4528 | 0.992
| 305 |  | 5.0407 | 4.9676 | 0.004272 | -1.4495 | 0.99204 | 315 | -1.4495 | 0.99204
| 320 |  | 5.0407 | 4.9676 | 0.0051288 | -1.4495 | 0.99204 | 330 | -1.5009 | 0.99191
| 335 |  | 5.0417 | 4.9647 | 0.0061124 | -1.5266 | 0.99185 | 345 | -1.5582 | 0.99179
| 350 |  | 5.0422 | 4.9628 | 0.0075395 | -1.5741 | 0.99176 | 360 | -1.5713 | 0.99169
| 365 |  | 5.0427 | 4.9635 | 0.0066201 | -1.5699 | 0.99166 | 375 | -1.5686 | 0.99166
| 380 |  | 5.0427 | 4.9636 | 0.0072203 | -1.568 | 0.99166 | 390 | -1.5707 | 0.99173
| 395 |  | 5.0422 | 4.9629 | 0.0072201 | -1.5721 | 0.99176 | 405 | -1.6158 | 0.99183
| 410 |  | 5.0417 | 4.9591 | 0.006627 | -1.6377 | 0.99187 | 420 | -1.6282 | 0.99174
| 425 |  | 5.0427 | 4.9608 | 0.0063328 | -1.6235 | 0.99167 | 435 | -1.6331 | 0.99187
| 440 |  | 5.0412 | 4.9586 | 0.0066531 | -1.6378 | 0.99197 | 450 | -1.5983 | 0.99203
| 455 |  | 5.0407 | 4.9611 | 0.0048987 | -1.5785 | 0.99206 | 465 | -1.6193 | 0.99193
| 470 |  | 5.0417 | 4.959 | 0.0064289 | -1.6397 | 0.99187 | 480 | -1.6505 | 0.992
| 485 |  | 5.0407 | 4.9572 | 0.006032 | -1.6559 | 0.99207 | 495 | -1.6347 | 0.99206
| 500 | 5.041 | 5.0407 | 4.9588 | 0.005182 | -1.6241 | 0.99206 | 510 | -1.6308 | 0.99213
| 515 |  | 5.0402 | 4.9578 | 0.0057512 | -1.6342 | 0.99216 | 525 | -1.6316 | 0.99216
| 530 |  | 5.0402 | 4.958 | 0.0075399 | -1.6302 | 0.99216
| 545 |  | 5.0377 | 4.9583 | 0.0057633 | -1.5755 | 0.99265
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 5.127 | 4.764 | 0.009 | -7.08 | 0.83 | 45
| 60 |  | 5.129 | 4.827 | 0.007 | -5.888 | 0.83 | 60
LEI | 75 |  | 5.126 | 4.873 | 0.007 | -4.936 | 0.83 | 75 | -4.936 | 0.83
| 90 |  | 5.127 | 4.917 | 0.007 | -4.096 | 0.83 | 90 | -4.096 | 0.83
| 105 |  | 5.125 | 4.949 | 0.007 | -3.434 | 0.83 | 105 | -3.434 | 0.83
| 120 |  | 5.125 | 4.972 | 0.008 | -2.985 | 0.83 | 120 | -2.985 | 0.83
| 135 |  | 5.122 | 4.979 | 0.007 | -2.792 | 0.83 | 135 | -2.792 | 0.83
| 150 |  | 5.122 | 4.985 | 0.008 | -2.675 | 0.83 | 150 | -2.675 | 0.83
| 165 |  | 5.12 | 4.993 | 0.009 | -2.48 | 0.83 | 165 | -2.48 | 0.83
| 180 |  | 5.117 | 4.988 | 0.008 | -2.521 | 0.83 | 180 | -2.521 | 0.83
| 195 |  | 5.116 | 4.988 | 0.009 | -2.502 | 0.83 | 195 | -2.502 | 0.83
| 210 |  | 5.115 | 4.985 | 0.008 | -2.542 | 0.83 | 210 | -2.542 | 0.83
| 225 |  | 5.115 | 4.986 | 0.01 | -2.522 | 0.83 | 225 | -2.522 | 0.83
| 240 |  | 5.112 | 4.984 | 0.007 | -2.504 | 0.83 | 240 | -2.504 | 0.83
| 255 |  | 5.109 | 4.984 | 0.008 | -2.447 | 0.83 | 255 | -2.447 | 0.83
| 270 |  | 5.112 | 4.982 | 0.01 | -2.543 | 0.83 | 270 | -2.543 | 0.83
| 285 |  | 5.108 | 4.972 | 0.01 | -2.662 | 0.83 | 285 | -2.662 | 0.83
| 300 |  | 5.106 | 4.946 | 0.008 | -3.134 | 0.83 | 300 | -3.134 | 0.83
| 315 |  | 5.107 | 4.895 | 0.008 | -4.151 | 0.83 | 315 | -4.151 | 0.83
| 330 | [6 empty] | 330
| 345 | [6 empty] | 345
| 360 | [6 empty] | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-L 8
| 8 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.12247
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 8.0557 | 7.467 | 0.013375 | -7.3074 | 0.64551 | 45
| 52.5 |  | 8.0595 | 7.665 | 0.0052705 | -4.8949 | 0.6452 | 60 |  |  | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 8.0563 | 7.724 | 0.0069921 | -4.1242 | 0.64546 | 75 | -3.7166 | 0.64524 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 8.0618 | 7.795 | 0.0070711 | -3.309 | 0.64502 | 90 | -2.8932 | 0.65135 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 8.0586 | 7.859 | 0.0073786 | -2.4774 | 0.65768 | 105 | -2.2378 | 0.65733 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 8.0672 | 7.906 | 0.0096609 | -1.9981 | 0.65698 | 120 | -1.6912 | 0.65746 |  | 2.1797 | 0.43881 | 0.74789 |  | 0.75453 | 1.4212 | 1.3634 |  | 1.3146 | 0.23221 |  | 0.24521 |  | 0.017734 |  |  |  |  | 0.68216 | 1.4702 | 0.16794 | 0.074817 |  | 0.091575
| 127.5 | 8.0549 | 8.0555 | 7.944 | 0.005164 | -1.3843 | 0.65793 | 135 | -1.4154 | 0.65752 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 8.0657 | 7.949 | 0.0056765 | -1.4465 | 0.65711 | 150 | -1.4717 | 0.65781 | 45
| 157.5 |  | 8.0485 | 7.928 | 0.0063246 | -1.4969 | 0.65851 | 165 | -1.7144 | 0.65805 | 60
| 172.5 |  | 8.0597 | 7.904 | 0.0069921 | -1.9319 | 0.65759 | 180 | -2.1423 | 0.65127 | 75 | 2.814 | 0.80528 |  |  | 1.0124 | 1.5284 |  |  | 1.624 |  |  | 0.10632
| 187.5 |  | 8.0627 | 7.873 | 0.0082327 | -2.3526 | 0.64495 | 195 |  |  | 90 | 2.7485 | 0.95091 | 1.2501 |  | 1.1503 | 1.3174 | 1.3376 |  | 1.4057 | 0.16205 |  | 0.086487 |  | 0.083013
| 202.5 |  | 8.0592 | 7.832 | 0.012293 | -2.8186 | 0.65764 | 210 |  |  | 105 | 2.6023 | 0.93842 | 1.075 |  | 1.1099 | 1.2152 | 1.4079 |  | 1.3169 | 0.017878 |  | 0.063885 |  | 0.05214
| 217.5 |  | 8.0718 | 7.617 | 0.0067495 | -5.6345 | 0.64422 | 225 |  |  | 120 | 2.4945 | 0.97375 | 1.0468 |  | 1.1708 | 1.0792 | 1.3296 |  | 1.134 | 0.043842 |  | 0.08375 |  | 0.1433 |  |  |  |  | 2.3182 | 0.1516 | 1.086 | 1.2864 |  | 1.1189
[8 empty] | 240 |  |  | 135 | 2.3647 | 0.8281 | 0.99397 |  | 1.1857 | 1.1349 | 1.2583 |  | 0.98398 | 0.057932 |  | 0.24545 |  | 0.21207 |  |  |  |  | 1.9748 | 0.090559 | 0.94349 | 1.0025 |  | 0.7731
[8 empty] | 255 |  |  | 150 | 2.1646 | 0.59886 | 0.66253 |  | 0.968 | 1.2212 | 1.4513 |  | 1.0422 | 0.0098064 |  | 0.27775 |  | 0.32414 |  |  |  |  | 1.5787 | 0.38242 | 0.81172 | 0.93323 |  | 0.58884
[8 empty] | 270 |  |  | 165 | 1.7478 | 0.30664 | 0.22882 |  | 0.48654 | 1.1922 | 1.5428 |  | 1.1968 | 0.10539 |  | 0.13422 |  | 0.26859 |  |  |  |  | 1.1085 | 0.50127 | 0.69107 | 0.90399 |  | 0.61818
[8 empty] | 285 |  |  | 180 | 1.1825 | 0.21108 | 0.23154 |  | 0.027438 | 1.2806 | 1.5015 |  | 1.2298 | 0.0055373 |  | 0.18766 |  | 0.20552 |  |  |  |  | 0.59551 | 0.52855 | 0.75552 | 0.85189 |  | 0.62956
[8 empty] | 300 |  |  | 195 | [5 empty] | 1.3467 | 1.5254 |  | 1.1905 | 0.044891 |  | 0.27441 |  | 0.25931 | [5 empty] | 0.79347 | 0.61032 | 0.6406 |  | 0.36842
[8 empty] | 315 |  |  | 210 | [5 empty] | 1.2833 | 1.527 |  | 1.2497 | 0.012434 |  | 0.17419 |  | 0.2102 | [5 empty] | 0.89544 | 0.4697 | 0.54596 |  | 0.32631
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.4808 | 1.6439 |  | 1.3521 | 0.073983 |  | 0.26888 |  | 0.22059 | [5 empty] | 0.98421 | 0.58878 | 0.58274 |  | 0.35048
[8 empty] | 345 |  |  | 240 | [5 empty] | 1.5912 | 1.9657 |  | 1.5906 | 0.07011 |  | 0.18471 |  | 0.28624 | [5 empty] | 1.4312 | 0.3525 | 0.47978 |  | 0.18291
[8 empty] | 360 |  |  | 255 | [9 empty] | 0.11069 |  | 0.14459 |  | 0.28616 | [6 empty] | 0.0076917 | 0.13506 |  | 0.15587
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.23751 |  | 0.076136 |  | 0.35044 | [6 empty] | 0.16319 | 0.086295 |  | 0.26879
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.21101 |  | 0.25907 |  | 0.52682 | [6 empty] | 0.19591 | 0.018857 |  | 0.51303
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.36153 |  | 0.61287 |  | 1.0932 | [6 empty] | 0.16219 | 0.22688 |  | 0.8741
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.50501 |  | 1.7599 |  | 2.5479 | [6 empty] | 0.30395 | 0.23251 |  | 2.3438
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.62405 | [10 empty] | 0.39404 | 0.26252
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.82445 | [10 empty] | 0.34886 | 0.5288
[8 empty] | 465 |  |  | 360 | [9 empty] | 1.1154 | [10 empty] | 0.55275 | 0.64334
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.4771 | [10 empty] | 1.3047 | 0.2003
[8 empty] | 495 |  |  | 390 | [21 empty] | 0.40222
[8 empty] | 510 |  |  | 405 | [21 empty] | 0.33875
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.21182
435
450
45
PTB | 55 | 7.9813 | 8.0211 | 7.4683 | 0.013917 | -6.8915 | 0.41 | 60
| 70 | 7.9813 | 8.0137 | 7.5174 | 0.011823 | -6.1935 | 0.41 | 75 | -5.8953 | 0.41
| 85 | 7.9813 | 8.0065 | 7.5822 | 0.011875 | -5.2989 | 0.41 | 90 | -5.0352 | 0.41
| 100 | 7.9813 | 8.0006 | 7.64 | 0.010954 | -4.5077 | 0.41 | 105 | -4.2788 | 0.41
| 115 | 7.9813 | 7.9955 | 7.69 | 0.0087706 | -3.8211 | 0.41 | 120 | -3.648 | 0.41
| 130 | 7.9813 | 7.9895 | 7.7257 | 0.0057275 | -3.3018 | 0.41 | 135 | -3.2705 | 0.41
| 145 | 7.9813 | 7.9829 | 7.7268 | 0.0072283 | -3.208 | 0.41 | 150 | -3.1703 | 0.41
| 160 | 7.9813 | 7.9813 | 7.7343 | 0.0069007 | -3.0951 | 0.41 | 165 | -3.0864 | 0.41
| 175 | 7.9813 | 7.9825 | 7.7375 | 0.0075154 | -3.069 | 0.41 | 180 | -3.0638 | 0.41
| 190 | 7.9813 | 7.9866 | 7.7427 | 0.0082741 | -3.0534 | 0.41 | 195 | -3.0465 | 0.41
| 205 | 7.9813 | 7.9909 | 7.7485 | 0.0090739 | -3.0328 | 0.41 | 210 | -3.0451 | 0.41
| 220 | 7.9813 | 7.994 | 7.7486 | 0.0089087 | -3.0699 | 0.41 | 225 | -3.148 | 0.41
| 235 | 7.9813 | 8.0052 | 7.7407 | 0.0076636 | -3.3042 | 0.41 | 240 | -3.3916 | 0.41
| 250 | 7.9813 | 8.0098 | 7.7241 | 0.0090701 | -3.5664 | 0.41 | 255
| 265 | 7.9813 | 8.0224 | 7.6863 | 0.0083887 | -4.1899 | 0.41 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 | [6 empty] | 45
| 60 |  | 8.1837 | 7.743 | 0.011286 | -5.3849 | 0.78188 | 60
| 75 |  | 8.1812 | 7.81 | 0.0091766 | -4.537 | 0.77884 | 75 | -4.537 | 0.77884
| 90 |  | 8.188 | 7.8715 | 0.010894 | -3.8652 | 0.77819 | 90 | -3.8652 | 0.77819
| 105 |  | 8.1834 | 7.9215 | 0.009333 | -3.2002 | 0.77766 | 105 | -3.2002 | 0.77766
| 120 |  | 8.1852 | 7.965 | 0.010513 | -2.69 | 0.77776 | 120 | -2.69 | 0.77776
| 135 |  | 8.1833 | 7.998 | 0.0069585 | -2.2642 | 0.77664 | 135 | -2.2642 | 0.77664
| 150 |  | 8.1817 | 8.011 | 0.010208 | -2.0862 | 0.77783 | 150 | -2.0862 | 0.77783
| 165 |  | 8.181 | 8.015 | 0.0076089 | -2.0289 | 0.77698 | 165 | -2.0289 | 0.77698
| 180 |  | 8.177 | 8.0195 | 0.0099868 | -1.9265 | 0.77816 | 180 | -1.9265 | 0.77816
| 195 |  | 8.1728 | 8.0215 | 0.0058714 | -1.8517 | 0.7773 | 195 | -1.8517 | 0.7773
| 210 |  | 8.1799 | 8.024 | 0.0088258 | -1.9063 | 0.77749 | 210 | -1.9063 | 0.77749
| 225 | 8.1765 | 8.1765 | 8.0265 | 0.0081273 | -1.8343 | 0.77708 | 225 | -1.8343 | 0.77708
| 240 |  | 8.1774 | 8.0155 | 0.0075915 | -1.9796 | 0.77733 | 240 | -1.9796 | 0.77733
| 255 |  | 8.1844 | 8.016 | 0.0088258 | -2.0574 | 0.77704 | 255 | -2.0574 | 0.77704
| 270 |  | 8.1795 | 8.0025 | 0.0096655 | -2.1637 | 0.77778 | 270 | -2.1637 | 0.77778
| 285 |  | 8.1816 | 8.005 | 0.0076089 | -2.1583 | 0.77693 | 285 | -2.1583 | 0.77693
| 300 |  | 8.1822 | 7.9915 | 0.0081273 | -2.3305 | 0.77708 | 300 | -2.3305 | 0.77708
| 315 |  | 8.1831 | 7.981 | 0.011653 | -2.4695 | 0.77832 | 315 | -2.4695 | 0.77832
| 330 |  | 8.1842 | 7.97 | 0.008584 | -2.617 | 0.7772 | 330 | -2.617 | 0.7772
| 345 |  | 8.1861 | 7.956 | 0.010463 | -2.8107 | 0.77785 | 345 | -2.8107 | 0.77785
| 360 |  | 8.1874 | 7.9335 | 0.0104 | -3.1009 | 0.77827 | 360 | -3.1009 | 0.77827
| 375 |  | 8.1948 | 7.9065 | 0.010894 | -3.5179 | 0.77879 | 375 | -3.5179 | 0.77879
| 390 |  | 8.1924 | 7.844 | 0.011877 | -4.2525 | 0.78328 | 390
| 405 | [6 empty] | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 7.9695 | 7.67 | 0.01 | -3.7585 | 0.58992 | 120 | -3.7585 | 0.58992
| 135 |  | 7.9651 | 7.71 | 0.04 | -3.2031 | 0.61001 | 135 | -3.2031 | 0.61001
| 150 |  | 7.9703 | 7.74 | 0.02 | -2.889 | 0.59848 | 150 | -2.889 | 0.59848
| 165 |  | 7.9663 | 7.75 | 0.03 | -2.7148 | 0.60528 | 165 | -2.7148 | 0.60528
| 180 |  | 7.9732 | 7.76 | 0.02 | -2.6745 | 0.59955 | 180 | -2.6745 | 0.59955
| 195 |  | 7.976 | 7.78 | 0.03 | -2.4579 | 0.60609 | 195 | -2.4579 | 0.60609
| 210 |  | 7.9795 | 7.79 | 0.04 | -2.3751 | 0.61383 | 210 | -2.3751 | 0.61383
| 225 |  | 7.9933 | 7.8 | 0.03 | -2.4186 | 0.60501 | 225 | -2.4186 | 0.60501
| 240 |  | 7.9861 | 7.8 | 0.03 | -2.3298 | 0.60609 | 240 | -2.3298 | 0.60609
| 255 |  | 7.9747 | 7.81 | 0.03 | -2.065 | 0.60855 | 255 | -2.065 | 0.60855
| 270 | 7.9787 | 7.9797 | 7.82 | 0.03 | -2.0014 | 0.60855 | 270 | -2.0014 | 0.60855
| 285 |  | 7.9766 | 7.82 | 0.03 | -1.9634 | 0.60902 | 285 | -1.9634 | 0.60902
| 300 |  | 7.9729 | 7.8 | 0.04 | -2.1685 | 0.61557 | 300 | -2.1685 | 0.61557
| 315 |  | 7.9728 | 7.8 | 0.03 | -2.167 | 0.60808 | 315 | -2.167 | 0.60808
| 330 |  | 7.9671 | 7.79 | 0.04 | -2.2234 | 0.61568 | 330 | -2.2234 | 0.61568
| 345 |  | 7.976 | 7.78 | 0.06 | -2.4579 | 0.63459 | 345 | -2.4579 | 0.63459
| 360 |  | 7.9732 | 7.77 | 0.04 | -2.549 | 0.61328 | 360 | -2.549 | 0.61328
| 375 |  | 7.9663 | 7.79 | 0.04 | -2.2127 | 0.61581 | 375 | -2.2127 | 0.61581
| 390 |  | 7.9703 | 7.78 | 0.06 | -2.3871 | 0.63546 | 390 | -2.3871 | 0.63546
| 405 |  | 7.9651 | 7.78 | 0.04 | -2.3242 | 0.61524 | 405 | -2.3242 | 0.61524
| 420 |  | 7.9695 | 7.79 | 0.02 | -2.2528 | 0.60239 | 420 | -2.2528 | 0.60239
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 | [6 empty] | 60
| 65 | [6 empty] | 75
| 80 |  | 8.077 | 7.6957 | 0.022776 | -4.7208 | 0.61973 | 90 | -4.0277 | 0.61986
| 95 |  | 8.071 | 7.7739 | 0.016572 | -3.6811 | 0.61992 | 105 | -3.2181 | 0.62008
| 110 |  | 8.066 | 7.8251 | 0.013249 | -2.9866 | 0.62016 | 120 | -2.6461 | 0.62031
| 125 |  | 8.062 | 7.8624 | 0.011755 | -2.4758 | 0.62038 | 135 | -2.3222 | 0.62041
| 140 |  | 8.061 | 7.88 | 0.010831 | -2.2454 | 0.62043 | 150 | -2.0763 | 0.62056
| 155 |  | 8.058 | 7.8975 | 0.010186 | -1.9918 | 0.62062 | 165 | -1.9233 | 0.62067
| 170 |  | 8.057 | 7.9048 | 0.0096414 | -1.889 | 0.62069 | 180 | -1.9321 | 0.62044
| 185 |  | 8.062 | 7.9045 | 0.0077814 | -1.9536 | 0.62031 | 195 | -1.8967 | 0.62061
| 200 |  | 8.056 | 7.9055 | 0.0077825 | -1.8682 | 0.62076 | 210 | -1.8938 | 0.62077
| 215 |  | 8.056 | 7.9024 | 0.0085014 | -1.9067 | 0.62077 | 225 | -1.9085 | 0.62082
| 230 |  | 8.055 | 7.9012 | 0.0091664 | -1.9094 | 0.62085 | 240 | -1.9094 | 0.62085
| 245 |  | 8.055 | 7.9012 | 0.0071544 | -1.9094 | 0.62085 | 255 | -1.9465 | 0.6208
| 260 |  | 8.056 | 7.8977 | 0.0093147 | -1.965 | 0.62078 | 270 | -1.9256 | 0.62087
| 275 |  | 8.054 | 7.9005 | 0.0085079 | -1.9059 | 0.62092 | 285 | -1.9469 | 0.6208
| 290 |  | 8.0565 | 7.898 | 0.0089929 | -1.9674 | 0.62074 | 300 | -1.9682 | 0.62074
| 305 |  | 8.0565 | 7.8979 | 0.0090758 | -1.9686 | 0.62074 | 315 | -1.963 | 0.62079
| 320 |  | 8.0555 | 7.8976 | 0.0099913 | -1.9602 | 0.62081 | 330 | -1.9916 | 0.62082
| 335 |  | 8.0555 | 7.8938 | 0.0093586 | -2.0073 | 0.62082 | 345 | -1.984 | 0.62076
| 350 |  | 8.0565 | 7.8976 | 0.0065198 | -1.9723 | 0.62074 | 360 | -1.9822 | 0.62071
| 365 |  | 8.057 | 7.8969 | 0.0093775 | -1.9871 | 0.6207 | 375 | -2.0358 | 0.62068
| 380 |  | 8.0575 | 7.8915 | 0.0081428 | -2.0602 | 0.62067 | 390 | -2.0264 | 0.62072
| 395 |  | 8.0565 | 7.8946 | 0.0087845 | -2.0096 | 0.62074 | 405 | -2.0253 | 0.62074
| 410 |  | 8.0565 | 7.8927 | 0.0083209 | -2.0331 | 0.62075 | 420 | -2.0677 | 0.6207
| 425 |  | 8.0575 | 7.8895 | 0.010006 | -2.085 | 0.62067 | 435 | -2.0687 | 0.62075
| 440 |  | 8.056 | 7.89 | 0.009118 | -2.0606 | 0.62079 | 450 | -2.0474 | 0.62081
| 455 |  | 8.0555 | 7.8911 | 0.0087507 | -2.0408 | 0.62082 | 465 | -2.0368 | 0.62085
| 470 |  | 8.055 | 7.8911 | 0.0076145 | -2.0348 | 0.62086 | 480 | -2.0323 | 0.62086
| 485 |  | 8.055 | 7.8914 | 0.0090555 | -2.031 | 0.62086 | 495 | -2.055 | 0.62086
| 500 | 8.055 | 8.055 | 7.8885 | 0.009217 | -2.067 | 0.62087 | 510 | -2.0647 | 0.62092
| 515 |  | 8.054 | 7.8878 | 0.0084126 | -2.0636 | 0.62094 | 525 | -2.0679 | 0.62099
| 530 |  | 8.053 | 7.8863 | 0.0079791 | -2.07 | 0.62102
| 545 |  | 8.052 | 7.8865 | 0.0084942 | -2.0554 | 0.62109
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 8.32 | 7.737 | 0.013 | -7.007 | 0.64 | 45
| 60 |  | 8.327 | 7.861 | 0.01 | -5.596 | 0.64 | 60
LEI | 75 |  | 8.331 | 7.944 | 0.01 | -4.645 | 0.64 | 75 | -4.645 | 0.64
| 90 |  | 8.323 | 7.994 | 0.011 | -3.953 | 0.64 | 90 | -3.953 | 0.64
| 105 |  | 8.331 | 8.059 | 0.01 | -3.265 | 0.64 | 105 | -3.265 | 0.64
| 120 |  | 8.323 | 8.092 | 0.011 | -2.775 | 0.64 | 120 | -2.775 | 0.64
| 135 |  | 8.318 | 8.109 | 0.011 | -2.513 | 0.64 | 135 | -2.513 | 0.64
| 150 |  | 8.319 | 8.122 | 0.01 | -2.368 | 0.64 | 150 | -2.368 | 0.64
| 165 |  | 8.313 | 8.133 | 0.009 | -2.165 | 0.64 | 165 | -2.165 | 0.64
| 180 |  | 8.312 | 8.136 | 0.009 | -2.117 | 0.64 | 180 | -2.117 | 0.64
| 195 |  | 8.309 | 8.132 | 0.011 | -2.13 | 0.64 | 195 | -2.13 | 0.64
| 210 |  | 8.307 | 8.134 | 0.01 | -2.083 | 0.64 | 210 | -2.083 | 0.64
| 225 |  | 8.304 | 8.129 | 0.01 | -2.107 | 0.64 | 225 | -2.107 | 0.64
| 240 |  | 8.308 | 8.128 | 0.012 | -2.167 | 0.64 | 240 | -2.167 | 0.64
| 255 |  | 8.303 | 8.12 | 0.01 | -2.204 | 0.64 | 255 | -2.204 | 0.64
| 270 |  | 8.299 | 8.113 | 0.012 | -2.241 | 0.64 | 270 | -2.241 | 0.64
| 285 |  | 8.301 | 8.1 | 0.012 | -2.421 | 0.64 | 285 | -2.421 | 0.64
| 300 |  | 8.299 | 8.054 | 0.014 | -2.952 | 0.64 | 300 | -2.952 | 0.64
| 315 |  | 8.294 | 7.941 | 0.012 | -4.256 | 0.64 | 315 | -4.256 | 0.64
| 330 | [6 empty] | 330
| 345 | [6 empty] | 345
| 360 | [6 empty] | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## T-L 12
| 12 m/s | [6 empty] | interpolation |  |  | U_st (%)
| d | vref | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | E(d) | U(E(d)) | 0.18779
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (%) | (%)
BEV/E+E | 37.5 |  | 12.0344 | 11.161 | 0.0099443 | -7.2579 | 0.63152 | 45
| 52.5 |  | 12.0284 | 11.435 | 0.093956 | -4.9335 | 0.80642 | 60 |  |  | color code for En: |  | 0 to 1 | 1 to 1.5 | 1.5 to 2 | > 2
| 67.5 |  | 12.0361 | 11.558 | 0.010328 | -3.9724 | 0.63974 | 75 | -3.535 | 0.6395 |  | lab pairs - open/open test section | [14 empty] | closed/closed test section |  |  | lab pairs - closed/open test section
| 82.5 |  | 12.0451 | 11.672 | 0.016193 | -3.0976 | 0.63926 | 90 | -2.6774 | 0.63938 |  | BEV/PTB | BEV/CMI | BEV/DWG | BEV/UL | BEV/LEI | PTB/CMI | PTB/DWG | PTB/UL | PTB/LEI | CMI/DWG | CMI/UL | CMI/LEI | DWG/UL | DWG/LEI | UL/LEI | DTI/CET | DTI/SHM | CET/SHM | CET/BEV | CET/PTB | CET/CMI | CET/DWG | CET/UL | CET/LEI | DTI/BEV | DTI/PTB | DTI/CMI | DTI/DWG | DTI/UL | DTI/LEI | SHM/BEV | SHM/PTB | SHM/CMI | SHM/DWG | SHM/UL | SHM/LEI
| 97.5 |  | 12.0408 | 11.769 | 0.011005 | -2.2571 | 0.63949 | 105 | -2.0242 | 0.63931 |  | lab to lab equivalence degrees En comparing calibartion data in the test section center | [14 empty] | En center |  |  | lab to lab equivalence degrees En comparing calibartion data in the test section center
| 112.5 |  | 12.0478 | 11.832 | 0.010328 | -1.7913 | 0.63912 | 120 | -1.4389 | 0.63952 |  | 2.6485 | 0.73379 | 1.035 |  | 0.53049 | 1.5367 | 1.6592 |  | 2.3599 | 0.16557 |  | 0.3111 |  | 0.57401 |  |  |  |  | 1.142 | 1.4819 | 0.2746 | 0.13087 |  | 0.6986
| 127.5 | 12.0413 | 12.0327 | 11.902 | 0.012293 | -1.0864 | 0.63992 | 135 | -1.1417 | 0.63948 | d (mm) | lab to lab equivalence degrees En for given insertion depth d | [14 empty] | En (d) |  |  | lab to lab equivalence degrees En for given insertion depth d
| 142.5 |  | 12.0492 | 11.905 | 0.010801 | -1.1969 | 0.63905 | 150 | -1.2694 | 0.63965 | 45
| 157.5 |  | 12.0264 | 11.865 | 0.011785 | -1.342 | 0.64026 | 165 | -1.5501 | 0.64004 | 60
| 172.5 |  | 12.0346 | 11.823 | 0.013375 | -1.7583 | 0.63982 | 180 | -2.0397 | 0.63961 | 75 | 3.2258 | 0.88636 |  |  | 0.55244 | 1.8742 |  |  | 2.9852 |  |  | 0.45605
| 187.5 |  | 12.0425 | 11.763 | 0.014944 | -2.321 | 0.6394 | 195 |  |  | 90 | 3.1501 | 1.0019 | 1.6819 |  | 0.75083 | 1.6793 | 1.4429 |  | 2.6618 | 0.46733 |  | 0.40115 |  | 1.0517
| 202.5 |  | 12.0308 | 11.715 | 0.010801 | -2.6252 | 0.64002 | 210 |  |  | 105 | 2.9254 | 0.99794 | 1.3495 |  | 0.80548 | 1.4851 | 1.5899 |  | 2.3424 | 0.17165 |  | 0.34724 |  | 0.625
| 217.5 |  | 12.0722 | 11.391 | 0.017288 | -5.6428 | 0.62954 | 225 |  |  | 120 | 2.8614 | 1.0814 | 1.3051 |  | 0.81047 | 1.3335 | 1.5721 |  | 2.2645 | 0.043446 |  | 0.43277 |  | 0.57071 |  |  |  |  | 2.8483 | 0.26323 | 1.4458 | 1.6608 |  | 2.2802
[8 empty] | 240 |  |  | 135 | 2.7601 | 0.84862 | 1.2531 |  | 0.81783 | 1.504 | 1.5231 |  | 2.1412 | 0.24105 |  | 0.176 |  | 0.50529 |  |  |  |  | 2.673 | 0.17929 | 1.5355 | 1.5331 |  | 2.0802
[8 empty] | 255 |  |  | 150 | 2.4904 | 0.64605 | 0.79584 |  | 0.55892 | 1.4925 | 1.7704 |  | 2.1465 | 0.04053 |  | 0.1913 |  | 0.27837 |  |  |  |  | 1.9558 | 0.37053 | 1.0995 | 1.2519 |  | 1.565
[8 empty] | 270 |  |  | 165 | 2.0757 | 0.33692 | 0.34632 |  | 0.14883 | 1.4711 | 1.8515 |  | 2.1693 | 0.041156 |  | 0.22842 |  | 0.223 |  |  |  |  | 1.2965 | 0.69993 | 0.82486 | 1.0229 |  | 1.2739
[8 empty] | 285 |  |  | 180 | 1.4847 | 0.18733 | 0.25161 |  | 0.73412 | 1.5348 | 1.9149 |  | 2.5567 | 0.03066 |  | 0.46032 |  | 0.51312 |  |  |  |  | 0.63813 | 0.82623 | 0.77432 | 0.94793 |  | 1.4776
[8 empty] | 300 |  |  | 195 | [5 empty] | 1.4109 | 1.9487 |  | 2.445 | 0.17322 |  | 0.49517 |  | 0.38037 | [5 empty] | 1.0326 | 0.5057 | 0.79804 |  | 1.1955
[8 empty] | 315 |  |  | 210 | [5 empty] | 1.7183 | 2.1029 |  | 2.5441 | 0.0023369 |  | 0.2742 |  | 0.32524 | [5 empty] | 1.2094 | 0.65504 | 0.77584 |  | 1.1188
[8 empty] | 330 |  |  | 225 | [5 empty] | 1.9953 | 2.4967 |  | 2.975 | 0.047001 |  | 0.33671 |  | 0.34503 | [5 empty] | 1.6876 | 0.53981 | 0.69113 |  | 1.0509
[8 empty] | 345 |  |  | 240 | [5 empty] | 2.17 | 2.7847 |  | 3.2674 | 0.10619 |  | 0.39223 |  | 0.33881 | [5 empty] | 1.9474 | 0.48954 | 0.69902 |  | 1.0512
[8 empty] | 360 |  |  | 255 | [9 empty] | 0.11153 |  | 0.42372 |  | 0.37014 | [6 empty] | 0.36935 | 0.56685 |  | 0.9503
[8 empty] | 375 |  |  | 270 | [9 empty] | 0.098522 |  | 0.21768 |  | 0.14002 | [6 empty] | 0.19105 | 0.34038 |  | 0.4882
[8 empty] | 390 |  |  | 285 | [9 empty] | 0.17616 |  | 0.042276 |  | 0.16467 | [6 empty] | 0.12517 | 0.35468 |  | 0.20248
[8 empty] | 405 |  |  | 300 | [9 empty] | 0.38838 |  | 0.25345 |  | 0.77768 | [6 empty] | 0.21278 | 0.20218 |  | 0.55682
[8 empty] | 420 |  |  | 315 | [9 empty] | 0.45795 |  | 1.6956 |  | 2.5871 | [6 empty] | 0.33906 | 0.13256 |  | 2.3972
[8 empty] | 435 |  |  | 330 | [9 empty] | 0.60207 | [10 empty] | 0.19806 | 0.46916
[8 empty] | 450 |  |  | 345 | [9 empty] | 0.79341 | [10 empty] | 0.26879 | 0.60717
[8 empty] | 465 |  |  | 360 | [9 empty] | 1.1437 | [10 empty] | 0.66295 | 0.54987
[8 empty] | 480 |  |  | 375 | [9 empty] | 1.5534 | [10 empty] | 0.98211 | 0.6574
[8 empty] | 495 |  |  | 390 | [21 empty] | 0.39434
[8 empty] | 510 |  |  | 405 | [21 empty] | 0.22655
[8 empty] | 525 |  |  | 420 | [21 empty] | 0.22739
435
450
45
PTB | 55 | 12.0203 | 12.0829 | 11.2194 | 0.017027 | -7.1464 | 0.44 | 60
| 70 | 12.0203 | 12.0697 | 11.2948 | 0.010874 | -6.42 | 0.44 | 75 | -6.1113 | 0.44
| 85 | 12.0203 | 12.0554 | 11.3931 | 0.015689 | -5.4937 | 0.44 | 90 | -5.1928 | 0.44
| 100 | 12.0203 | 12.0465 | 11.4935 | 0.017876 | -4.5909 | 0.44 | 105 | -4.3601 | 0.44
| 115 | 12.0203 | 12.0437 | 11.5742 | 0.01447 | -3.8984 | 0.44 | 120 | -3.7241 | 0.44
| 130 | 12.0203 | 12.0309 | 11.6248 | 0.011219 | -3.3755 | 0.44 | 135 | -3.346 | 0.44
| 145 | 12.0203 | 12.0249 | 11.6296 | 0.013467 | -3.2868 | 0.44 | 150 | -3.2587 | 0.44
| 160 | 12.0203 | 12.0203 | 11.6354 | 0.0083808 | -3.2025 | 0.44 | 165 | -3.2088 | 0.44
| 175 | 12.0203 | 12.0286 | 11.6412 | 0.0090893 | -3.2214 | 0.44 | 180 | -3.2256 | 0.44
| 190 | 12.0203 | 12.0357 | 11.6464 | 0.0086984 | -3.234 | 0.44 | 195 | -3.2296 | 0.44
| 205 | 12.0203 | 12.0392 | 11.6514 | 0.011127 | -3.2209 | 0.44 | 210 | -3.3327 | 0.44
| 220 | 12.0203 | 12.0523 | 11.6237 | 0.012452 | -3.5562 | 0.44 | 225 | -3.6131 | 0.44
| 235 | 12.0203 | 12.0629 | 11.6133 | 0.013218 | -3.7269 | 0.44 | 240 | -3.8367 | 0.44
| 250 | 12.0203 | 12.0771 | 11.5872 | 0.010214 | -4.0563 | 0.44 | 255
| 265 | 12.0203 | 12.0889 | 11.525 | 0.015029 | -4.6646 | 0.44 | 270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
CMI | 45 | [6 empty] | 45
| 60 |  | 12.3338 | 11.6895 | 0.016051 | -5.2238 | 0.75912 | 60
| 75 |  | 12.3266 | 11.7805 | 0.017313 | -4.4302 | 0.75877 | 75 | -4.4302 | 0.75877
| 90 |  | 12.3308 | 11.876 | 0.015694 | -3.6882 | 0.75756 | 90 | -3.6882 | 0.75756
| 105 |  | 12.3245 | 11.951 | 0.012096 | -3.0304 | 0.75679 | 105 | -3.0304 | 0.75679
| 120 |  | 12.3248 | 12.013 | 0.014546 | -2.5298 | 0.7572 | 120 | -2.5298 | 0.7572
| 135 |  | 12.3135 | 12.0675 | 0.014464 | -1.9981 | 0.75784 | 135 | -1.9981 | 0.75784
| 150 |  | 12.3105 | 12.074 | 0.012732 | -1.9214 | 0.75762 | 150 | -1.9214 | 0.75762
| 165 |  | 12.3122 | 12.0795 | 0.014318 | -1.8903 | 0.75791 | 165 | -1.8903 | 0.75791
| 180 |  | 12.3139 | 12.086 | 0.012732 | -1.8506 | 0.75738 | 180 | -1.8506 | 0.75738
| 195 |  | 12.3177 | 12.0755 | 0.01191 | -1.9662 | 0.75699 | 195 | -1.9662 | 0.75699
| 210 |  | 12.3076 | 12.087 | 0.014179 | -1.7923 | 0.75812 | 210 | -1.7923 | 0.75812
| 225 | 12.3119 | 12.3119 | 12.087 | 0.012183 | -1.8266 | 0.75683 | 225 | -1.8266 | 0.75683
| 240 |  | 12.3215 | 12.088 | 0.0076777 | -1.8954 | 0.75598 | 240 | -1.8954 | 0.75598
| 255 |  | 12.3228 | 12.0845 | 0.012344 | -1.9341 | 0.75673 | 255 | -1.9341 | 0.75673
| 270 |  | 12.3123 | 12.0775 | 0.016819 | -1.9073 | 0.75856 | 270 | -1.9073 | 0.75856
| 285 |  | 12.3145 | 12.071 | 0.011653 | -1.9776 | 0.75719 | 285 | -1.9776 | 0.75719
| 300 |  | 12.324 | 12.0525 | 0.015517 | -2.2033 | 0.75753 | 300 | -2.2033 | 0.75753
| 315 |  | 12.3238 | 12.042 | 0.015079 | -2.2869 | 0.75745 | 315 | -2.2869 | 0.75745
| 330 |  | 12.3217 | 12.0215 | 0.0074516 | -2.4366 | 0.7562 | 330 | -2.4366 | 0.7562
| 345 |  | 12.3225 | 11.998 | 0.013219 | -2.6337 | 0.75748 | 345 | -2.6337 | 0.75748
| 360 |  | 12.3309 | 11.9615 | 0.011367 | -2.996 | 0.75682 | 360 | -2.996 | 0.75682
| 375 |  | 12.3338 | 11.916 | 0.013139 | -3.3873 | 0.75781 | 375 | -3.3873 | 0.75781
| 390 |  | 12.3322 | 11.852 | 0.016416 | -3.8938 | 0.7607 | 390
| 405 | [6 empty] | 405
420
435
450
465
480
495
510
525
45
60
75
90
105
CETIAT | 120 |  | 11.9585 | 11.49 | 0.03 | -3.9179 | 0.55972 | 120 | -3.9179 | 0.55972
| 135 |  | 11.9559 | 11.54 | 0.04 | -3.4788 | 0.56598 | 135 | -3.4788 | 0.56598
| 150 |  | 11.9567 | 11.6 | 0.04 | -2.9833 | 0.56876 | 150 | -2.9833 | 0.56876
| 165 |  | 11.9615 | 11.64 | 0.04 | -2.6878 | 0.57021 | 165 | -2.6878 | 0.57021
| 180 |  | 11.9615 | 11.65 | 0.06 | -2.6039 | 0.58081 | 180 | -2.6039 | 0.58081
| 195 |  | 11.9639 | 11.67 | 0.05 | -2.4565 | 0.57598 | 195 | -2.4565 | 0.57598
| 210 |  | 11.9809 | 11.69 | 0.05 | -2.4277 | 0.57533 | 210 | -2.4277 | 0.57533
| 225 |  | 11.9815 | 11.7 | 0.05 | -2.3498 | 0.57574 | 225 | -2.3498 | 0.57574
| 240 |  | 11.9842 | 11.7 | 0.06 | -2.3711 | 0.58101 | 240 | -2.3711 | 0.58101
| 255 |  | 11.9743 | 11.7 | 0.03 | -2.2905 | 0.5683 | 255 | -2.2905 | 0.5683
| 270 | 11.9695 | 11.9705 | 11.72 | 0.046667 | -2.0928 | 0.57609 | 270 | -2.0928 | 0.57609
| 285 |  | 11.9712 | 11.72 | 0.02 | -2.0983 | 0.56699 | 285 | -2.0983 | 0.56699
| 300 |  | 11.9588 | 11.72 | 0.04 | -1.9971 | 0.57427 | 300 | -1.9971 | 0.57427
| 315 |  | 11.954 | 11.72 | 0.05 | -1.9574 | 0.57928 | 315 | -1.9574 | 0.57928
| 330 |  | 11.9585 | 11.69 | 0.04 | -2.2451 | 0.57288 | 330 | -2.2451 | 0.57288
| 345 |  | 11.9639 | 11.68 | 0.05 | -2.3729 | 0.57646 | 345 | -2.3729 | 0.57646
| 360 |  | 11.9615 | 11.68 | 0.05 | -2.3531 | 0.57669 | 360 | -2.3531 | 0.57669
| 375 |  | 11.9615 | 11.67 | 0.04 | -2.437 | 0.57164 | 375 | -2.437 | 0.57164
| 390 |  | 11.9567 | 11.69 | 0.05 | -2.2306 | 0.57761 | 390 | -2.2306 | 0.57761
| 405 |  | 11.9559 | 11.7 | 0.05 | -2.1406 | 0.57815 | 405 | -2.1406 | 0.57815
| 420 |  | 11.9585 | 11.7 | 0.05 | -2.1618 | 0.57791 | 420 | -2.1618 | 0.57791
435
450
465
480
495
510
525
45
DTI | [7 empty] | 60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
DWG | 50 | [6 empty] | 60
| 65 | [6 empty] | 75
| 80 |  | 12.097 | 11.5017 | 0.04129 | -4.9211 | 0.55451 | 90 | -4.1357 | 0.55481
| 95 |  | 12.081 | 11.6288 | 0.036717 | -3.7431 | 0.55497 | 105 | -3.1948 | 0.55533
| 110 |  | 12.066 | 11.7136 | 0.02093 | -2.9206 | 0.55551 | 120 | -2.5714 | 0.5557
| 125 |  | 12.058 | 11.769 | 0.017293 | -2.3967 | 0.5558 | 135 | -2.2291 | 0.55575
| 140 |  | 12.059 | 11.8003 | 0.014727 | -2.1453 | 0.55573 | 150 | -1.9603 | 0.55592
| 155 |  | 12.052 | 11.8269 | 0.017463 | -1.8677 | 0.55602 | 165 | -1.8508 | 0.55592
| 170 |  | 12.055 | 11.8329 | 0.012764 | -1.8424 | 0.55588 | 180 | -1.8213 | 0.55584
| 185 |  | 12.056 | 11.8377 | 0.0095002 | -1.8107 | 0.55583 | 195 | -1.8003 | 0.55601
| 200 |  | 12.05 | 11.8337 | 0.012835 | -1.795 | 0.5561 | 210 | -1.79 | 0.5561
| 215 |  | 12.05 | 11.8346 | 0.013667 | -1.7876 | 0.5561 | 225 | -1.7816 | 0.55613
| 230 |  | 12.049 | 11.8347 | 0.012292 | -1.7786 | 0.55615 | 240 | -1.7937 | 0.55621
| 245 |  | 12.047 | 11.83 | 0.011547 | -1.8013 | 0.55624 | 255 | -1.8273 | 0.55625
| 260 |  | 12.047 | 11.8253 | 0.013397 | -1.8403 | 0.55625 | 270 | -1.8128 | 0.55631
| 275 |  | 12.045 | 11.8283 | 0.0098575 | -1.7991 | 0.55633 | 285 | -1.8088 | 0.55631
| 290 |  | 12.0458 | 11.8273 | 0.011059 | -1.8137 | 0.5563 | 300 | -1.8311 | 0.55625
| 305 |  | 12.0474 | 11.8258 | 0.011457 | -1.8398 | 0.55623 | 315 | -1.8481 | 0.55623
| 320 |  | 12.0474 | 11.8243 | 0.015905 | -1.8522 | 0.55623 | 330 | -1.8603 | 0.55621
| 335 |  | 12.0481 | 11.8235 | 0.014452 | -1.8643 | 0.5562 | 345 | -1.8734 | 0.55611
| 350 |  | 12.0511 | 11.8248 | 0.013816 | -1.8779 | 0.55606 | 360 | -1.9006 | 0.55606
| 365 |  | 12.0511 | 11.8207 | 0.01047 | -1.9119 | 0.55606 | 375 | -1.8984 | 0.55604
| 380 |  | 12.0518 | 11.8238 | 0.012737 | -1.8917 | 0.55603 | 390 | -1.9059 | 0.55604
| 395 |  | 12.0514 | 11.8209 | 0.01199 | -1.913 | 0.55605 | 405 | -1.9539 | 0.55605
| 410 |  | 12.0514 | 11.8135 | 0.013282 | -1.9744 | 0.55606 | 420 | -1.9746 | 0.55604
| 425 |  | 12.0518 | 11.8138 | 0.012538 | -1.9746 | 0.55604 | 435 | -1.9754 | 0.5561
| 440 |  | 12.0498 | 11.8117 | 0.014173 | -1.9758 | 0.55613 | 450 | -1.9433 | 0.55612
| 455 |  | 12.0501 | 11.8179 | 0.011228 | -1.927 | 0.55611 | 465 | -1.9388 | 0.5561
| 470 |  | 12.0504 | 11.8161 | 0.012347 | -1.9447 | 0.5561 | 480 | -1.9694 | 0.55611
| 485 |  | 12.0501 | 11.8113 | 0.011605 | -1.9818 | 0.55612 | 495 | -1.9841 | 0.55619
| 500 | 12.048 | 12.0477 | 11.8085 | 0.014542 | -1.9852 | 0.55623 | 510 | -2.004 | 0.55616
| 515 |  | 12.0501 | 11.8075 | 0.011471 | -2.0134 | 0.55612 | 525 | -2.0009 | 0.55614
| 530 |  | 12.0494 | 11.8091 | 0.011133 | -1.9947 | 0.55615
| 545 |  | 12.0477 | 11.8071 | 0.012467 | -1.9968 | 0.55623
UL-LMPS | [7 empty] | 45
60
75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
45
60
SHMU | [7 empty] | 75
90
105
120
135
150
165
180
195
210
225
240
255
270
285
300
315
330
345
360
375
390
405
420
435
450
465
480
495
510
525
| 45 |  | 12.031 | 11.311 | 0.013 | -5.985 | 0.52 | 45
| 60 |  | 12.025 | 11.443 | 0.013 | -4.84 | 0.52 | 60
LEI | 75 |  | 12.02 | 11.539 | 0.012 | -4.002 | 0.52 | 75 | -4.002 | 0.52
| 90 |  | 12.017 | 11.619 | 0.015 | -3.312 | 0.52 | 90 | -3.312 | 0.52
| 105 |  | 12.016 | 11.691 | 0.015 | -2.705 | 0.52 | 105 | -2.705 | 0.52
| 120 |  | 12.008 | 11.753 | 0.015 | -2.124 | 0.52 | 120 | -2.124 | 0.52
| 135 |  | 12.004 | 11.784 | 0.014 | -1.833 | 0.52 | 135 | -1.833 | 0.52
| 150 |  | 11.995 | 11.786 | 0.016 | -1.742 | 0.52 | 150 | -1.742 | 0.52
| 165 |  | 11.993 | 11.792 | 0.014 | -1.676 | 0.52 | 165 | -1.676 | 0.52
| 180 |  | 11.982 | 11.812 | 0.013 | -1.419 | 0.52 | 180 | -1.419 | 0.52
| 195 |  | 11.988 | 11.808 | 0.015 | -1.502 | 0.52 | 195 | -1.502 | 0.52
| 210 |  | 11.988 | 11.804 | 0.011 | -1.535 | 0.52 | 210 | -1.535 | 0.52
| 225 |  | 11.981 | 11.8 | 0.012 | -1.511 | 0.52 | 225 | -1.511 | 0.52
| 240 |  | 11.98 | 11.797 | 0.016 | -1.528 | 0.52 | 240 | -1.528 | 0.52
| 255 |  | 11.972 | 11.788 | 0.015 | -1.537 | 0.52 | 255 | -1.537 | 0.52
| 270 |  | 11.977 | 11.773 | 0.016 | -1.703 | 0.52 | 270 | -1.703 | 0.52
| 285 |  | 11.971 | 11.739 | 0.019 | -1.938 | 0.52 | 285 | -1.938 | 0.52
| 300 |  | 11.964 | 11.672 | 0.022 | -2.441 | 0.52 | 300 | -2.441 | 0.52
| 315 |  | 11.967 | 11.503 | 0.015 | -3.877 | 0.52 | 315 | -3.877 | 0.52
| 320 |  | 11.963 | 11.408 | 0.015 | -4.639 | 0.52 | 330
| 345 | [6 empty] | 345
| 360 | [6 empty] | 360
375
390
405
420
435
450
465
480
495
510
525
Note on M1: jgers: Uncertainty component related to meter stability during the measurement campaign - determined from repeated calibrations at CMI

## v dependencies
| Schiltknecht MINI | Schiltknecht MACRO | Testo 0635 9540 (small probe, 16 mm) | Testo 0635 9340 (large probe, 107 mm)
BEV/E+E
PTB
CMI
CETIAT
DTI
DWG
UL-LMPS
SHMU
LEI

## stability
| Schiltknecht MINI | [6 empty] | Schiltknecht MACRO | [6 empty] | Testo 0635 9540 (small probe, 16 mm) | [6 empty] | Testo 0635 9340 (large probe, 107 mm)
| vREF | vMUT | E | U(E) | E | U(E) | calibration | vREF | vMUT | E | U(E) | E | U(E) | calibration | vREF | vMUT | E | U(E) | E | U(E) | calibration | vREF | vMUT | E | U(E) | E | U(E) | calibration
| m/s | m/s | m/s | m/s | % | % | date | m/s | m/s | m/s | m/s | % | % | date | m/s | m/s | m/s | m/s | % | % | date | m/s | m/s | m/s | m/s | % | % | date
Aug-Sep 2022 | 0.54 | 0.53 | -0.01 | 0.012 | -1.89 | 2.15 | 2022-09-07 | 0.54 | 0.423 | -0.117 | 0.012 | -21.74 | 2.15 | 2022-09-07 | 0.533 | 0.642 | 0.109 | 0.012 | 20.44 | 2.18 | 2022-08-19 | 0.541 | 0.51 | -0.031 | 0.012 | -5.75 | 2.15 | 2022-09-09
| 2.005 | 2.007 | 0.002 | 0.016 | 0.09 | 0.8 |  | 2 | 1.889 | -0.112 | 0.016 | -5.58 | 0.8 |  | 1.997 | 2.089 | 0.092 | 0.016 | 4.61 | 0.8 |  | 1.986 | 1.938 | -0.048 | 0.016 | -2.43 | 0.81
| 4.975 | 5.008 | 0.033 | 0.025 | 0.66 | 0.5 |  | 5.059 | 4.898 | -0.16 | 0.026 | -3.17 | 0.51 |  | 5.066 | 5.164 | 0.098 | 0.025 | 1.93 | 0.5 |  | 5.057 | 4.973 | -0.084 | 0.038 | -1.67 | 0.75
| 8.036 | 8.056 | 0.019 | 0.04 | 0.24 | 0.5 |  | 8.061 | 7.946 | -0.114 | 0.04 | -1.42 | 0.5 |  | 8.126 | 8.168 | 0.042 | 0.041 | 0.52 | 0.5 |  | 8.033 | 7.888 | -0.145 | 0.06 | -1.8 | 0.74
| 12.066 | 12.055 | -0.011 | 0.06 | -0.09 | 0.5 |  | 12.053 | 11.943 | -0.11 | 0.06 | -0.91 | 0.5 |  | 12.006 | 11.953 | -0.053 | 0.06 | -0.44 | 0.5 |  | 12.216 | 12.027 | -0.188 | 0.09 | -1.54 | 0.74
| 20.054 | 19.779 | -0.274 | 0.1 | -1.37 | 0.5 |  | 20.091 | 19.797 | -0.294 | 0.1 | -1.46 | 0.5 |  | 20.063 | 19.775 | -0.288 | 0.1 | -1.44 | 0.5
Jun 2023 | 0.547 | 0.546 | -0.001 | 0.012 | -0.23 | 2.13 | 2023-06-09 | 0.549 | 0.53 | -0.019 | 0.012 | -3.41 | 2.12 | 2023-06-09 | 0.547 | 0.677 | 0.13 | 0.012 | 23.7 | 2.13 | 2023-06-15 | 0.544 | 0.54 | -0.004 | 0.012 | -0.75 | 2.14 | 2023-06-20
| 2.038 | 2.042 | 0.004 | 0.016 | 0.2 | 0.79 |  | 2.093 | 1.986 | -0.107 | 0.016 | -5.12 | 0.78 |  | 2.006 | 2.125 | 0.118 | 0.016 | 5.89 | 0.8 |  | 1.981 | 1.947 | -0.034 | 0.016 | -1.71 | 0.82
| 4.997 | 5.033 | 0.036 | 0.025 | 0.72 | 0.5 |  | 5.177 | 5.033 | -0.144 | 0.026 | -2.78 | 0.5 |  | 5.09 | 5.201 | 0.111 | 0.025 | 2.18 | 0.5 |  | 5.081 | 5.004 | -0.077 | 0.038 | -1.52 | 0.75
| 8.075 | 8.101 | 0.026 | 0.04 | 0.32 | 0.5 |  | 8.161 | 8.051 | -0.11 | 0.041 | -1.34 | 0.5 |  | 8.074 | 8.111 | 0.037 | 0.04 | 0.46 | 0.5 |  | 8.15 | 8.003 | -0.147 | 0.06 | -1.81 | 0.74
| 12.041 | 12.05 | 0.009 | 0.06 | 0.08 | 0.5 |  | 12.103 | 11.982 | -0.121 | 0.06 | -1 | 0.5 |  | 12.119 | 12.048 | -0.07 | 0.061 | -0.58 | 0.5 |  | 12.174 | 11.981 | -0.193 | 0.09 | -1.59 | 0.74
| 20.318 | 20.023 | -0.295 | 0.102 | -1.45 | 0.5 |  | 20.214 | 19.907 | -0.307 | 0.101 | -1.52 | 0.5 |  | 20.142 | 19.836 | -0.306 | 0.101 | -1.52 | 0.5
Dec 2024 | 0.547 | 0.554 | 0.006 | 0.012 | 1.15 | 2.13 | 2024-12-04 | 0.546 | 0.531 | -0.015 | 0.012 | -2.82 | 2.13 | 2024-12-05 | 0.544 | 0.612 | 0.068 | 0.012 | 12.52 | 2.14 | 2024-12-09 | 0.543 | 0.52 | -0.023 | 0.012 | -4.23 | 2.14 | 2024-12-09
| 2.039 | 2.044 | 0.005 | 0.016 | 0.24 | 0.79 |  | 2.009 | 1.908 | -0.1 | 0.016 | -4.99 | 0.8 |  | 2.007 | 2.049 | 0.042 | 0.016 | 2.08 | 0.8 |  | 2.003 | 1.968 | -0.035 | 0.016 | -1.74 | 0.8
| 5.071 | 5.105 | 0.034 | 0.025 | 0.67 | 0.5 |  | 5.078 | 4.946 | -0.132 | 0.026 | -2.6 | 0.51 |  | 5.084 | 5.169 | 0.085 | 0.025 | 1.67 | 0.5 |  | 5.061 | 4.983 | -0.078 | 0.038 | -1.54 | 0.75
| 8.079 | 8.091 | 0.012 | 0.04 | 0.15 | 0.5 |  | 8.095 | 7.983 | -0.112 | 0.04 | -1.38 | 0.5 |  | 8.077 | 8.091 | 0.014 | 0.04 | 0.18 | 0.5 |  | 8.085 | 7.927 | -0.158 | 0.06 | -1.95 | 0.75
| 12.029 | 11.996 | -0.034 | 0.06 | -0.28 | 0.5 |  | 12.026 | 11.928 | -0.098 | 0.06 | -0.81 | 0.5 |  | 12.033 | 11.961 | -0.072 | 0.06 | -0.6 | 0.5 |  | 12.096 | 11.883 | -0.214 | 0.09 | -1.77 | 0.74
| 20.006 | 19.623 | -0.383 | 0.1 | -1.92 | 0.5 |  | 20.091 | 19.793 | -0.298 | 0.1 | -1.48 | 0.5 |  | 20.052 | 19.846 | -0.205 | 0.1 | -1.02 | 0.5
Aug 2025 | 0.545 | 0.538 | -0.007 | 0.012 | -1.29 | 2.14 | 2025-08-19 | 0.542 | 0.418 | -0.124 | 0.012 | -22.85 | 2.15 | 2025-08-19 | 0.543 | 0.661 | 0.118 | 0.012 | 21.71 | 2.14 | 2025-08-20
| 2 | 1.991 | -0.009 | 0.016 | -0.43 | 0.8 |  | 1.996 | 1.882 | -0.114 | 0.016 | -5.71 | 0.8 |  | 1.996 | 2.106 | 0.11 | 0.016 | 5.53 | 0.8
| 5.081 | 5.104 | 0.022 | 0.025 | 0.44 | 0.5 |  | 5.071 | 4.914 | -0.156 | 0.026 | -3.08 | 0.51 |  | 5.079 | 5.197 | 0.118 | 0.025 | 2.32 | 0.5
| 8.07 | 8.061 | -0.009 | 0.04 | -0.11 | 0.5 |  | 8.082 | 7.959 | -0.123 | 0.04 | -1.52 | 0.5 |  | 8.082 | 8.134 | 0.052 | 0.04 | 0.64 | 0.5
| 12.029 | 11.976 | -0.053 | 0.06 | -0.44 | 0.5 |  | 12.004 | 11.875 | -0.129 | 0.06 | -1.07 | 0.5 |  | 12.02 | 11.941 | -0.079 | 0.06 | -0.66 | 0.5
| 20.055 | 19.684 | -0.371 | 0.1 | -1.85 | 0.5 |  | 20.037 | 19.718 | -0.319 | 0.1 | -1.59 | 0.5 |  | 20.053 | 19.751 | -0.302 | 0.1 | -1.51 | 0.5
| vNOM | uDRIFT | U_st |  |  |  |  | vNOM | uDRIFT | U_st |  |  |  |  | vNOM | uDRIFT | U_st |  |  |  |  | vNOM | uDRIFT | U_st
| m/s | % | % |  |  |  |  | m/s | % | % |  |  |  |  | m/s | % | % |  |  |  |  | m/s | % | %
| 0.5 | 0.87757 | 2.4821 |  |  |  |  | 0.5 | 5.7822 | 16.3544 |  |  |  |  | 0.6 | 3.2274 | 9.1284 |  |  |  |  | 0.5 | 1.4434 | 4.0825
| 2 | 0.19341 | 0.54705 |  |  |  |  | 2 | 0.20785 | 0.58788 |  |  |  |  | 2 | 1.0999 | 3.1109 |  |  |  |  | 2 | 0.20785 | 0.58788
| 5 | 0.080829 | 0.22862 |  |  |  |  | 5 | 0.16454 | 0.4654 |  |  |  |  | 5 | 0.18764 | 0.53072 |  |  |  |  | 5 | 0.043301 | 0.12247
| 8 | 0.12413 | 0.35109 |  |  |  |  | 8 | 0.051962 | 0.14697 |  |  |  |  | 8 | 0.13279 | 0.37559 |  |  |  |  | 8 | 0.043301 | 0.12247
| 12 | 0.15011 | 0.42458 |  |  |  |  | 12 | 0.075056 | 0.21229 |  |  |  |  | 12 | 0.063509 | 0.17963 |  |  |  |  | 12 | 0.066395 | 0.18779
| 20 | 0.15877 | 0.44907 |  |  |  |  | 20 | 0.037528 | 0.10614 |  |  |  |  | 20 | 0.14434 | 0.40825

## DWG - Vane anemometers
| TSR MINI, 8 m/s | [6 empty] | TSR MACRO, 8 m/s | [6 empty] | RM Young Gill Propeller 27106DT, 8 m/s
| d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d))
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%)
DWG wall | 140 | 8.1784 | 8.1696 | 8.5201 | 0.051327 | 4.2912 | 0.52265
| 155 | 8.1787 | 8.1696 | 8.3619 | 0.048659 | 2.3541 | 0.51292 | 155 | 8.1801 | 8.1721 | 8.3091 | 0.018875 | 1.6768 | 0.50897
| 170 | 8.1784 | 8.1696 | 8.3037 | 0.031812 | 1.6422 | 0.50901 | 170 | 8.1812 | 8.1721 | 8.2638 | 0.018818 | 1.1222 | 0.50626 | 170 | 8.1792 | 8.1652 | 8.1761 | 0.0070965 | 0.1334 | 0.50159
| 185 | 8.1796 | 8.1696 | 8.2937 | 0.031481 | 1.5194 | 0.50847 | 185 | 8.1835 | 8.1721 | 8.2591 | 0.016576 | 1.0642 | 0.50609 | 185 | 8.1805 | 8.1652 | 8.1769 | 0.0060834 | 0.14345 | 0.5017
| 200 | 8.1809 | 8.1696 | 8.2847 | 0.029216 | 1.4096 | 0.50796 | 200 | 8.1815 | 8.1721 | 8.2487 | 0.020172 | 0.93726 | 0.50537 | 200 | 8.1819 | 8.1652 | 8.1764 | 0.007136 | 0.13745 | 0.50178
| 215 | 8.1786 | 8.1696 | 8.2339 | 0.027119 | 0.78774 | 0.50468 | 215 | 8.1824 | 8.1721 | 8.2588 | 0.017433 | 1.0612 | 0.50602 | 215 | 8.1807 | 8.1652 | 8.1747 | 0.0071923 | 0.11604 | 0.50159
| 230 | 8.1779 | 8.1696 | 8.2296 | 0.024298 | 0.73518 | 0.50433 | 230 | 8.1856 | 8.1721 | 8.2261 | 0.017732 | 0.66072 | 0.50421 | 230 | 8.179 | 8.1652 | 8.1725 | 0.0061248 | 0.089421 | 0.50134
| 245 | 8.1767 | 8.1696 | 8.2118 | 0.022693 | 0.51649 | 0.50315 | 245 | 8.1823 | 8.1721 | 8.2145 | 0.012622 | 0.51907 | 0.50326 | 245 | 8.1784 | 8.1652 | 8.1702 | 0.00661 | 0.061238 | 0.50117
| 260 | 8.1775 | 8.1696 | 8.1993 | 0.024455 | 0.36355 | 0.50245 | 260 | 8.1808 | 8.1721 | 8.205 | 0.012891 | 0.40324 | 0.50259 | 260 | 8.1777 | 8.1652 | 8.1693 | 0.0063596 | 0.050099 | 0.50107
| 275 | 8.1761 | 8.1696 | 8.1868 | 0.028655 | 0.21147 | 0.50166 | 275 | 8.1812 | 8.1721 | 8.1986 | 0.013111 | 0.32467 | 0.50222 | 275 | 8.1787 | 8.1652 | 8.1703 | 0.0066351 | 0.062953 | 0.5012
| 290 | 8.1764 | 8.1696 | 8.1789 | 0.025577 | 0.11454 | 0.50115 | 290 | 8.1797 | 8.1721 | 8.1909 | 0.011071 | 0.23064 | 0.50165 | 290 | 8.1782 | 8.1652 | 8.1697 | 0.0059964 | 0.055167 | 0.50112
| 305 | 8.1763 | 8.1696 | 8.1761 | 0.024702 | 0.079643 | 0.50096 | 305 | 8.1822 | 8.1721 | 8.19 | 0.011159 | 0.21935 | 0.50175 | 305 | 8.1761 | 8.1652 | 8.169 | 0.0061507 | 0.046183 | 0.50095
| 320 | 8.1772 | 8.1696 | 8.1786 | 0.026362 | 0.11043 | 0.50119 | 320 | 8.1804 | 8.1721 | 8.1846 | 0.010647 | 0.15307 | 0.5013 | 320 | 8.1757 | 8.1652 | 8.1688 | 0.0073335 | 0.044519 | 0.50093
| 335 | 8.1764 | 8.1696 | 8.1733 | 0.025085 | 0.045614 | 0.5008 | 335 | 8.1813 | 8.1721 | 8.1853 | 0.011518 | 0.16178 | 0.5014 | 335 | 8.1769 | 8.1652 | 8.1718 | 0.0057797 | 0.080317 | 0.50116
| 350 | 8.1747 | 8.1696 | 8.1711 | 0.025345 | 0.018913 | 0.50057 | 350 | 8.1788 | 8.1721 | 8.182 | 0.010442 | 0.12061 | 0.50104 | 350 | 8.1754 | 8.1652 | 8.1711 | 0.0054379 | 0.072409 | 0.50102
| 365 | 8.1761 | 8.1696 | 8.1753 | 0.026448 | 0.070381 | 0.50093 | 365 | 8.1774 | 8.1721 | 8.1797 | 0.011175 | 0.093042 | 0.50082 | 365 | 8.1746 | 8.1652 | 8.1692 | 0.00629 | 0.048875 | 0.50087
| 380 | 8.1755 | 8.1696 | 8.174 | 0.026609 | 0.054657 | 0.50082 | 380 | 8.1789 | 8.1721 | 8.1768 | 0.01045 | 0.057892 | 0.50073 | 380 | 8.1741 | 8.1652 | 8.1695 | 0.0072021 | 0.052834 | 0.50087
| 395 | 8.1747 | 8.1696 | 8.1686 | 0.02732 | -0.011747 | 0.50044 | 395 | 8.178 | 8.1721 | 8.178 | 0.011343 | 0.072809 | 0.50076 | 395 | 8.1756 | 8.1652 | 8.1715 | 0.0066344 | 0.076901 | 0.50107
| 410 | 8.1738 | 8.1696 | 8.1725 | 0.024884 | 0.036009 | 0.50059 | 410 | 8.1776 | 8.1721 | 8.1759 | 0.011062 | 0.046936 | 0.5006 | 410 | 8.1711 | 8.1652 | 8.1667 | 0.005819 | 0.017759 | 0.50049
| 425 | 8.1759 | 8.1696 | 8.1772 | 0.02608 | 0.093536 | 0.50102 | 425 | 8.1773 | 8.1721 | 8.1756 | 0.0093754 | 0.042291 | 0.50055 | 425 | 8.1715 | 8.1652 | 8.1674 | 0.0059084 | 0.026869 | 0.50056
| 440 | 8.1729 | 8.1696 | 8.1776 | 0.02843 | 0.098672 | 0.5009 | 440 | 8.1785 | 8.1721 | 8.179 | 0.01098 | 0.084833 | 0.50084 | 440 | 8.1722 | 8.1652 | 8.1681 | 0.0056368 | 0.035522 | 0.50064
| 455 | 8.1747 | 8.1696 | 8.1703 | 0.027224 | 0.0093727 | 0.50055 | 455 | 8.1773 | 8.1721 | 8.1757 | 0.010851 | 0.044664 | 0.50057 | 455 | 8.1694 | 8.1652 | 8.1649 | 0.005676 | -0.0041296 | 0.50028
| 470 | 8.1717 | 8.1696 | 8.1727 | 0.025687 | 0.038497 | 0.50049 | 470 | 8.174 | 8.1721 | 8.1737 | 0.011054 | 0.019199 | 0.50024 | 470 | 8.1682 | 8.1652 | 8.1647 | 0.006056 | -0.0055852 | 0.5002
| 485 | 8.1717 | 8.1696 | 8.17 | 0.025616 | 0.0057743 | 0.50032 | 485 | 8.176 | 8.1721 | 8.1758 | 0.011177 | 0.045106 | 0.5005 | 485 | 8.1682 | 8.1652 | 8.1664 | 0.0062185 | 0.014581 | 0.5003
| 500 | 8.1741 | 8.1696 | 8.1717 | 0.028298 | 0.025686 | 0.50061 | 500 | 8.1754 | 8.1721 | 8.1741 | 0.0099707 | 0.02488 | 0.50035 | 500 | 8.167 | 8.1652 | 8.1653 | 0.0064227 | 0.0016965 | 0.50017
| 515 | 8.1686 | 8.1696 | 8.1692 | 0.028007 | -0.0048177 | 0.50011 | 515 | 8.1729 | 8.1721 | 8.1701 | 0.0084101 | -0.024936 | 0.49994 | 515 | 8.166 | 8.1652 | 8.1644 | 0.0059318 | -0.010403 | 0.50004
| 530 | 8.1712 | 8.1696 | 8.17 | 0.025399 | 0.0052747 | 0.50029 | 530 | 8.1723 | 8.1721 | 8.1717 | 0.010274 | -0.0044798 | 0.50002 | 530 | 8.1651 | 8.1652 | 8.1646 | 0.0055642 | -0.0074517 | 0.49999
| 545 | 8.168 | 8.1696 | 8.173 | 0.025552 | 0.042571 | 0.50028 | 545 | 8.1704 | 8.1721 | 8.169 | 0.009582 | -0.037885 | 0.49973 | 545 | 8.1652 | 8.1652 | 8.1647 | 0.006103 | -0.0057357 | 0.50001
| 560 | 8.1682 | 8.1696 | 8.1711 | 0.028333 | 0.018926 | 0.50021 | 560 | 8.1693 | 8.1721 | 8.1705 | 0.0090822 | -0.019504 | 0.49975 | 560 | 8.1631 | 8.1652 | 8.1641 | 0.0064118 | -0.013905 | 0.49985
| 575 | 8.1651 | 8.1696 | 8.1665 | 0.026567 | -0.03722 | 0.49972 | 575 | 8.1679 | 8.1721 | 8.1686 | 0.0093098 | -0.043124 | 0.49955 | 575 | 8.1623 | 8.1652 | 8.164 | 0.0060201 | -0.014986 | 0.49979
| 590 | 8.1654 | 8.1696 | 8.1705 | 0.027452 | 0.011853 | 0.49999 | 590 | 8.1694 | 8.1721 | 8.1723 | 0.011873 | 0.0027593 | 0.49988 | 590 | 8.1619 | 8.1652 | 8.1645 | 0.0050851 | -0.0082244 | 0.49979
| 605 | 8.1653 | 8.1696 | 8.1704 | 0.027735 | 0.010743 | 0.49999 | 605 | 8.1677 | 8.1721 | 8.17 | 0.0084811 | -0.026201 | 0.49962 | 605 | 8.1632 | 8.1652 | 8.1666 | 0.0068917 | 0.017441 | 0.50002
| 620 | 8.1625 | 8.1696 | 8.1645 | 0.024076 | -0.062316 | 0.4994 | 620 | 8.1657 | 8.1721 | 8.1698 | 0.0097603 | -0.028009 | 0.49949 | 620 | 8.1607 | 8.1652 | 8.1638 | 0.0063731 | -0.016957 | 0.49969
| 635 | 8.1636 | 8.1696 | 8.1686 | 0.023942 | -0.012139 | 0.49972 | 635 | 8.1645 | 8.1721 | 8.1691 | 0.010832 | -0.036216 | 0.49938 | 635 | 8.1579 | 8.1652 | 8.1629 | 0.0054067 | -0.027806 | 0.49945
| 650 | 8.1621 | 8.1696 | 8.1668 | 0.025617 | -0.034513 | 0.49954 | 650 | 8.1639 | 8.1721 | 8.1694 | 0.010995 | -0.033018 | 0.49936 | 650 | 8.1567 | 8.1652 | 8.1617 | 0.0069424 | -0.042665 | 0.49933
| 665 | 8.1617 | 8.1696 | 8.1718 | 0.023849 | 0.026987 | 0.4998 | 665 | 8.1644 | 8.1721 | 8.1724 | 0.011709 | 0.0037005 | 0.49958 | 665 | 8.1556 | 8.1652 | 8.1614 | 0.0066475 | -0.046673 | 0.49923
| 680 | 8.1589 | 8.1696 | 8.166 | 0.028141 | -0.04417 | 0.49933 | 680 | 8.165 | 8.1721 | 8.1729 | 0.010422 | 0.0094649 | 0.49964 | 680 | 8.1535 | 8.1652 | 8.1602 | 0.0055694 | -0.060756 | 0.49902
| 695 | 8.159 | 8.1696 | 8.1649 | 0.025679 | -0.056944 | 0.49923 | 695 | 8.162 | 8.1721 | 8.171 | 0.0098991 | -0.013483 | 0.49934 | 695 | 8.1539 | 8.1652 | 8.1621 | 0.0072473 | -0.038084 | 0.49919
| 710 | 8.1581 | 8.1696 | 8.164 | 0.025604 | -0.068055 | 0.49912 | 710 | 8.1621 | 8.1721 | 8.1734 | 0.0099184 | 0.01544 | 0.49949 | 710 | 8.1501 | 8.1652 | 8.1597 | 0.0069626 | -0.067608 | 0.4988
| 725 | 8.1594 | 8.1696 | 8.1685 | 0.025951 | -0.01258 | 0.49948 | 725 | 8.1633 | 8.1721 | 8.175 | 0.0087951 | 0.035788 | 0.49966 | 725 | 8.1496 | 8.1652 | 8.1611 | 0.0058385 | -0.050687 | 0.49884
| 740 | 8.1596 | 8.1696 | 8.173 | 0.0279 | 0.041991 | 0.49979 | 740 | 8.1632 | 8.1721 | 8.1735 | 0.010159 | 0.017387 | 0.49957 | 740 | 8.1492 | 8.1652 | 8.1619 | 0.005722 | -0.040416 | 0.49886
| 755 | 8.156 | 8.1696 | 8.1628 | 0.026394 | -0.082304 | 0.49893 | 755 | 8.1591 | 8.1721 | 8.1709 | 0.0096457 | -0.014252 | 0.49916 | 755 | 8.1476 | 8.1652 | 8.1611 | 0.0061396 | -0.050837 | 0.49872
| 770 | 8.1578 | 8.1696 | 8.1685 | 0.023975 | -0.01296 | 0.49936 | 770 | 8.1591 | 8.1721 | 8.1733 | 0.010178 | 0.014964 | 0.49931 | 770 | 8.1452 | 8.1652 | 8.1597 | 0.0055125 | -0.067202 | 0.49848
| 785 | 8.1561 | 8.1696 | 8.169 | 0.028677 | -0.0073502 | 0.49935 | 785 | 8.1591 | 8.1721 | 8.1775 | 0.011917 | 0.065933 | 0.49957 | 785 | 8.1429 | 8.1652 | 8.1591 | 0.0065356 | -0.074255 | 0.49832
| 800 | 8.156 | 8.1696 | 8.1675 | 0.026982 | -0.025052 | 0.49922 | 800 | 8.1612 | 8.1721 | 8.1805 | 0.010151 | 0.10338 | 0.49987 | 800 | 8.1434 | 8.1652 | 8.1597 | 0.0058943 | -0.067476 | 0.49837
| 815 | 8.1569 | 8.1696 | 8.1699 | 0.027395 | 0.0041759 | 0.49943 | 815 | 8.1609 | 8.1721 | 8.1802 | 0.010533 | 0.099275 | 0.49984 | 815 | 8.1427 | 8.1652 | 8.1583 | 0.0061196 | -0.084575 | 0.49825
[8 empty] | 830 | 8.1603 | 8.1721 | 8.184 | 0.012499 | 0.14591 | 0.50005 | 830 | 8.1427 | 8.1652 | 8.1499 | 0.0058353 | -0.18801 | 0.49773
DWG open | 140 | 8.1804 | 8.1731 | 8.4765 | 0.015034 | 3.712 | 0.51908
| 155 | 8.1808 | 8.1731 | 8.5067 | 0.014914 | 4.0807 | 0.52095 | 155 | 8.1807 | 8.171 | 8.3504 | 0.010426 | 2.1954 | 0.51161
| 170 | 8.1829 | 8.1731 | 8.5293 | 0.014088 | 4.3574 | 0.52246 | 170 | 8.179 | 8.171 | 8.3548 | 0.010172 | 2.2494 | 0.51177 | 170 | 8.1915 | 8.1738 | 8.1894 | 0.0088393 | 0.19017 | 0.50213
| 185 | 8.1805 | 8.1731 | 8.5422 | 0.016274 | 4.5153 | 0.52311 | 185 | 8.1813 | 8.171 | 8.3469 | 0.0099509 | 2.1527 | 0.51143 | 185 | 8.1926 | 8.1738 | 8.1923 | 0.0084599 | 0.22607 | 0.50237
| 200 | 8.1804 | 8.1731 | 8.5363 | 0.018551 | 4.4441 | 0.52277 | 200 | 8.1834 | 8.171 | 8.3337 | 0.010467 | 1.9917 | 0.51076 | 200 | 8.192 | 8.1738 | 8.1911 | 0.0079319 | 0.21142 | 0.50225
| 215 | 8.183 | 8.1731 | 8.5003 | 0.03234 | 4.0036 | 0.5209 | 215 | 8.1811 | 8.171 | 8.3116 | 0.010993 | 1.7205 | 0.50926 | 215 | 8.1899 | 8.1738 | 8.1897 | 0.0074068 | 0.19439 | 0.50203
| 230 | 8.1822 | 8.1731 | 8.4417 | 0.036588 | 3.2864 | 0.51733 | 230 | 8.1823 | 8.171 | 8.297 | 0.012074 | 1.5428 | 0.50845 | 230 | 8.189 | 8.1738 | 8.185 | 0.0081684 | 0.13713 | 0.5017
| 245 | 8.1836 | 8.1731 | 8.3693 | 0.032653 | 2.4 | 0.51291 | 245 | 8.1812 | 8.171 | 8.2782 | 0.014691 | 1.3126 | 0.50725 | 245 | 8.1905 | 8.1738 | 8.186 | 0.0079342 | 0.14944 | 0.50185
| 260 | 8.1833 | 8.1731 | 8.3305 | 0.029666 | 1.9259 | 0.51048 | 260 | 8.1819 | 8.171 | 8.2599 | 0.01491 | 1.0887 | 0.50618 | 260 | 8.1912 | 8.1738 | 8.1857 | 0.0070988 | 0.14571 | 0.50186
| 275 | 8.1808 | 8.1731 | 8.29 | 0.028129 | 1.4296 | 0.50782 | 275 | 8.1831 | 8.171 | 8.2485 | 0.016466 | 0.94815 | 0.50556 | 275 | 8.1911 | 8.1738 | 8.1854 | 0.0063163 | 0.14212 | 0.50182
| 290 | 8.1818 | 8.1731 | 8.2586 | 0.024852 | 1.0456 | 0.50592 | 290 | 8.1831 | 8.171 | 8.2385 | 0.014452 | 0.82579 | 0.50493 | 290 | 8.1886 | 8.1738 | 8.1825 | 0.0071809 | 0.1061 | 0.5015
| 305 | 8.1803 | 8.1731 | 8.2301 | 0.025479 | 0.69731 | 0.50409 | 305 | 8.1819 | 8.171 | 8.2279 | 0.013612 | 0.69619 | 0.5042 | 305 | 8.189 | 8.1738 | 8.1822 | 0.0070752 | 0.10268 | 0.50151
| 320 | 8.181 | 8.1731 | 8.2178 | 0.024407 | 0.54718 | 0.50337 | 320 | 8.1819 | 8.171 | 8.22 | 0.0146 | 0.60018 | 0.50373 | 320 | 8.1864 | 8.1738 | 8.1802 | 0.0071313 | 0.078145 | 0.50122
| 335 | 8.18 | 8.1731 | 8.2086 | 0.024808 | 0.43446 | 0.50274 | 335 | 8.1841 | 8.171 | 8.2122 | 0.013311 | 0.50484 | 0.50337 | 335 | 8.1902 | 8.1738 | 8.1838 | 0.007907 | 0.12224 | 0.50169
| 350 | 8.1799 | 8.1731 | 8.1971 | 0.024824 | 0.2934 | 0.50204 | 350 | 8.1803 | 8.171 | 8.2016 | 0.014311 | 0.37447 | 0.5025 | 350 | 8.184 | 8.1738 | 8.1782 | 0.0062251 | 0.053134 | 0.50094
| 365 | 8.1801 | 8.1731 | 8.1898 | 0.02695 | 0.2039 | 0.50163 | 365 | 8.1809 | 8.171 | 8.1947 | 0.013565 | 0.29073 | 0.50211 | 365 | 8.187 | 8.1738 | 8.1817 | 0.0070656 | 0.096744 | 0.50135
| 380 | 8.1806 | 8.1731 | 8.1871 | 0.025262 | 0.17049 | 0.50147 | 380 | 8.1799 | 8.171 | 8.1895 | 0.01044 | 0.22657 | 0.50171 | 380 | 8.184 | 8.1738 | 8.1783 | 0.0056486 | 0.054668 | 0.50094
| 395 | 8.1772 | 8.1731 | 8.179 | 0.027115 | 0.07236 | 0.5008 | 395 | 8.1793 | 8.171 | 8.1864 | 0.011286 | 0.18869 | 0.50149 | 395 | 8.1842 | 8.1738 | 8.1788 | 0.0072763 | 0.061456 | 0.50101
| 410 | 8.1759 | 8.1731 | 8.1747 | 0.023059 | 0.019682 | 0.50041 | 410 | 8.1777 | 8.171 | 8.183 | 0.013264 | 0.14742 | 0.50119 | 410 | 8.1822 | 8.1738 | 8.177 | 0.0062962 | 0.038526 | 0.50075
| 425 | 8.1769 | 8.1731 | 8.1768 | 0.025631 | 0.044763 | 0.50062 | 425 | 8.1782 | 8.171 | 8.1784 | 0.010221 | 0.091109 | 0.50093 | 425 | 8.18 | 8.1738 | 8.1747 | 0.0054838 | 0.011165 | 0.50047
| 440 | 8.1778 | 8.1731 | 8.1691 | 0.025556 | -0.04959 | 0.5002 | 440 | 8.1787 | 8.171 | 8.1794 | 0.011667 | 0.10251 | 0.50102 | 440 | 8.1816 | 8.1738 | 8.1766 | 0.0071393 | 0.034414 | 0.50071
| 455 | 8.1763 | 8.1731 | 8.1726 | 0.026113 | -0.0065158 | 0.50033 | 455 | 8.1748 | 8.171 | 8.177 | 0.011153 | 0.073821 | 0.50063 | 455 | 8.1791 | 8.1738 | 8.1745 | 0.006419 | 0.0089783 | 0.50042
| 470 | 8.1757 | 8.1731 | 8.1734 | 0.024648 | 0.0035383 | 0.50032 | 470 | 8.1751 | 8.171 | 8.1761 | 0.011073 | 0.06315 | 0.5006 | 470 | 8.178 | 8.1738 | 8.1733 | 0.0068651 | -0.0057319 | 0.50028
| 485 | 8.173 | 8.1731 | 8.1706 | 0.025243 | -0.03079 | 0.5 | 485 | 8.1761 | 8.171 | 8.1754 | 0.011612 | 0.053551 | 0.50061 | 485 | 8.1783 | 8.1738 | 8.1739 | 0.008006 | 0.00091915 | 0.50036
| 500 | 8.1748 | 8.1731 | 8.1714 | 0.024373 | -0.02115 | 0.50015 | 500 | 8.1744 | 8.171 | 8.1729 | 0.01134 | 0.023726 | 0.50036 | 500 | 8.1762 | 8.1738 | 8.1731 | 0.005642 | -0.0092081 | 0.50014
| 515 | 8.1715 | 8.1731 | 8.1724 | 0.025892 | -0.0091214 | 0.50002 | 515 | 8.1738 | 8.171 | 8.1725 | 0.0095397 | 0.018363 | 0.50029 | 515 | 8.1785 | 8.1738 | 8.1761 | 0.0066732 | 0.028335 | 0.50049
| 530 | 8.1733 | 8.1731 | 8.1724 | 0.02424 | -0.0089223 | 0.50012 | 530 | 8.1725 | 8.171 | 8.1712 | 0.010104 | 0.0026173 | 0.50013 | 530 | 8.1751 | 8.1738 | 8.1745 | 0.0067285 | 0.0082322 | 0.50018
| 545 | 8.1726 | 8.1731 | 8.1692 | 0.024677 | -0.048079 | 0.49988 | 545 | 8.1716 | 8.171 | 8.1704 | 0.010526 | -0.0067604 | 0.50003 | 545 | 8.1723 | 8.1738 | 8.172 | 0.0065216 | -0.021851 | 0.49985
| 560 | 8.1727 | 8.1731 | 8.1736 | 0.025116 | 0.0063369 | 0.50017 | 560 | 8.1724 | 8.171 | 8.1719 | 0.010369 | 0.011745 | 0.50017 | 560 | 8.1736 | 8.1738 | 8.1747 | 0.0068797 | 0.010566 | 0.5001
| 575 | 8.1723 | 8.1731 | 8.1718 | 0.02633 | -0.016421 | 0.50004 | 575 | 8.1701 | 8.171 | 8.1664 | 0.0091896 | -0.055814 | 0.49969 | 575 | 8.1721 | 8.1738 | 8.1745 | 0.0062096 | 0.0083616 | 0.49999
| 590 | 8.1683 | 8.1731 | 8.1679 | 0.025983 | -0.0639 | 0.49956 | 590 | 8.17 | 8.171 | 8.1671 | 0.0098943 | -0.047051 | 0.49973 | 590 | 8.1695 | 8.1738 | 8.1719 | 0.0065161 | -0.022853 | 0.49967
| 605 | 8.1692 | 8.1731 | 8.1685 | 0.025522 | -0.056952 | 0.49964 | 605 | 8.167 | 8.171 | 8.1659 | 0.0082887 | -0.062186 | 0.49946 | 605 | 8.1696 | 8.1738 | 8.1736 | 0.0069743 | -0.0029861 | 0.49979
| 620 | 8.1667 | 8.1731 | 8.1693 | 0.027481 | -0.046733 | 0.49956 | 620 | 8.1684 | 8.171 | 8.1697 | 0.011673 | -0.016075 | 0.49979 | 620 | 8.1716 | 8.1738 | 8.1755 | 0.00577 | 0.021083 | 0.50001
| 635 | 8.1686 | 8.1731 | 8.175 | 0.027542 | 0.023283 | 0.50003 | 635 | 8.1666 | 8.171 | 8.1671 | 0.0097491 | -0.047439 | 0.49952 | 635 | 8.1681 | 8.1738 | 8.1732 | 0.00597 | -0.0074928 | 0.49966
| 650 | 8.1692 | 8.1731 | 8.1695 | 0.025955 | -0.043892 | 0.49971 | 650 | 8.1664 | 8.171 | 8.1682 | 0.0088468 | -0.033662 | 0.49957 | 650 | 8.1668 | 8.1738 | 8.1724 | 0.006087 | -0.017469 | 0.49953
| 665 | 8.1682 | 8.1731 | 8.1755 | 0.025607 | 0.029152 | 0.50001 | 665 | 8.1642 | 8.171 | 8.1672 | 0.0098306 | -0.046169 | 0.49938 | 665 | 8.1665 | 8.1738 | 8.1736 | 0.0069347 | -0.0021035 | 0.4996
| 680 | 8.164 | 8.1731 | 8.1688 | 0.025787 | -0.052743 | 0.49934 | 680 | 8.1649 | 8.171 | 8.1696 | 0.012083 | -0.016475 | 0.49958 | 680 | 8.1665 | 8.1738 | 8.1754 | 0.0072274 | 0.01892 | 0.49971
| 695 | 8.1641 | 8.1731 | 8.1695 | 0.025077 | -0.044453 | 0.49938 | 695 | 8.1653 | 8.171 | 8.1718 | 0.01005 | 0.010191 | 0.49973 | 695 | 8.1621 | 8.1738 | 8.1727 | 0.0062169 | -0.013045 | 0.49927
| 710 | 8.1629 | 8.1731 | 8.169 | 0.026354 | -0.050858 | 0.49929 | 710 | 8.1638 | 8.171 | 8.1697 | 0.0098368 | -0.016 | 0.4995 | 710 | 8.1635 | 8.1738 | 8.1743 | 0.0059612 | 0.0059536 | 0.49945
| 725 | 8.1628 | 8.1731 | 8.1662 | 0.024311 | -0.085097 | 0.49909 | 725 | 8.1651 | 8.171 | 8.1717 | 0.0092585 | 0.0088691 | 0.49971 | 725 | 8.1591 | 8.1738 | 8.1725 | 0.0089504 | -0.015769 | 0.49912
| 740 | 8.1619 | 8.1731 | 8.1687 | 0.028491 | -0.054376 | 0.49924 | 740 | 8.1621 | 8.171 | 8.1709 | 0.0082722 | -0.0010538 | 0.49947 | 740 | 8.1597 | 8.1738 | 8.1742 | 0.0069132 | 0.0049296 | 0.49922
| 755 | 8.1622 | 8.1731 | 8.1672 | 0.024008 | -0.072706 | 0.49911 | 755 | 8.1606 | 8.171 | 8.1697 | 0.0091046 | -0.015632 | 0.4993 | 755 | 8.1588 | 8.1738 | 8.1751 | 0.0062704 | 0.015915 | 0.49921
| 770 | 8.1615 | 8.1731 | 8.1707 | 0.025373 | -0.029934 | 0.4993 | 770 | 8.1626 | 8.171 | 8.1735 | 0.01013 | 0.031195 | 0.49967 | 770 | 8.1577 | 8.1738 | 8.1754 | 0.006733 | 0.01951 | 0.49917
| 785 | 8.1604 | 8.1731 | 8.1737 | 0.025631 | 0.0073274 | 0.49942 | 785 | 8.162 | 8.171 | 8.1758 | 0.0098924 | 0.059382 | 0.49977 | 785 | 8.158 | 8.1738 | 8.177 | 0.0054974 | 0.038684 | 0.49926
| 800 | 8.1604 | 8.1731 | 8.1677 | 0.02434 | -0.066862 | 0.49904 | 800 | 8.1602 | 8.171 | 8.1753 | 0.0086258 | 0.05268 | 0.49962 | 800 | 8.1556 | 8.1738 | 8.1749 | 0.0058547 | 0.013727 | 0.499
| 815 | 8.1587 | 8.1731 | 8.1701 | 0.026499 | -0.037404 | 0.4991 | 815 | 8.1591 | 8.171 | 8.1769 | 0.0091733 | 0.072893 | 0.49966 | 815 | 8.1529 | 8.1738 | 8.1711 | 0.0068396 | -0.032802 | 0.49861
[8 empty] | 830 | 8.1598 | 8.171 | 8.1797 | 0.010604 | 0.10712 | 0.49988 | 830 | 8.1531 | 8.1738 | 8.1647 | 0.0058854 | -0.11198 | 0.49822

## DWG - Thies FCA
| Thies FCA with Rotor #1; 8 m/s | [6 empty] | Thies FCA with Rotor #2; 8 m/s | [6 empty] | Thies FCA with Rotor #3; 8 m/s
| d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d))
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%)
DWG wall | 260 | 8.2051 | 8.1855 | 8.0382 | 0.066754 | -1.7987 | 0.49331 | 260 | 8.1922 | 8.1708 | 8.1668 | 0.087673 | -0.049237 | 0.50298 | 260 | 8.1916 | 8.1734 | 8.1748 | 0.064015 | 0.016322 | 0.50221
| 275 | 8.2045 | 8.1855 | 8.0791 | 0.070285 | -1.2998 | 0.49589 | 275 | 8.1936 | 8.1708 | 8.1817 | 0.085197 | 0.13344 | 0.50386 | 275 | 8.1954 | 8.1734 | 8.1803 | 0.060377 | 0.083661 | 0.50267
| 290 | 8.2028 | 8.1855 | 8.0392 | 0.076543 | -1.7867 | 0.49358 | 290 | 8.1926 | 8.1708 | 8.1604 | 0.080791 | -0.12732 | 0.50233 | 290 | 8.191 | 8.1734 | 8.1743 | 0.064445 | 0.010717 | 0.50216
| 305 | 8.2047 | 8.1855 | 8.0277 | 0.075088 | -1.9278 | 0.49294 | 305 | 8.1906 | 8.1708 | 8.1887 | 0.098842 | 0.21948 | 0.50474 | 305 | 8.1915 | 8.1734 | 8.1979 | 0.065405 | 0.30009 | 0.50367
| 320 | 8.2036 | 8.1855 | 8.0454 | 0.078081 | -1.7107 | 0.49407 | 320 | 8.1871 | 8.1708 | 8.1577 | 0.082799 | -0.16046 | 0.50191 | 320 | 8.1923 | 8.1734 | 8.2002 | 0.062622 | 0.32793 | 0.50377
| 335 | 8.2021 | 8.1855 | 8.0621 | 0.074628 | -1.5075 | 0.49487 | 335 | 8.1877 | 8.1708 | 8.1587 | 0.094939 | -0.14831 | 0.50254 | 335 | 8.1895 | 8.1734 | 8.1985 | 0.074656 | 0.30661 | 0.5039
| 350 | 8.2014 | 8.1855 | 8.0606 | 0.074977 | -1.526 | 0.49474 | 350 | 8.1877 | 8.1708 | 8.1678 | 0.088082 | -0.037032 | 0.5028 | 350 | 8.1871 | 8.1734 | 8.1799 | 0.070379 | 0.078747 | 0.50247
| 365 | 8.2001 | 8.1855 | 8.0806 | 0.078494 | -1.2811 | 0.49602 | 365 | 8.1839 | 8.1708 | 8.1628 | 0.088707 | -0.097417 | 0.50227 | 365 | 8.1884 | 8.1734 | 8.1799 | 0.054131 | 0.07882 | 0.50204
| 380 | 8.1958 | 8.1855 | 8.0865 | 0.075912 | -1.2094 | 0.49602 | 380 | 8.1839 | 8.1708 | 8.1736 | 0.085182 | 0.034283 | 0.50278 | 380 | 8.1843 | 8.1734 | 8.1848 | 0.068576 | 0.13919 | 0.50253
| 395 | 8.1977 | 8.1855 | 8.1179 | 0.078313 | -0.82544 | 0.49816 | 395 | 8.18 | 8.1708 | 8.1585 | 0.089923 | -0.14976 | 0.50183 | 395 | 8.1856 | 8.1734 | 8.1656 | 0.069005 | -0.095623 | 0.50145
| 410 | 8.1965 | 8.1855 | 8.1328 | 0.081065 | -0.64387 | 0.49909 | 410 | 8.1821 | 8.1708 | 8.1725 | 0.080722 | 0.020657 | 0.50242 | 410 | 8.181 | 8.1734 | 8.1817 | 0.063595 | 0.10105 | 0.50198
| 425 | 8.1941 | 8.1855 | 8.1465 | 0.07336 | -0.47628 | 0.49949 | 425 | 8.1778 | 8.1708 | 8.1683 | 0.086929 | -0.030948 | 0.50217 | 425 | 8.1817 | 8.1734 | 8.1957 | 0.059476 | 0.27208 | 0.50274
| 440 | 8.1946 | 8.1855 | 8.1681 | 0.080704 | -0.21234 | 0.50111 | 440 | 8.178 | 8.1708 | 8.1698 | 0.094165 | -0.011627 | 0.50259 | 440 | 8.1789 | 8.1734 | 8.1941 | 0.064125 | 0.25278 | 0.50262
| 455 | 8.1922 | 8.1855 | 8.1746 | 0.078567 | -0.13337 | 0.50128 | 455 | 8.1758 | 8.1708 | 8.1655 | 0.086459 | -0.06494 | 0.50185 | 455 | 8.1791 | 8.1734 | 8.1792 | 0.057494 | 0.070699 | 0.50153
| 470 | 8.1886 | 8.1855 | 8.1746 | 0.076871 | -0.13294 | 0.501 | 470 | 8.1754 | 8.1708 | 8.1551 | 0.092144 | -0.19248 | 0.50144 | 470 | 8.1774 | 8.1734 | 8.1873 | 0.073171 | 0.16933 | 0.50242
| 485 | 8.1907 | 8.1855 | 8.1916 | 0.077725 | 0.075117 | 0.50219 | 485 | 8.1741 | 8.1708 | 8.181 | 0.09867 | 0.12561 | 0.50325 | 485 | 8.1773 | 8.1734 | 8.1889 | 0.063698 | 0.18963 | 0.5022
| 500 | 8.1871 | 8.1855 | 8.1869 | 0.076418 | 0.016872 | 0.50164 | 500 | 8.1734 | 8.1708 | 8.182 | 0.082001 | 0.13757 | 0.50252 | 500 | 8.1752 | 8.1734 | 8.1785 | 0.071968 | 0.06163 | 0.50171
| 515 | 8.1866 | 8.1855 | 8.1848 | 0.082803 | -0.0087325 | 0.50173 | 515 | 8.1726 | 8.1708 | 8.1794 | 0.10089 | 0.10529 | 0.50317 | 515 | 8.1737 | 8.1734 | 8.1667 | 0.067698 | -0.08268 | 0.50075
| 530 | 8.1839 | 8.1855 | 8.1959 | 0.078094 | 0.12702 | 0.50205 | 530 | 8.1692 | 8.1708 | 8.1663 | 0.088294 | -0.055385 | 0.50157 | 530 | 8.1732 | 8.1734 | 8.1739 | 0.076135 | 0.0056663 | 0.50146
| 545 | 8.1837 | 8.1855 | 8.1948 | 0.085336 | 0.11347 | 0.50227 | 545 | 8.17 | 8.1708 | 8.1839 | 0.085369 | 0.16008 | 0.50256 | 545 | 8.171 | 8.1734 | 8.1787 | 0.071038 | 0.064858 | 0.50144
| 560 | 8.1833 | 8.1855 | 8.1895 | 0.08061 | 0.049338 | 0.50173 | 560 | 8.1683 | 8.1708 | 8.1751 | 0.08806 | 0.053241 | 0.50204 | 560 | 8.1707 | 8.1734 | 8.1825 | 0.070355 | 0.11114 | 0.50162
| 575 | 8.1808 | 8.1855 | 8.1946 | 0.069711 | 0.11093 | 0.50147 | 575 | 8.1648 | 8.1708 | 8.1977 | 0.090528 | 0.32952 | 0.50332 | 575 | 8.1661 | 8.1734 | 8.173 | 0.06983 | -0.0045724 | 0.50074
| 590 | 8.1804 | 8.1855 | 8.1767 | 0.084328 | -0.10771 | 0.50094 | 590 | 8.1669 | 8.1708 | 8.1836 | 0.089639 | 0.15728 | 0.50255 | 590 | 8.17 | 8.1734 | 8.1599 | 0.064418 | -0.16489 | 0.5
| 605 | 8.178 | 8.1855 | 8.1763 | 0.078006 | -0.11222 | 0.5005 | 605 | 8.165 | 8.1708 | 8.1249 | 0.092067 | -0.56135 | 0.49897 | 605 | 8.1661 | 8.1734 | 8.1729 | 0.06752 | -0.0065488 | 0.50066
| 620 | 8.1765 | 8.1855 | 8.1835 | 0.07543 | -0.024361 | 0.50074 | 620 | 8.1635 | 8.1708 | 8.1671 | 0.089471 | -0.045185 | 0.50133 | 620 | 8.164 | 8.1734 | 8.1841 | 0.083844 | 0.13021 | 0.50183
| 635 | 8.1754 | 8.1855 | 8.196 | 0.070876 | 0.12831 | 0.50127 | 635 | 8.1642 | 8.1708 | 8.1887 | 0.099715 | 0.21884 | 0.50316 | 635 | 8.166 | 8.1734 | 8.1635 | 0.072531 | -0.12171 | 0.50025
| 650 | 8.1728 | 8.1855 | 8.178 | 0.074714 | -0.091696 | 0.50016 | 650 | 8.1617 | 8.1708 | 8.1416 | 0.07923 | -0.35764 | 0.49923 | 650 | 8.1626 | 8.1734 | 8.1847 | 0.063878 | 0.13755 | 0.50104
| 665 | 8.1735 | 8.1855 | 8.1883 | 0.083365 | 0.034785 | 0.50117 | 665 | 8.1626 | 8.1708 | 8.1615 | 0.094262 | -0.1132 | 0.50115 | 665 | 8.1626 | 8.1734 | 8.1848 | 0.066893 | 0.1395 | 0.50115
| 680 | 8.1737 | 8.1855 | 8.1848 | 0.072531 | -0.0076355 | 0.50055 | 680 | 8.1628 | 8.1708 | 8.1843 | 0.096449 | 0.16608 | 0.50265 | 680 | 8.1618 | 8.1734 | 8.1643 | 0.065477 | -0.11127 | 0.49981
| 695 | 8.1703 | 8.1855 | 8.1821 | 0.073162 | -0.040887 | 0.50021 | 695 | 8.1582 | 8.1708 | 8.1509 | 0.08604 | -0.24384 | 0.49987 | 695 | 8.1592 | 8.1734 | 8.1794 | 0.06463 | 0.073797 | 0.50054
| 710 | 8.1706 | 8.1855 | 8.195 | 0.077109 | 0.11685 | 0.50115 | 710 | 8.1588 | 8.1708 | 8.1727 | 0.096395 | 0.023114 | 0.5017 | 710 | 8.1577 | 8.1734 | 8.1768 | 0.064141 | 0.041212 | 0.50027
| 725 | 8.1686 | 8.1855 | 8.201 | 0.074508 | 0.19004 | 0.5013 | 725 | 8.158 | 8.1708 | 8.1517 | 0.098634 | -0.2333 | 0.50048 | 725 | 8.1579 | 8.1734 | 8.1635 | 0.061131 | -0.12193 | 0.49937
| 740 | 8.1666 | 8.1855 | 8.2001 | 0.075242 | 0.17851 | 0.50114 | 740 | 8.1578 | 8.1708 | 8.1619 | 0.097389 | -0.10847 | 0.50103 | 740 | 8.1566 | 8.1734 | 8.1826 | 0.059955 | 0.11238 | 0.50043
| 755 | 8.1653 | 8.1855 | 8.1947 | 0.083691 | 0.1132 | 0.50108 | 755 | 8.1547 | 8.1708 | 8.1814 | 0.097729 | 0.13043 | 0.50205 | 755 | 8.1556 | 8.1734 | 8.1742 | 0.059888 | 0.0091406 | 0.49984
| 770 | 8.1625 | 8.1855 | 8.1842 | 0.074272 | -0.015795 | 0.49989 | 770 | 8.1541 | 8.1708 | 8.1644 | 0.087379 | -0.078048 | 0.50049 | 770 | 8.1541 | 8.1734 | 8.1642 | 0.057366 | -0.11282 | 0.49908
| 785 | 8.1613 | 8.1855 | 8.1898 | 0.089757 | 0.052894 | 0.50079 | 785 | 8.1511 | 8.1708 | 8.1804 | 0.087824 | 0.11825 | 0.50131 | 785 | 8.152 | 8.1734 | 8.1733 | 0.065292 | -0.0014165 | 0.49975
| 800 | 8.1617 | 8.1855 | 8.1934 | 0.087814 | 0.096344 | 0.50095 | 800 | 8.1519 | 8.1708 | 8.1586 | 0.085811 | -0.14893 | 0.49994 | 800 | 8.1517 | 8.1734 | 8.1726 | 0.068322 | -0.010247 | 0.49979
| 815 | 8.1581 | 8.1855 | 8.1906 | 0.074366 | 0.06319 | 0.50002 | 815 | 8.15 | 8.1708 | 8.1728 | 0.088466 | 0.024429 | 0.50079 | 815 | 8.1478 | 8.1734 | 8.1897 | 0.068092 | 0.19894 | 0.50058
| 830 | 8.1599 | 8.1855 | 8.1877 | 0.079882 | 0.026747 | 0.50017 | 830 | 8.1495 | 8.1708 | 8.1732 | 0.090026 | 0.029078 | 0.50085 | 830 | 8.1488 | 8.1734 | 8.1905 | 0.06326 | 0.20914 | 0.50057
| 845 | 8.1586 | 8.1855 | 8.1983 | 0.084543 | 0.15694 | 0.50092 | 845 | 8.1467 | 8.1708 | 8.1553 | 0.095251 | -0.1892 | 0.49984 | 845 | 8.1458 | 8.1734 | 8.1885 | 0.068155 | 0.18426 | 0.50039
DWG open | 260 | 8.1768 | 8.1597 | 7.4348 | 0.083362 | -8.8829 | 0.47666 | 260 | 8.171 | 8.1513 | 7.926 | 0.081506 | -2.7639 | 0.49479 | 260 | 8.1784 | 8.1615 | 8.0582 | 0.064284 | -1.2652 | 0.49875
| 275 | 8.1752 | 8.1597 | 7.4372 | 0.078632 | -8.8542 | 0.47652 | 275 | 8.1683 | 8.1513 | 7.9157 | 0.09414 | -2.8903 | 0.49456 | 275 | 8.1771 | 8.1615 | 8.0794 | 0.059701 | -1.0054 | 0.49983
| 290 | 8.1727 | 8.1597 | 7.4575 | 0.07818 | -8.6056 | 0.47764 | 290 | 8.1657 | 8.1513 | 7.927 | 0.070367 | -2.7524 | 0.49411 | 290 | 8.1776 | 8.1615 | 8.0742 | 0.068345 | -1.0698 | 0.49981
| 305 | 8.1726 | 8.1597 | 7.4751 | 0.075553 | -8.3889 | 0.47866 | 305 | 8.1647 | 8.1513 | 7.9626 | 0.088294 | -2.3155 | 0.49697 | 305 | 8.1769 | 8.1615 | 8.0599 | 0.069173 | -1.2439 | 0.49892
| 320 | 8.1704 | 8.1597 | 7.4978 | 0.086207 | -8.1116 | 0.48042 | 320 | 8.1653 | 8.1513 | 7.956 | 0.088345 | -2.396 | 0.49661 | 320 | 8.1773 | 8.1615 | 8.106 | 0.069436 | -0.6797 | 0.5018
| 335 | 8.172 | 8.1597 | 7.5246 | 0.074987 | -7.7833 | 0.48174 | 335 | 8.162 | 8.1513 | 7.956 | 0.077698 | -2.396 | 0.49595 | 335 | 8.1774 | 8.1615 | 8.0854 | 0.075084 | -0.93145 | 0.50074
| 350 | 8.1717 | 8.1597 | 7.5698 | 0.074019 | -7.2287 | 0.48457 | 350 | 8.163 | 8.1513 | 7.9705 | 0.085352 | -2.2177 | 0.49725 | 350 | 8.1751 | 8.1615 | 8.1009 | 0.079906 | -0.74264 | 0.50173
| 365 | 8.1712 | 8.1597 | 7.6289 | 0.08903 | -6.5052 | 0.48893 | 365 | 8.1618 | 8.1513 | 7.9888 | 0.082193 | -1.9936 | 0.49818 | 365 | 8.1738 | 8.1615 | 8.1002 | 0.056491 | -0.75069 | 0.50082
| 380 | 8.1689 | 8.1597 | 7.6684 | 0.075114 | -6.0207 | 0.49072 | 380 | 8.1624 | 8.1513 | 8.0113 | 0.080397 | -1.7176 | 0.49953 | 380 | 8.1716 | 8.1615 | 8.0969 | 0.066567 | -0.79167 | 0.50079
| 395 | 8.1696 | 8.1597 | 7.7136 | 0.085643 | -5.4661 | 0.49406 | 395 | 8.159 | 8.1513 | 8.0137 | 0.088123 | -1.6889 | 0.49978 | 395 | 8.173 | 8.1615 | 8.1014 | 0.060173 | -0.73605 | 0.50095
| 410 | 8.1692 | 8.1597 | 7.7378 | 0.075976 | -5.1697 | 0.49519 | 410 | 8.16 | 8.1513 | 8.039 | 0.087355 | -1.3785 | 0.50137 | 410 | 8.1711 | 8.1615 | 8.1182 | 0.063771 | -0.53041 | 0.50198
| 425 | 8.1661 | 8.1597 | 7.7969 | 0.079681 | -4.4461 | 0.4989 | 425 | 8.1577 | 8.1513 | 8.0422 | 0.083584 | -1.3382 | 0.50127 | 425 | 8.1678 | 8.1615 | 8.1016 | 0.070119 | -0.73313 | 0.50097
| 440 | 8.1681 | 8.1597 | 7.8244 | 0.078696 | -4.1083 | 0.50073 | 440 | 8.157 | 8.1513 | 8.0493 | 0.078821 | -1.252 | 0.50149 | 440 | 8.1688 | 8.1615 | 8.1176 | 0.067104 | -0.53699 | 0.50191
| 455 | 8.1651 | 8.1597 | 7.8325 | 0.08453 | -4.0095 | 0.5013 | 455 | 8.1573 | 8.1513 | 8.0502 | 0.07637 | -1.2405 | 0.50147 | 455 | 8.1653 | 8.1615 | 8.1112 | 0.070728 | -0.6153 | 0.50143
| 470 | 8.1609 | 8.1597 | 7.8429 | 0.073057 | -3.882 | 0.50125 | 470 | 8.1542 | 8.1513 | 8.068 | 0.091298 | -1.0221 | 0.50301 | 470 | 8.1651 | 8.1615 | 8.1136 | 0.068165 | -0.58676 | 0.50147
| 485 | 8.1634 | 8.1597 | 7.8372 | 0.08601 | -3.9521 | 0.50155 | 485 | 8.1533 | 8.1513 | 8.0671 | 0.08806 | -1.0336 | 0.50273 | 485 | 8.1631 | 8.1615 | 8.1181 | 0.072376 | -0.53114 | 0.50178
| 500 | 8.1609 | 8.1597 | 7.8517 | 0.084159 | -3.7736 | 0.50228 | 500 | 8.1533 | 8.1513 | 8.0615 | 0.09759 | -1.1025 | 0.50284 | 500 | 8.161 | 8.1615 | 8.116 | 0.054551 | -0.55749 | 0.50095
| 515 | 8.1607 | 8.1597 | 7.8572 | 0.083151 | -3.7067 | 0.50254 | 515 | 8.1503 | 8.1513 | 8.075 | 0.07897 | -0.93583 | 0.50266 | 515 | 8.1621 | 8.1615 | 8.123 | 0.061004 | -0.47113 | 0.50164
| 530 | 8.1612 | 8.1597 | 7.885 | 0.085142 | -3.3656 | 0.50443 | 530 | 8.1524 | 8.1513 | 8.0605 | 0.077672 | -1.114 | 0.50185 | 530 | 8.1628 | 8.1615 | 8.1156 | 0.06477 | -0.56261 | 0.50134
| 545 | 8.1575 | 8.1597 | 7.8949 | 0.085977 | -3.2445 | 0.50486 | 545 | 8.1482 | 8.1513 | 8.0774 | 0.099049 | -0.90709 | 0.50359 | 545 | 8.1596 | 8.1615 | 8.1132 | 0.057725 | -0.59115 | 0.50079
| 560 | 8.1518 | 8.1597 | 7.9313 | 0.085005 | -2.7983 | 0.50679 | 560 | 8.1485 | 8.1513 | 8.0769 | 0.091902 | -0.91283 | 0.50324 | 560 | 8.1594 | 8.1615 | 8.1215 | 0.06903 | -0.49016 | 0.50165
| 575 | 8.1533 | 8.1597 | 7.9659 | 0.072471 | -2.3743 | 0.50859 | 575 | 8.1461 | 8.1513 | 8.0694 | 0.078395 | -1.0048 | 0.50205 | 575 | 8.1561 | 8.1615 | 8.1238 | 0.062416 | -0.46162 | 0.50136
| 590 | 8.1522 | 8.1597 | 7.9896 | 0.080926 | -2.0843 | 0.51034 | 590 | 8.1459 | 8.1513 | 8.09 | 0.078287 | -0.75188 | 0.5033 | 590 | 8.1561 | 8.1615 | 8.1081 | 0.067353 | -0.65336 | 0.50056
| 605 | 8.1522 | 8.1597 | 8.0221 | 0.083391 | -1.6859 | 0.5125 | 605 | 8.1479 | 8.1513 | 8.0844 | 0.096123 | -0.82086 | 0.50386 | 605 | 8.1555 | 8.1615 | 8.1265 | 0.089714 | -0.42868 | 0.50253
| 620 | 8.1506 | 8.1597 | 8.04 | 0.071811 | -1.4659 | 0.51311 | 620 | 8.1441 | 8.1513 | 8.1064 | 0.093267 | -0.55068 | 0.50486 | 620 | 8.1548 | 8.1615 | 8.1331 | 0.059022 | -0.34745 | 0.50177
| 635 | 8.1498 | 8.1597 | 8.0481 | 0.072941 | -1.3671 | 0.51361 | 635 | 8.1436 | 8.1513 | 8.0999 | 0.10391 | -0.63116 | 0.50493 | 635 | 8.1566 | 8.1615 | 8.1226 | 0.067431 | -0.47625 | 0.50148
| 650 | 8.1488 | 8.1597 | 8.0491 | 0.068628 | -1.3544 | 0.51346 | 650 | 8.1429 | 8.1513 | 8.1139 | 0.089576 | -0.4587 | 0.50507 | 650 | 8.153 | 8.1615 | 8.1231 | 0.067774 | -0.4704 | 0.5013
| 665 | 8.1481 | 8.1597 | 8.0832 | 0.075677 | -0.93684 | 0.51583 | 665 | 8.1414 | 8.1513 | 8.0919 | 0.093706 | -0.72888 | 0.50382 | 665 | 8.1532 | 8.1615 | 8.1204 | 0.061715 | -0.5026 | 0.50096
| 680 | 8.1455 | 8.1597 | 8.0796 | 0.080991 | -0.98147 | 0.51568 | 680 | 8.1412 | 8.1513 | 8.0877 | 0.094976 | -0.78062 | 0.5036 | 680 | 8.1505 | 8.1615 | 8.1175 | 0.067887 | -0.53919 | 0.50081
| 695 | 8.1453 | 8.1597 | 8.0811 | 0.084571 | -0.96234 | 0.51586 | 695 | 8.1379 | 8.1513 | 8.1013 | 0.088975 | -0.61391 | 0.50395 | 695 | 8.1504 | 8.1615 | 8.126 | 0.072732 | -0.43454 | 0.5015
| 710 | 8.1437 | 8.1597 | 8.0897 | 0.081779 | -0.85716 | 0.5162 | 710 | 8.1388 | 8.1513 | 8.1116 | 0.080883 | -0.48744 | 0.50429 | 710 | 8.1493 | 8.1615 | 8.1313 | 0.069396 | -0.3694 | 0.50164
| 725 | 8.1422 | 8.1597 | 8.0926 | 0.076093 | -0.8221 | 0.51607 | 725 | 8.1376 | 8.1513 | 8.1027 | 0.093519 | -0.59667 | 0.50422 | 725 | 8.1481 | 8.1615 | 8.1235 | 0.062418 | -0.46527 | 0.50085
| 740 | 8.1412 | 8.1597 | 8.0882 | 0.079617 | -0.87628 | 0.51585 | 740 | 8.1347 | 8.1513 | 8.1167 | 0.098368 | -0.42421 | 0.50513 | 740 | 8.1489 | 8.1615 | 8.1187 | 0.067206 | -0.52382 | 0.50076
| 755 | 8.1415 | 8.1597 | 8.1022 | 0.075906 | -0.70417 | 0.51662 | 755 | 8.1344 | 8.1513 | 8.1027 | 0.084288 | -0.59667 | 0.50362 | 755 | 8.1463 | 8.1615 | 8.1351 | 0.06601 | -0.3233 | 0.50161
| 770 | 8.1374 | 8.1597 | 8.0913 | 0.084576 | -0.83803 | 0.51601 | 770 | 8.1343 | 8.1513 | 8.0943 | 0.091524 | -0.70014 | 0.50341 | 770 | 8.1439 | 8.1615 | 8.1347 | 0.069296 | -0.32769 | 0.50151
| 785 | 8.1371 | 8.1597 | 8.1118 | 0.072966 | -0.58623 | 0.51685 | 785 | 8.1346 | 8.1513 | 8.1167 | 0.095015 | -0.42421 | 0.50497 | 785 | 8.1426 | 8.1615 | 8.1295 | 0.063388 | -0.39136 | 0.50092
| 800 | 8.1347 | 8.1597 | 8.1188 | 0.078745 | -0.50017 | 0.51734 | 800 | 8.1359 | 8.1513 | 8.1303 | 0.09575 | -0.2575 | 0.50594 | 800 | 8.1421 | 8.1615 | 8.1491 | 0.06289 | -0.15205 | 0.50207
| 815 | 8.1355 | 8.1597 | 8.098 | 0.08268 | -0.75516 | 0.51624 | 815 | 8.1308 | 8.1513 | 8.106 | 0.086604 | -0.55643 | 0.5037 | 815 | 8.1417 | 8.1615 | 8.1292 | 0.061659 | -0.39575 | 0.50079
| 830 | 8.1343 | 8.1597 | 8.1155 | 0.080255 | -0.54161 | 0.51717 | 830 | 8.1306 | 8.1513 | 8.1266 | 0.09021 | -0.30349 | 0.50512 | 830 | 8.1405 | 8.1615 | 8.1483 | 0.063638 | -0.16156 | 0.50195
| 845 | 8.1315 | 8.1597 | 8.0993 | 0.086511 | -0.73923 | 0.51623 | 845 | 8.1275 | 8.1513 | 8.1041 | 0.10055 | -0.57942 | 0.50416 | 845 | 8.1376 | 8.1615 | 8.1525 | 0.061205 | -0.11033 | 0.50195

## DWG - Thies Classic
| Thies Classic with Rotor #1; 8 m/s | [6 empty] | Thies Classic with Rotor #4; 8 m/s | [6 empty] | Thies Classic with Rotor #3; 8 m/s
| d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d)) | d | vref_ac (d) | v'ref (d) | vMUT (d) | sMUT (d) | E(d) | U(E(d))
| (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%) | (mm) | (m/s) | (m/s) | (m/s) | (m/s) | (%) | (%)
DWG wall | 200 | 8.1808 | 8.1531 | 7.575 | 0.083446 | -7.0896 | 0.468 | 200 | 8.2178 | 8.1764 | 7.7798 | 0.08245 | -4.8509 | 0.47992 | 200 | 8.1359 | 8.1032 | 7.8989 | 0.05967 | -2.521 | 0.49028
| 215 | 8.1788 | 8.1531 | 7.7111 | 0.076498 | -5.4207 | 0.47594 | 215 | 8.2179 | 8.1764 | 7.8796 | 0.088438 | -3.6301 | 0.4863 | 215 | 8.1347 | 8.1032 | 7.9508 | 0.052429 | -1.8806 | 0.49321
| 230 | 8.1794 | 8.1531 | 7.822 | 0.088581 | -4.0602 | 0.48329 | 230 | 8.2175 | 8.1764 | 7.9239 | 0.08608 | -3.0884 | 0.48889 | 230 | 8.1327 | 8.1032 | 7.9496 | 0.058073 | -1.8954 | 0.49318
| 245 | 8.1777 | 8.1531 | 7.9079 | 0.093551 | -3.0069 | 0.48868 | 245 | 8.2164 | 8.1764 | 7.9831 | 0.085146 | -2.364 | 0.4924 | 245 | 8.1334 | 8.1032 | 7.9824 | 0.061328 | -1.4904 | 0.49535
| 260 | 8.1779 | 8.1531 | 7.9818 | 0.083465 | -2.1011 | 0.49276 | 260 | 8.2128 | 8.1764 | 7.9993 | 0.091231 | -2.166 | 0.49345 | 260 | 8.1331 | 8.1032 | 7.9788 | 0.061039 | -1.5355 | 0.4951
| 275 | 8.176 | 8.1531 | 8.028 | 0.083411 | -1.5338 | 0.49548 | 275 | 8.2112 | 8.1764 | 8.0179 | 0.085082 | -1.9385 | 0.49422 | 275 | 8.132 | 8.1032 | 8.0068 | 0.057306 | -1.1892 | 0.49665
| 290 | 8.1745 | 8.1531 | 8.0974 | 0.086088 | -0.6831 | 0.49975 | 290 | 8.2104 | 8.1764 | 8.0301 | 0.092419 | -1.7899 | 0.49524 | 290 | 8.1299 | 8.1032 | 8.0156 | 0.066075 | -1.0817 | 0.49734
| 305 | 8.1746 | 8.1531 | 8.115 | 0.081201 | -0.46713 | 0.50063 | 305 | 8.2095 | 8.1764 | 8.0603 | 0.087291 | -1.42 | 0.49681 | 305 | 8.1289 | 8.1032 | 8.0204 | 0.059381 | -1.0217 | 0.49736
| 320 | 8.1738 | 8.1531 | 8.1327 | 0.078064 | -0.24995 | 0.50155 | 320 | 8.2075 | 8.1764 | 8.0698 | 0.086701 | -1.3038 | 0.49725 | 320 | 8.1269 | 8.1032 | 8.0263 | 0.061364 | -0.94914 | 0.49767
| 335 | 8.1715 | 8.1531 | 8.1614 | 0.079514 | 0.10277 | 0.50323 | 335 | 8.205 | 8.1764 | 8.1132 | 0.088321 | -0.77294 | 0.49982 | 335 | 8.1269 | 8.1032 | 8.031 | 0.064582 | -0.89052 | 0.49806
| 350 | 8.1726 | 8.1531 | 8.1889 | 0.081714 | 0.43905 | 0.50506 | 350 | 8.2024 | 8.1764 | 8.1072 | 0.084996 | -0.84672 | 0.49915 | 350 | 8.1253 | 8.1032 | 8.0552 | 0.060195 | -0.59205 | 0.49931
| 365 | 8.1697 | 8.1531 | 8.215 | 0.074155 | 0.75978 | 0.50619 | 365 | 8.202 | 8.1764 | 8.1138 | 0.083644 | -0.76559 | 0.49948 | 365 | 8.1226 | 8.1032 | 8.0539 | 0.06418 | -0.60806 | 0.4992
| 380 | 8.1695 | 8.1531 | 8.2264 | 0.075437 | 0.899 | 0.50692 | 380 | 8.1966 | 8.1764 | 8.1306 | 0.084961 | -0.56091 | 0.50022 | 380 | 8.123 | 8.1032 | 8.0548 | 0.064836 | -0.59781 | 0.4993
| 395 | 8.1679 | 8.1531 | 8.2517 | 0.073929 | 1.2094 | 0.50832 | 395 | 8.1982 | 8.1764 | 8.1391 | 0.083166 | -0.45638 | 0.50077 | 395 | 8.1187 | 8.1032 | 8.0731 | 0.061472 | -0.37132 | 0.50005
| 410 | 8.1632 | 8.1531 | 8.2531 | 0.076594 | 1.2273 | 0.50822 | 410 | 8.1934 | 8.1764 | 8.1421 | 0.095493 | -0.41985 | 0.50121 | 410 | 8.1177 | 8.1032 | 8.0546 | 0.068895 | -0.59985 | 0.4991
| 425 | 8.1634 | 8.1531 | 8.2613 | 0.076783 | 1.3279 | 0.50874 | 425 | 8.1916 | 8.1764 | 8.1575 | 0.091242 | -0.23109 | 0.50185 | 425 | 8.117 | 8.1032 | 8.0902 | 0.066889 | -0.16088 | 0.50118
| 440 | 8.1614 | 8.1531 | 8.2686 | 0.083959 | 1.417 | 0.50934 | 440 | 8.1876 | 8.1764 | 8.1661 | 0.081494 | -0.12634 | 0.5017 | 440 | 8.1155 | 8.1032 | 8.078 | 0.072133 | -0.31087 | 0.50053
| 455 | 8.1587 | 8.1531 | 8.2264 | 0.08214 | 0.89921 | 0.50652 | 455 | 8.1874 | 8.1764 | 8.1735 | 0.079502 | -0.036205 | 0.50206 | 455 | 8.1119 | 8.1032 | 8.1041 | 0.067224 | 0.010392 | 0.50173
| 470 | 8.1578 | 8.1531 | 8.2342 | 0.083137 | 0.99484 | 0.50698 | 470 | 8.1847 | 8.1764 | 8.1856 | 0.072944 | 0.11276 | 0.50239 | 470 | 8.1097 | 8.1032 | 8.103 | 0.061932 | -0.0027998 | 0.50136
| 485 | 8.1589 | 8.1531 | 8.1983 | 0.082678 | 0.55508 | 0.50483 | 485 | 8.1807 | 8.1764 | 8.176 | 0.076195 | -0.0055516 | 0.50168 | 485 | 8.1084 | 8.1032 | 8.1151 | 0.070162 | 0.14684 | 0.5023
| 500 | 8.1561 | 8.1531 | 8.1562 | 0.08727 | 0.038746 | 0.50229 | 500 | 8.1772 | 8.1764 | 8.1724 | 0.081062 | -0.04976 | 0.50143 | 500 | 8.107 | 8.1032 | 8.1105 | 0.07203 | 0.090103 | 0.502
| 515 | 8.1545 | 8.1531 | 8.1158 | 0.096492 | -0.45752 | 0.50014 | 515 | 8.1755 | 8.1764 | 8.1479 | 0.080745 | -0.34903 | 0.49983 | 515 | 8.1064 | 8.1032 | 8.094 | 0.07685 | -0.11311 | 0.50113
| 530 | 8.1508 | 8.1531 | 8.1052 | 0.096183 | -0.58692 | 0.49926 | 530 | 8.1761 | 8.1764 | 8.1508 | 0.086294 | -0.31396 | 0.50027 | 530 | 8.1049 | 8.1032 | 8.084 | 0.076191 | -0.23729 | 0.50039
| 545 | 8.1488 | 8.1531 | 8.0971 | 0.096836 | -0.68641 | 0.49867 | 545 | 8.1703 | 8.1764 | 8.1367 | 0.094848 | -0.48533 | 0.49945 | 545 | 8.0997 | 8.1032 | 8.071 | 0.072238 | -0.39789 | 0.49912
| 560 | 8.1449 | 8.1531 | 8.0649 | 0.094909 | -1.0818 | 0.49637 | 560 | 8.1687 | 8.1764 | 8.1174 | 0.0891 | -0.7225 | 0.49791 | 560 | 8.1002 | 8.1032 | 8.0563 | 0.063498 | -0.57881 | 0.49795
| 575 | 8.1437 | 8.1531 | 8.0321 | 0.098902 | -1.4836 | 0.4945 | 575 | 8.1659 | 8.1764 | 8.1006 | 0.085683 | -0.92739 | 0.49657 | 575 | 8.0962 | 8.1032 | 8.0382 | 0.066523 | -0.80274 | 0.49669
| 590 | 8.1417 | 8.1531 | 8.0202 | 0.0969 | -1.6295 | 0.49356 | 590 | 8.1634 | 8.1764 | 8.1133 | 0.088871 | -0.77225 | 0.49733 | 590 | 8.0967 | 8.1032 | 8.0296 | 0.061724 | -0.90843 | 0.49604
| 605 | 8.1398 | 8.1531 | 8.008 | 0.095047 | -1.7789 | 0.49261 | 605 | 8.162 | 8.1764 | 8.0943 | 0.092433 | -1.0047 | 0.49626 | 605 | 8.0926 | 8.1032 | 8.0151 | 0.071477 | -1.0877 | 0.49523
| 620 | 8.137 | 8.1531 | 7.9761 | 0.089318 | -2.1704 | 0.49022 | 620 | 8.1578 | 8.1764 | 8.0888 | 0.088735 | -1.072 | 0.4955 | 620 | 8.0923 | 8.1032 | 8.0124 | 0.057981 | -1.1209 | 0.4946
| 635 | 8.1366 | 8.1531 | 7.9693 | 0.10084 | -2.2539 | 0.49042 | 635 | 8.1571 | 8.1764 | 8.076 | 0.083595 | -1.2284 | 0.49446 | 635 | 8.0908 | 8.1032 | 8.0086 | 0.066625 | -1.168 | 0.49454
| 650 | 8.1347 | 8.1531 | 7.951 | 0.094845 | -2.478 | 0.48882 | 650 | 8.156 | 8.1764 | 8.0969 | 0.083326 | -0.97303 | 0.49564 | 650 | 8.0894 | 8.1032 | 8.0087 | 0.064682 | -1.1661 | 0.4944
| 665 | 8.1317 | 8.1531 | 7.9399 | 0.096463 | -2.6145 | 0.48805 | 665 | 8.1545 | 8.1764 | 8.0723 | 0.090629 | -1.274 | 0.49438 | 665 | 8.0866 | 8.1032 | 8.0021 | 0.066608 | -1.2482 | 0.49389
| 680 | 8.1301 | 8.1531 | 7.9527 | 0.099847 | -2.457 | 0.4889 | 680 | 8.1553 | 8.1764 | 8.0699 | 0.08348 | -1.3027 | 0.49397 | 680 | 8.0871 | 8.1032 | 7.9988 | 0.064217 | -1.2887 | 0.49364
| 695 | 8.1267 | 8.1531 | 7.9384 | 0.098439 | -2.6323 | 0.48776 | 695 | 8.1515 | 8.1764 | 8.0808 | 0.091142 | -1.1699 | 0.49474 | 695 | 8.0827 | 8.1032 | 8.0153 | 0.06406 | -1.0846 | 0.49438
| 710 | 8.1262 | 8.1531 | 7.9517 | 0.09982 | -2.4703 | 0.4886 | 710 | 8.1517 | 8.1764 | 8.0733 | 0.088253 | -1.2613 | 0.49426 | 710 | 8.079 | 8.1032 | 7.994 | 0.059825 | -1.3477 | 0.49271
| 725 | 8.1246 | 8.1531 | 7.9176 | 0.084146 | -2.8877 | 0.48571 | 725 | 8.15 | 8.1764 | 8.0808 | 0.078113 | -1.1697 | 0.4941 | 725 | 8.0794 | 8.1032 | 7.9912 | 0.071224 | -1.3826 | 0.49294
| 740 | 8.1243 | 8.1531 | 7.935 | 0.08867 | -2.6744 | 0.48694 | 740 | 8.1469 | 8.1764 | 8.0893 | 0.089012 | -1.0653 | 0.49489 | 740 | 8.0813 | 8.1032 | 7.9878 | 0.071824 | -1.424 | 0.49287
| 755 | 8.1198 | 8.1531 | 7.935 | 0.10496 | -2.6746 | 0.48749 | 755 | 8.1461 | 8.1764 | 8.109 | 0.089092 | -0.82484 | 0.49604 | 755 | 8.078 | 8.1032 | 8.0137 | 0.058307 | -1.1047 | 0.49381
| 770 | 8.117 | 8.1531 | 7.9155 | 0.1017 | -2.9131 | 0.48596 | 770 | 8.1432 | 8.1764 | 8.1093 | 0.080559 | -0.82089 | 0.49551 | 770 | 8.0772 | 8.1032 | 8.0041 | 0.057264 | -1.2229 | 0.49315
| 785 | 8.1191 | 8.1531 | 7.9505 | 0.1016 | -2.4849 | 0.4882 | 785 | 8.1425 | 8.1764 | 8.1074 | 0.088755 | -0.84476 | 0.49571 | 785 | 8.0769 | 8.1032 | 8.0108 | 0.063928 | -1.1405 | 0.49374
| 800 | 8.117 | 8.1531 | 7.9406 | 0.098493 | -2.6054 | 0.48732 | 800 | 8.1413 | 8.1764 | 8.1173 | 0.087732 | -0.72345 | 0.49619 | 800 | 8.0753 | 8.1032 | 8.0055 | 0.067699 | -1.2055 | 0.49345
| 815 | 8.1168 | 8.1531 | 7.9312 | 0.085809 | -2.7214 | 0.48613 | 815 | 8.1362 | 8.1764 | 8.1178 | 0.083858 | -0.71729 | 0.49574 | 815 | 8.0704 | 8.1032 | 8.0196 | 0.061702 | -1.0314 | 0.49382
| 830 | 8.1132 | 8.1531 | 7.9385 | 0.094449 | -2.6321 | 0.48676 | 830 | 8.1358 | 8.1764 | 8.1367 | 0.092292 | -0.48539 | 0.49724 | 830 | 8.0709 | 8.1032 | 8.0227 | 0.078593 | -0.99349 | 0.49464