## Intro
FINAL REPORT OF THE EURAMET P1061 v2
Comparison of Air Temperature Calibrations
Annex 3: | Summary of the results reported by the participants
Results reported by the participants to the coordinator are summarised in this file.

## Air_ascending
RESULTS OF THE EURAMET P1061
Comparison of Air Temperature Calibrations
- Measurement in ascending order
|  |  |  | LOCAL REFERENCE |  | TRANSFER STANDARDS
|  |  |  | Reference temperature (°C) |  | ASL Probe A |  | ASL Probe B |  | HMT335
|  | No | Nominal measurement point (°C) | Mean value | Standard uncertainty | Mean value (ohm) | Standard uncertainty (ohm) * | Mean value (ohm) | Standard uncertainty (ohm) * | Mean value (°C) | Standard uncertainty (°C) *
Lab: | MKEH (Hungary) | 1 | -5 | -4.917 | 0.031 | 97.997 | 0.0005 | 98.045 | 0.0005 | -4.9 | 0.02
Date: | March 2009 | 2 | 10 | 10.028 | 0.024 | 103.837 | 0.0004 | 103.999 | 0.0005 | 9.9 | 0.02
|  | 3 | 20 | 20.031 | 0.018 | 107.737 | 0.0003 | 107.967 | 0.0003 | 19.9 | 0.01
|  | 4 | 30 | 30.19 | 0.018 | 111.677 | 0.0002 | 111.988 | 0.0003 | 30.1 | 0.01
|  | 5 | 40 | 40.033 | 0.023 | 115.487 | 0.0002 | 115.87 | 0.0005 | 39.9 | 0.01
Lab: | Centro Español de Metrología | 1 | 10 | 10.111 | 0.12 | 103.888 | 0.002 | 104.044 | 0.002
Date: | June 2009 | 2 | 20 | 20.024 | 0.12 | 107.747 | 0.002 | 107.974 | 0.002 | 20.054 | 0.006
|  | 3 | 30 | 30.049 | 0.12 | 111.633 | 0.002 | 111.937 | 0.002 | 30.017 | 0.01
|  | 4 | 40 | 40.033 | 0.12 | 115.49 | 0.002 | 115.874 | 0.002 | 39.968 | 0.01
|  | 5 | 50 | 49.987 | 0.12 | 119.311 | 0.002 | 119.784 | 0.002 | 49.855 | 0.01
Lab: | INTA | 1 | -40 | -39.8079 | 0.0244 | 84.2567 | 0.0025682 | 84.0393 | 0.0014586 | -39.9279 | 0.011179
Date: | May 2009 | 2 | 10 | 10.0146 | 0.021099 | 103.833 | 0.0028973 | 103.9929 | 0.0029063 | 9.9805 | 0.012053
|  | 3 | 50 | 49.98 | 0.016649 | 119.3195 | 0.00062034 | 119.7804 | 0.00036471 | 49.9587 | 0.0074818
|  | 4 | 90 | 89.7485 | 0.028885 | 134.5512 | 0.00058992 | 135.308 | 0.00036291 | 89.6808 | 0.0078817
|  | 5 | 150 | 149.06 | 0.047456 | 156.8998 | 0.00068259 | 158.0978 | 0.00037868 | 148.8558 | 0.0069423
Lab: | Humidity Laboratory GUM Poland | 1 | -40 | -39.7902 | 0.057 | 84.2778 | 0.0002452 | 84.0658 | 0.00021643 | -39.8799 | 0.0010403
Date: | July 2009 | 2 | -5 | -5.4425 | 0.025 | 97.8122 | 0.00044779 | 97.8601 | 0.00042463 | -5.4655 | 0.0014123
|  | 3 | 20 | 20.0328 | 0.031 | 107.7516 | 0.0023423 | 107.9903 | 0.0022353 | 20.0451 | 0.0038218
|  | 4 | 50 | 50.3733 | 0.021 | 119.4788 | 0.00078683 | 119.9455 | 0.00078852 | 50.3521 | 0.0026339
|  | 5 | 90 | 90.152 | 0.035 | 134.6979 | 0.00051041 | 135.4644 | 0.00050645 | 90.0607 | 0.0014389
Lab: | CMI OI Prague, laboratory No. 1033 | 1 | -30 | -29.1171 | 0.073726 | 88.5142 | 0.0011 | 88.3443 | 0.0008 | -29.051 | 0.003
Date: | January 2010 | 2 | 10 | 10.4105 | 0.054416 | 104.0014 | 0.0005 | 104.1579 | 0.0003 | 10.385 | 0.003
|  | 3 | 40 | 39.6055 | 0.0695 | 115.3208 | 0.0005 | 115.7111 | 0.0006 | 39.536 | 0.002
|  | 4 | 80 | 80.0972 | 0.1313 | 130.8541 | 0.0007 | 131.573 | 0.0006 | 79.926 | 0.002
|  | 5 | 150 | 150.4089 | 0.18579 | 157.3638 | 0.0011 | 158.661 | 0.0011 | 149.935 | 0.006
Lab: | CMI OI Brno, Czech Republic, 6036 | 1 | -30 | -29.911 | 0.084825 | 88.189 | 0.007 | 88.032 | 0.007 | -29.99 | 0.016
Date: | January 2010 | 2 | 10 | 10.212 | 0.057396 | 103.921 | 0.002 | 104.076 | 0.002 | 10.18 | 0.006
|  | 3 | 40 | 40.113 | 0.058432 | 115.521 | 0.004 | 115.909 | 0.004 | 40.1 | 0.009
|  | 4 | 80 | 79.999 | 0.087845 | 130.84 | 0.002 | 131.51 | 0.002 | 80 | 0.003
|  | 5 | 150 | 150.066 | 0.13946 | 157.277 | 0.013 | 158.483 | 0.014 | 150.11 | 0.032
Lab: | Istituto Nazionale di Ricerca Metrologica - INRIM | 1 | -20 | -20.5986 | 0.017 | 91.8581 | 0.0003 | 91.7817 | 0.0001 | -20.65 | 0.001
Date: | February 2010 | 2 | 10 | 9.7445 | 0.017 | 103.7451 | 0.0003 | 103.8984 | 0.0002 | 9.74 | 0.001
|  | 3 | 30 | 29.8808 | 0.017 | 111.574 | 0.0006 | 111.8774 | 0.0003 | 29.89 | 0.001
|  | 4 | 50 | 49.7787 | 0.019 | 119.2597 | 0.0008 | 119.7123 | 0.0005 | 49.781 | 0.001
|  | 5 | 80 | 79.6292 | 0.019 | 130.7036 | 0.0001 | 131.379 | 4e-05 | 79.591 | 0.001
Lab: | Testing Centre, University of Tartu (TCUT) | 1 | -10 | -9.7 | 0.047 | 96.18 | 0.0017 | 96.12 | 0.0015 | -9.69 | 0.0022
Date: | November 2010 | 2 | 10 | 10.1 | 0.047 | 103.91 | 0.0005 | 103.84 | 0.0016 | 10.11 | 0.002
|  | 3 | 20 | 20.08 | 0.047 | 107.78 | 0.0016 | 107.72 | 0.002 | 20.1 | 0.0018
|  | 4 | 50 | 50.08 | 0.052 | 119.38 | 0.0005 | 119.31 | 0.0005 | 50.12 | 0.0028
|  | 5 | 80 | 79.86 | 0.059 | 130.8 | 0.0016 | 130.71 | 0.0016 | 79.87 | 0.0029
Lab: | MALTA STANDARDS AUTHORITY | 1 | 10 | 10.285 | 0.087 | 103.965 | 0.001 | 103.911 | 0.001 | 10.4 | 0.01
Date: | March 2011 | 2 | 20 | 20.184 | 0.087 | 107.814 | 0.001 | 107.754 | 0.001 | 20.24 | 0.01
|  | 3 | 30 | 29.979 | 0.087 | 111.606 | 0.001 | 111.543 | 0.001 | 29.95 | 0.01
|  | 4 | 40 | 39.787 | 0.087 | 115.402 | 0.001 | 115.332 | 0.001 | 39.75 | 0.01
|  | 5 | 50 | 49.611 | 0.087 | 119.104 | 0.001 | 119.102 | 0.001 | 49.62 | 0.01
Lab: | Hellenic Institute of Metrology (EIM) | 1 | -40 | -39.377 | 0.13 | 84.4439 | 0.0016966 | 84.4163 | 0.0016876 | -39.4685 | 0.0031014
Date: | February 2011 | 2 | -10 | -10.0085 | 0.13 | 96.0122 | 0.0021756 | 95.968 | 0.0023255 | -10.0455 | 0.0033619
|  | 3 | 20 | 19.8985 | 0.1 | 107.6831 | 0.0024336 | 107.6337 | 0.0021778 | 19.873 | 0.0041739
|  | 4 | 50 | 50.0435 | 0.1 | 119.3157 | 0.0013458 | 119.2747 | 0.0014587 | 49.9525 | 0.0033931
|  | 5 | 80 | 80.233 | 0.17 | 130.8528 | 0.0013603 | 130.8328 | 0.0012949 | 80.0485 | 0.0023255
Lab: | Bulgarian Institute of Metrology | 1 | -20 | -19.82 | 0.0754 | 92.198 | 0.022 | 92.149 | 0.015 | -19.86 | 0.047
Date: | February 2011 | 2 | 5 | 4.893 | 0.0635 | 101.811 | 0.023 | 101.813 | 0.021 | 4.88 | 0.032
|  | 3 | 20 | 19.843 | 0.0795 | 107.664 | 0.032 | 107.608 | 0.059 | 19.823 | 0.06
|  | 4 | 40 | 39.769 | 0.0595 | 115.389 | 0.001 | 115.301 | 0.002 | 39.76 | 0.021
|  | 5 | 60 | 60.221 | 0.06 | 123.261 | 0.006 | 123.15 | 0.002 | 60.25 | 0.024
Lab: | BRML - INM, Bucharest, Romania | 1 | -20 | -20.035 | 0.035 | 92.069 | 0.002 | 92.017 | 0.002 | -20.044 | 0.005
Date: | April 2011 | 2 | 10 | 9.918 | 0.03 | 103.819 | 0.002 | 103.768 | 0.002 | 9.965 | 0.005
|  | 3 | 30 | 29.997 | 0.015 | 111.61 | 0.002 | 111.544 | 0.002 | 30.02 | 0.005
|  | 4 | 50 | 49.995 | 0.02 | 119.328 | 0.002 | 119.26 | 0.002 | 49.975 | 0.005
|  | 5 | 100 | 99.963 | 0.08 | 138.393 | 0.002 | 138.324 | 0.002 | 99.841 | 0.005
Lab: | National Metrology Laboratory NSAI Ireland | 1 | -40 | -39.972 | 0.1 | 84.211 | 0.0116 | 84.177 | 0.0117 | -40.05 | 0.018
Date: | September 2011 | 2 | -10 | -10.008 | 0.091 | 96.02 | 0.0102 | 95.974 | 0.0106 | -10.04 | 0.016
|  | 3 | 30 | 30.057 | 0.093 | 111.646 | 0.0144 | 111.564 | 0.0147 | 30.08 | 0.023
|  | 4 | 60 | 60.087 | 0.08 | 123.216 | 0.0028 | 123.121 | 0.003 | 60.05 | 0.008
|  | 5 | 100 | 99.987 | 0.18 | 138.395 | 0.0028 | 138.342 | 0.0031 | 99.8 | 0.009
Lab: | NPL | 1 | -40 | -40.019 | 0.04 | 84.189 | 0.0003 | 84.1818 | 0.0005 | -40.0163 | 0.007
Date: | May 2011 | 2 | 10 | 10.016 | 0.04 | 103.849 | 0.0008 | 103.7994 | 0.001 | 10.004 | 0.007
|  | 3 | 30 | 30.026 | 0.04 | 111.628 | 0.0018 | 111.565 | 0.001 | 30.016 | 0.007
|  | 4 | 80 | 80.062 | 0.04 | 130.873 | 0 | 130.7708 | 0.0007 | 80.018 | 0.007
|  | 5 | 100 | 100.065 | 0.04 | 138.482 | 0.0006 | 138.3616 | 0.0005 | 99.999 | 0.011
Lab: | Danish Technological Institute | 1 | -10 | -8.676 | 0.056 | 96.5305 | 0.00067691 | 96.508 | 0.00056401 | -8.5 | 0.0028868
Date: | June 2011 | 2 | 10 | 10.5 | 0.019391 | 104.0313 | 0.00028868 | 103.9901 | 0.00028868 | 10.553 | 0.0028868
|  | 3 | 30 | 29.736 | 0.021973 | 111.5202 | 0.00028868 | 111.4424 | 0.00033451 | 29.687 | 0.0028868
|  | 4 | 50 | 49.308 | 0.037376 | 119.0868 | 0.00028868 | 118.9827 | 0.00066502 | 49.147 | 0.0028868
|  | 5 | 70 | 68.473 | 0.064519 | 126.4495 | 0.00028868 | 126.3114 | 0.00028868 | 68.21 | 0
Lab: | MIRS/UL-FE/LMK | 1 | -40 | -37.182 | 0.05 | 85.3423 | 0.00089667 | 85.3487 | 0.00109 | -37.23 | 0.01
Date: | July 2011 | 2 | -20 | -19.1733 | 0.05 | 92.4325 | 0.00073667 | 92.4253 | 0.00043333 | -19.2 | 0.01
|  | 3 | 40 | 40.4037 | 0.05 | 115.6389 | 0.00061333 | 115.5815 | 0.00047 | 40.4 | 0.01
|  | 4 | 70 | 69.6529 | 0.065 | 126.8918 | 0.00092333 | 126.8068 | 0.00056 | 69.62 | 0.01
|  | 5 | 100 | 96.1545 | 0.08 | 137.0079 | 0.00086333 | 136.8962 | 0.00087667 | 96.11 | 0.01
Lab: | MIRS/UL-FE/LMK | 1 | -7 | -6.9214 | 0.025 | 97.2044 | 0.00017433 | 97.1726 | 0.000236 | -6.98 | 0.01
Date: | July 2011 | 2 | 20 | 20.0156 | 0.03 | 107.7375 | 0.00010567 | 107.6824 | 9.4667e-05 | 20.01 | 0.01
|  | 3 | 40 | 39.8896 | 0.04 | 115.4556 | 0.00012167 | 115.3842 | 0.00016733 | 39.91 | 0.01
|  | 4 | 60 | 59.7711 | 0.05 | 123.1315 | 6.2333e-05 | 123.0445 | 0.00032533 | 59.83 | 0.01
|  | 5 | 70 | 69.7344 | 0.06 | 126.9612 | 0.00015133 | 126.8666 | 0.00015867 | 69.83 | 0.01
Lab: | Justervesenet | 1 | -40 | -39.517 | 0.068232 | 84.3925 | 0.024 | 84.373 | 0.026 | -39.606 | 0.037
Date: | September 2011 | 2 | 0 | 0.004 | 0.065511 | 99.9419 | 0.011 | 99.894 | 0.011 | -0.035 | 0.008
|  | 3 | 20 | 19.932 | 0.065511 | 107.7046 | 0.009 | 107.642 | 0.009 | 19.909 | 0.013
|  | 4 | 40 | 40.02 | 0.065511 | 115.4813 | 0.001 | 115.4099 | 0.002 | 39.992 | 0.011
|  | 5 | 80 | 80.211 | 0.06627 | 130.8976 | 0.007 | 130.802 | 0.005 | 80.104 | 0.013
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | -10 | -9.1996 | 0.035512 | 96.3334 | 0.0010022 | 96.3004 | 0.002045 | -9.1609 | 0.0030712
Date: | October 2011 | 2 | 10 | 10.3039 | 0.03614 | 103.965 | 0.0004714 | 103.9025 | 0.00045338 | 10.2683 | 0.0026418
|  | 3 | 30 | 29.9478 | 0.03614 | 111.5763 | 0.00091954 | 111.5266 | 0.0015362 | 29.9158 | 0.0018306
|  | 4 | 50 | 49.605 | 0.03614 | 119.1505 | 0.00034157 | 119.1179 | 0.00037859 | 49.5607 | 0.0027489
|  | 5 | 70 | 69.2995 | 0.03614 | 126.6903 | 0.00021344 | 126.6685 | 0.00063683 | 69.1979 | 0.0033844
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | -40 | -39.654 | 0.047302 | 84.3272 | 0.0017751 | 84.3133 | 0.0016333 | -39.7698 | 0.0040656
Date: | October 2011 | 2 | -10 | -9.8764 | 0.047302 | 96.0605 | 0.002437 | 96.0365 | 0.0013271 | -9.9618 | 0.0051914
|  | 3 | 30 | 30.5196 | 0.047776 | 111.8133 | 0.0020442 | 111.748 | 0.0011155 | 30.489 | 0.0026667
|  | 4 | 70 | 70.1973 | 0.047776 | 127.0982 | 0.0014892 | 126.9962 | 0.0013149 | 70.1467 | 0.0041259
|  | 5 | 100 | 99.56 | 0.047776 | 138.276 | 0.00090676 | 138.1494 | 0.0011662 | 99.4609 | 0.0031143
Lab: | Physikalisch-Technische Bundesanstalt Braunschweig (PTB) | 1 | 10 | 10.058 | 0.01 | 103.869 | 0.008 | 103.799 | 0.008 | 10.045 | 0.02
Date: | January 2012 | 2 | 20 | 20.04 | 0.01 | 107.749 | 0.008 | 107.692 | 0.008 | 20.02 | 0.02
|  | 3 | 30 | 30.005 | 0.01 | 111.609 | 0.008 | 111.565 | 0.008 | 29.98 | 0.02
|  | 4 | 40 | 40.017 | 0.01 | 115.476 | 0.008 | 115.449 | 0.008 | 39.97 | 0.02
|  | 5 | 50 | 49.982 | 0.015 | 119.313 | 0.008 | 119.312 | 0.01 | 49.92 | 0.02
Lab: | MIKES | 1 | -40 | -40.4485 | 0.078797 | 84.0425 | 0 | 83.751 | 0 | 0 | 0
Date: | January to March 2009 | 3 | -20 | -20.2342 | 0.078797 | 92.0084 | 0 | 91.8854 | 0 | 0 | 0
|  | 6 | 10 | 10.1017 | 0.078797 | 103.8994 | 0 | 104.0286 | 0 | 0 | 0
|  | 7 | 20 | 19.7531 | 0.078797 | 107.6395 | 0 | 107.8597 | 0 | 0 | 0
|  | 8 | 30 | 30.2444 | 0.078797 | 111.6938 | 0 | 112.002 | 0 | 0 | 0
|  | 9 | 40 | 40.2901 | 0.078797 | 115.5892 | 0 | 115.9682 | 0 | 0 | 0
|  | 10 | 50 | 50.3165 | 0.078797 | 119.4806 | 0 | 119.9439 | 0 | 0 | 0
|  | 11 | 60 | 60.4195 | 0.078797 | 123.3531 | 0 | 123.888 | 0 | 0 | 0
|  | 13 | 80 | 80.5037 | 0.078797 | 131.0569 | 0 | 131.721 | 0 | 0 | 0
|  | 14 | 90 | 90.5086 | 0.078797 | 134.896 | 0 | 135.6511 | 0 | 0 | 0
Lab: | MIKES | 1 | -36 | -35.8837 | 0.02 | 85.8193 | 0.0031484 | 0 | 0 | -35.9989 | 0
Date: | February 2012 to June 2013 | 2 | -20 | -20.0474 | 0.015 | 92.0687 | 0.00015275 | 0 | 0 | -20.139 | 0
|  | 3 | 0 | 0.084247 | 0.02 | 99.968 | 0.00049441 | 0 | 0 | -0.00014643 | 0
|  | 4 | 20 | 20.0481 | 0.02 | 107.7505 | 0.00022361 | 0 | 0 | 19.9662 | 0
|  | 5 | 50 | 50.1255 | 0.02 | 119.3819 | 0.00050442 | 0 | 0 | 50.0376 | 0
|  | 6 | 80 | 80.1386 | 0.025 | 130.8876 | 0.00022111 | 0 | 0 | 80.0172 | 0
|  | 7 | 90 | 90.0966 | 0.025 | 134.6846 | 0.00056174 | 0 | 0 | 89.9632 | 0
Lab: | MIKES | 1 | -40 | -40.2536 | 0.05 | 84.049 | 0.00307 |  |  | -40.5 | 0
Date: | July to August 2009 | 3 | -20 | -20.1374 | 0.05 | 92.0154 | 0.00092 |  |  | -20.3 | 0
|  | 6 | 10 | 10.0926 | 0.04 | 103.8635 | 0.00307 |  |  | 10 | 0
|  | 7 | 20 | 19.9964 | 0.04 | 107.7287 | 0.002 |  |  | 19.9 | 0
|  | 8 | 30 | 30.2325 | 0.03 | 111.709 | 0 |  |  | 30.17 | 0.04583
|  | 9 | 40 | 40.2892 | 0.02 | 115.6003 | 0.00142 |  |  | 40.2 | 0
|  | 10 | 50 | 50.4377 | 0.02 | 119.503 | 0 |  |  | 50.3 | 0
|  | 11 | 60 | 60.5269 | 0.02 | 123.389 | 0 |  |  | 60.4 | 0
|  | 13 | 80 | 80.704 | 0.03 | 131.118 | 0 |  |  | 80.6 | 0
|  | 14 | 90 | 90.7495 | 0.04 | 134.954 | 0 |  |  | 90.6 | 0
|  | 15 | 120 | 120.7051 | 0.07 | 146.2657 | 0.00149 |  |  | 120.5 | 0
|  | 16 | 150 | 150.8327 | 0.15 | 157.5477 | 0.00155 |  |  | 150.6 | 0
|  | 17 | 180 | 180.8556 | 0.3 | 168.7149 | 0.00221 |  |  | 180.5 | 0
Results obtained for the HMT335 measured separately from the PRTs are marked with red font. (These results shown here were adjusted to correspond the reference temperature values in column G.)

## Air_descending
RESULTS OF THE EURAMET P1061
Comparison of Air Temperature Calibrations
- Measurement in descending order
|  |  |  | LOCAL REFERENCE |  | TRANSFER STANDARDS
|  |  |  | Reference temperature (°C) |  | ASL Probe A |  | ASL Probe B |  | HMT335
|  | No | Nominal measurement point (°C) | Mean value | Standard uncertainty | Mean value (ohm) | Standard uncertainty (ohm) * | Mean value (ohm) | Standard uncertainty (ohm) * | Mean value (°C) | Standard uncertainty (°C) *
Lab: | MKEH (Hungary) | 1 | 40 | 40.084 | 0.001 | 115.506 | 0.0002 | 115.889 | 0.0005 | 40 | 0.01
|  | 2 | 20 | 20.187 | 0.001 | 107.793 | 0.0003 | 108.028 | 0.0003 | 20.1 | 0.01
|  | 3 | -5 | -4.939 | 0.001 | 97.989 | 0.0005 | 98.035 | 0.0005 | -4.9 | 0.02
Lab: | Centro Español de Metrología | 1 | 50 | 49.999 | 0.12 | 119.324 | 0.002 | 119.786 | 0.002 | 49.861 | 0.01
|  | 2 | 30 | 30.055 | 0.12 | 111.632 | 0.002 | 111.939 | 0.002 | 30.055 | 0.009
|  | 3 | 10 | 10.053 | 0.12 | 103.868 | 0.002 | 104.02 | 0.002 | 10.064 | 0.006
Lab: | INTA | 1 | 150 | 149.06 | 0.047456 | 156.8998 | 0.00068259 | 158.0978 | 0.00037868 | 148.8558 | 0.0069423
|  | 2 | 50 | 50.0977 | 0.016844 | 119.3645 | 0.00045447 | 119.8248 | 0.00030516 | 50.0712 | 0.0059944
|  | 3 | -40 | -39.7317 | 0.027072 | 84.2878 | 0.0063743 | 84.0707 | 0.0045507 | -39.8534 | 0.01408
Lab: | Humidity Laboratory GUM Poland | 1 | 90 | 90.1895 | 0.042 | 134.7106 | 0.00025627 | 135.4767 | 0.00026062 | 90.0876 | 0.0012163
|  | 2 | 20 | 19.9722 | 0.031 | 107.7331 | 0.0011382 | 107.9675 | 0.0012004 | 19.9816 | 0.0016721
|  | 3 | -40 | -39.8147 | 0.059 | 84.2672 | 0.00037002 | 84.0528 | 0.0003403 | -39.9119 | 0.0011459
Lab: | CMI OI Prague, laboratory No. 1033 | 1 | 150 | 150.3977 | 0.18579 | 157.3643 | 0.0008 | 158.6609 | 0.0008 | 149.883 | 0.004
|  | 2 | 80 | 79.3572 | 0.1313 | 130.5873 | 0.0011 | 131.2832 | 0.0007 | 79.223 | 0.003
|  | 3 | -30 | -29.3453 | 0.073726 | 88.4166 | 0.001 | 88.2416 | 0.0007 | -29.229 | 0.004
Lab: | CMI OI Brno, Czech Republic, 6036 | 1 | 150 | 149.832 | 0.13946 | 157.2 | 0.006 | 158.403 | 0.006 | 149.9 | 0.012
|  | 2 | 40 | 39.735 | 0.058432 | 115.385 | 0.002 | 115.749 | 0.001 | 39.73 | 0.004
|  | 3 | -30 | -30.043 | 0.084825 | 88.144 | 0.002 | 87.973 | 0.002 | -30.11 | 0.004
Lab: | Istituto Nazionale di Ricerca Metrologica - INRIM | 1 | 80 | 79.5361 | 0.019 | 130.6692 | 5e-05 | 131.3441 | 4e-05 | 79.498 | 0.001
|  | 2 | 30 | 29.896 | 0.017 | 111.5785 | 0.0004 | 111.8807 | 0.0002 | 29.891 | 0.001
|  | 3 | -20 | -20.3436 | 0.017 | 91.9578 | 0.0004 | 91.8806 | 0.0008 | -20.409 | 0.001
Lab: | Testing Centre, University of Tartu (TCUT) | 1 | 80 | 79.84 | 0.059 | 130.79 | 0.0013 | 130.71 | 0.001 | 79.86 | 0.0027
|  | 2 | 20 | 20.08 | 0.047 | 107.8 | 0.0005 | 107.73 | 0.0005 | 20.12 | 0.0023
|  | 3 | -10 | -9.76 | 0.047 | 96.15 | 0.0013 | 96.1 | 0.0005 | -9.75 | 0.002
Lab: | MALTA STANDARDS AUTHORITY | 1 | 50 | 49.611 | 0.087 | 119.104 | 0.001 | 119.102 | 0.001 | 49.62 | 0.01
|  | 2 | 40 | 39.868 | 0.087 | 115.428 | 0.001 | 115.362 | 0.001 | 39.81 | 0.01
|  | 3 | 30 | 30.05 | 0.087 | 111.635 | 0.001 | 111.573 | 0.001 | 30.04 | 0.01
Lab: | Hellenic Institute of Metrology (EIM) | 1 | -40 | -39.3995 | 0.13 | 84.4383 | 0.0022535 | 84.3942 | 0.0024609 | -39.4355 | 0.0040701
|  | 2 | 20 | 19.8955 | 0.1 | 107.6903 | 0.0018523 | 107.6385 | 0.0017925 | 19.913 | 0.0034868
|  | 3 | 80 | 80.294 | 0.17 | 130.8748 | 0.0023406 | 130.8489 | 0.0022917 | 80.076 | 0.004724
Lab: | Bulgarian Institute of Metrology | 1 | 60 | 60.1895 | 0.0605 | 123.254 | 0.008 | 123.14 | 0.006 | 60.21 | 0.013
|  | 2 | 20 | 19.724 | 0.0755 | 107.627 | 0.032 | 107.57 | 0.045 | 19.7 | 0.06
|  | 3 | -20 | -19.751 | 0.0833 | 92.229 | 0.022 | 92.179 | 0.011 | -19.8 | 0.044
Lab: | BRML - INM, Bucharest, Romania | 1 | 10 | 9.891 | 0.03 | 103.814 | 0.002 | 103.766 | 0.002 | 9.942 | 0.005
|  | 2 | 30 | 29.929 | 0.015 | 111.589 | 0.002 | 111.508 | 0.002 | 29.953 | 0.005
|  | 3 | 50 | 49.867 | 0.02 | 119.282 | 0.002 | 119.208 | 0.002 | 49.856 | 0.005
Lab: | National Metrology Laboratory NSAI Ireland | 1 | 100 | 99.987 | 0.18 | 138.395 | 0.0028 | 138.3416 | 0.0031 | 99.8 | 0.009
|  | 2 | 30 | 30.066 | 0.093 | 111.65 | 0.0105 | 111.5601 | 0.0106 | 30.08 | 0.015
|  | 3 | -40 | -39.992 | 0.08 | 84.199 | 0.011 | 84.1655 | 0.0109 | -40.07 | 0.014
Lab: | NPL | 1 | 100 | 100.066 | 0.04 | 138.483 | 0.0004 | 138.3636 | 0.0005 | 100 | 0.01
|  | 2 | 30 | 30.015 | 0.04 | 111.624 | 0.002 | 111.5618 | 0.002 | 30.011 | 0.011
|  | 3 | -40 | -40.019 | 0.04 | 84.191 | 0.0007 | 84.1792 | 0.0009 | -40.128 | 0.007
Lab: | Danish Technological Institute | 1 | 70 | 68.4731 | 0.064519 | 126.4495 | 0.00029 | 126.3114 | 0.00029 | 68.21 | 0
|  | 2 | 30 | 29.7835 | 0.019837 | 111.5362 | 0.00044595 | 111.4639 | 0.00029 | 29.7167 | 0.0028868
|  | 3 | -10 | -8.7408 | 0.056 | 96.5078 | 0.0010154 | 96.4864 | 0.0003384 | -8.5433 | 0.0028868
Lab: | MIRS/UL-FE/LMK | 1 | 100 | 99.5138 | 0.08 | 138.269 | 0.00036 | 138.1804 | 0.00085 | 99.47 | 0.01
|  | 2 | 40 | 40.3413 | 0.05 | 115.6288 | 0.000225 | 115.558 | 0.00041 | 40.31 | 0.01
|  | 3 | -40 | -40.0855 | 0.05 | 84.2093 | 0.00041333 | 84.1925 | 0.0010233 | -40.12 | 0.01
Lab: | MIRS/UL-FE/LMK | 1 | 70 | 69.7362 | 0.06 | 126.9601 | 8.4333e-05 | 126.8655 | 0.00056333 | 69.83 | 0.01
|  | 2 | 40 | 39.9612 | 0.04 | 115.4807 | 0.00034 | 115.4099 | 0.000187 | 39.99 | 0.01
|  | 3 | -5 | -4.2501 | 0.025 | 98.2466 | 0.000268 | 98.2079 | 0.00033333 | -4.33 | 0.01
Lab: | Justervesenet | 1 | 80 | 80.173 | 0.06627 | 130.886 | 0.005 | 130.789 | 0.006 | 80.07 | 0.013
|  | 2 | 20 | 19.932 | 0.065511 | 107.704 | 0.009 | 107.644 | 0.008 | 19.91 | 0.014
|  | 3 | -40 | -38.13 | 0.068232 | 84.941 | 0.036 | 84.918 | 0.033 | -38.25 | 0.046
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | 70 | 69.3123 | 0.03614 | 126.6955 | 0.00034157 | 126.6754 | 0.0001633 | 69.2121 | 0.0029077
|  | 2 | 30 | 29.94 | 0.03614 | 111.5822 | 0.00024944 | 111.5322 | 0.0002 | 29.9166 | 0.0037689
|  | 3 | -10 | -9.332 | 0.035512 | 96.2767 | 0.00053852 | 96.2413 | 0.0016536 | -9.3609 | 0.003965
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | 100 | 99.5732 | 0.047776 | 138.2809 | 0.0016895 | 138.154 | 0.0015916 | 99.4764 | 0.0042063
|  | 2 | 30 | 30.5196 | 0.047776 | 111.8132 | 0.0019425 | 111.7496 | 0.002202 | 30.4905 | 0.0038622
|  | 3 | -40 | -39.6794 | 0.047302 | 84.3111 | 0.0016763 | 84.3068 | 0.0016786 | -39.8214 | 0.0038274
Lab: | Physikalisch-Technische Bundesanstalt Braunschweig (PTB) | 1 | 50 | 49.982 | 0.015 | 119.308 | 0.008 | 119.305 | 0.008 | 49.92 | 0.02
|  | 2 | 30 | 30.004 | 0.01 | 111.61 | 0.008 | 111.566 | 0.008 | 29.98 | 0.02
|  | 3 | 10 | 9.95 | 0.01 | 103.829 | 0.008 | 103.76 | 0.01 | 9.94 | 0.02
Results obtained for the HMT335 measured separately from the PRTs are marked with red font. (These results shown here were adjusted to correspond the reference temperature values in column G.)

## Liquid
RESULTS OF THE EURAMET P1061
Comparison of Air Temperature Calibrations
- Measurements in liquid bath
|  |  |  | LOCAL REFERENCE |  | TRANSFER STANDARDS
|  |  |  | Reference temperature (°C) |  | ASL Probe A |  | ASL Probe B
|  | No | Nominal measurement point (°C) | Mean value | Standard uncertainty | Mean value (ohm) | Standard uncertainty (ohm) * | Mean value (ohm) | Standard uncertainty (ohm) *
Lab: | MKEH (Hungary) | 1 | 10 | 10.223 | 0.001 | 103.922 | 0.0004 | 104.085 | 0.0004
|  | 2 | 80 | 80.145 | 0.002 | 130.886 | 0.0007 | 131.563 | 0.0007
|  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lab: | Centro Español de Metrología | 1 | 10 | 9.882 | 0.015 | 103.789 | 0.001 | 103.948 | 0.001
|  | 2 | 80 | 80.006 | 0.015 | 130.844 | 0.001 | 131.525 | 0.001
|  | 5 | 10 | 10.098 | 0.015 | 103.872 | 0.001 | 104.033 | 0.001
|  | 6 | 80 | 80.093 | 0.015 | 130.875 | 0.001 | 131.556 | 0.001
Lab: | INTA | 1 | 10 | 9.9489 | 0.01 | 103.819 | 0.00057735 | 103.977 | 0.00057735
|  | 2 | 80 | 80.0472 | 0.01 | 130.865 | 0.00057735 | 131.543 | 0.00057735
|  | 5 | 10 | 9.983 | 0.01 | 103.8292 | 0.0006433 | 103.9883 | 0.00030438
|  | 6 | 80 | 79.9164 | 0.01 | 130.8115 | 0.00038416 | 131.4896 | 0.00025952
Lab: | GUM | 1 | 10 | 10.0001 | 0.00555 | 103.8323 | 0.00022111 | 103.9921 | 0.00015986
|  | 2 | 80 | 79.9999 | 0.0056 | 130.8401 | 8.8192e-05 | 131.5199 | 8.8192e-05
|  | 5 | 10 | 10.0003 | 0.0056 | 103.8329 | 8.8192e-05 | 103.9922 | 0.00011547
|  | 6 | 80 | 80.0007 | 0.0056 | 130.8404 | 0.0001291 | 131.519 | 9.4281e-05
Lab: | CMI OI Prague, laboratory No. 1033 | 1 | 10 | 10.0042 | 0.0062793 | 103.8365 | 0.00027 | 103.9938 | 0.00025
|  | 2 | 80 | 80.0056 | 0.0098609 | 130.8448 | 0.0002 | 131.5218 | 0.0002
|  | 5 | 10 | 10.0015 | 0.0062793 | 103.8328 | 0.00049 | 103.9933 | 0.00015
|  | 6 | 80 | 80.013 | 0.0098609 | 130.8415 | 0.00031 | 131.5226 | 0.00016
Lab: | CMI OI Brno, Czech Republic, 6036 | 1 | 10 | 10.014 | 0.016296 | 103.84 | 0.0002 | 104 | 0.0001
|  | 2 | 80 | 80.018 | 0.016296 | 130.846 | 0.0002 | 131.524 | 0.0002
|  | 5 | 10 | 10.008 | 0.016296 | 103.839 | 0.0001 | 103.997 | 0.0001
|  | 6 | 80 | 80.006 | 0.016296 | 130.84 | 0.0001 | 131.518 | 0.0001
Lab: | Istituto Nazionale di Ricerca Metrologica - INRIM | 1 | 10 | 10.055 | 0.003 | 103.856 | 0.0003 | 104.016 | 0.0003
|  | 2 | 80 | 80.3 | 0.003 | 130.9559 | 0.0002 | 131.636 | 0.0002
|  | 5 | 10 | 9.753 | 0.003 | 103.7354 | 0.00011947 | 0 | 0
|  | 6 | 80 | 79.978 | 0.003 | 130.8329 | 5.8824e-05 | 0 | 0
|  | 7 | 10 | 9.754 | 0.003 | 0 | 0 | 103.8964 | 0.00012304
|  | 8 | 80 | 79.985 | 0.003 | 0 | 0 | 131.5128 | 0.0001875
Lab: | Metrosert Estonia | 1 | 0 | 10.0019 | 0.0051673 | 103.8349 | 0.0010169 | 103.7901 | 0.0010028
|  | 2 | 0 | 80.0032 | 0.0051673 | 130.8421 | 0.0011324 | 130.7362 | 0.0016368
Lab: | MALTA STANDARDS AUTHORITY | 1 | 10 | 10.0034 | 0.0054 | 103.835 | 0.001 | 103.795 | 0.001
|  | 2 | 80 | 80.0065 | 0.0054 | 130.843 | 0.001 | 130.743 | 0.001
Lab: | Hellenic Institute of Metrology (EIM) | 1 | 10 | 10.409 | 0.00090757 | 103.9907 | 0.00053852 | 103.9434 | 0.0006
|  | 2 | 80 | 80.121 | 0.0011764 | 130.8898 | 0.00013334 | 130.7833 | 0.00036878
|  | 5 | 10 | 10.4217 | 0.0014641 | 103.994 | 0.0005164 | 103.9511 | 9.9999e-05
|  | 6 | 80 | 80.1093 | 0.00049015 | 130.8835 | 0.00030732 | 130.7784 | 0.0003055
Lab: | Bulgarian Institute of Metrology | 1 | 10 | 10.195 | 0.0035 | 103.91 | 0.0012 | 103.865 | 0.001
|  | 2 | 80 | 80.0561 | 0.005 | 130.864 | 0.001 | 130.756 | 0.002
Lab: | BRML - INM, Bucharest, Romania | 1 | 10 | 10 | 0.02 | 103.833 | 0.001 | 103.792 | 0.001
|  | 2 | 80 | 80 | 0.02 | 130.842 | 0.001 | 130.744 | 0.001
|  | 5 | 10 | 10 | 0.02 | 103.834 | 0.001 | 103.793 | 0.001
|  | 6 | 80 | 79.999 | 0.02 | 130.842 | 0.001 | 130.744 | 0.001
Lab: | National Metrology Laboratory NSAI Ireland | 1 | 10 | 10.0013 | 0.002 | 103.831 | 0.0001 | 103.791 | 0.0016
|  | 2 | 80 | 80.3608 | 0.003 | 130.977 | 0.0001 | 130.878 | 0.0004
|  | 5 | 10 | 10.0013 | 0.002 | 103.832 | 0.0005 | 103.793 | 0.0003
|  | 6 | 80 | 80.1811 | 0.003 | 130.912 | 0.0004 | 130.809 | 0.0004
Lab: | NPL | 1 | 10 | 10.162 | 0.0025 | 103.895 | 0.0009 | 103.852 | 0.0022
|  | 2 | 80 | 80.095 | 0.0025 | 130.875 | 0.0006 | 130.772 | 0.0008
|  | 5 | 10
|  | 6 | 80
Lab: | Danish Technological Institute | 1 | 20 | 20.024 | 0.01 | 107.7327 | 0.00029 | 107.6877 | 0.00038232
|  | 2 | 80 | 79.5707 | 0.0073 | 130.6759 | 0.00029 | 130.5811 | 0.00029
|  | 5 | 10 | 10.5065 | 0.010066 | 104.031 | 0.00029 | 103.9906 | 0.00029
|  | 6 | 80 | 79.5732 | 0.0076467 | 130.676 | 0.00029 | 130.5777 | 0.00029
Lab: | MIRS/UL-FE/LMK | 1 | 10 | 9.9791 | 0.003 | 103.8278 | 0.00018767 | 103.7828 | 0.000331
|  | 2 | 80 | 80.0727 | 0.003 | 130.8709 | 0.00035333 | 130.7649 | 0.00037667
|  | 5 | 10 | 9.9833 | 0.003 | 103.8275 | 5.4667e-05 | 103.7804 | 0.0003
|  | 6 | 80 | 79.9665 | 0.003 | 130.8299 | 0.00035 | 130.7241 | 0.00056
Lab: | MIRS/UL-FE/LMK | 1 | 10 | 9.9791 | 0.003 | 103.8278 | 0.00018767 | 103.7828 | 0.000331
|  | 2 | 80 | 80.0727 | 0.003 | 130.8709 | 0.00035333 | 130.7649 | 0.00037667
|  | 5 | 10 | 9.9833 | 0.003 | 103.8275 | 5.4667e-05 | 103.7804 | 0.0003
|  | 6 | 80 | 79.9665 | 0.003 | 130.8299 | 0.00035 | 130.7241 | 0.00056
Lab: | Justervesenet | 1 | 10 | 10.0039 | 0.003239 | 103.8369 | 0.00055515 | 103.798 | 0.00018411
|  | 2 | 80 | 80.0046 | 0.0035683 | 130.8432 | 0.00064221 | 130.7439 | 0.00049403
|  | 5 | 10 | 9.9979 | 0.003239 | 103.8328 | 0.00037582 | 103.7943 | 0.00048099
|  | 6 | 80 | 80.0001 | 0.0035683 | 130.8404 | 0.00053016 | 130.7412 | 0.00053652
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | 10 | 9.9933 | 0.00065064 | 103.8279 | 0.00023333 | 103.7894 | 0.00081921
|  | 2 | 80 | 79.9798 | 0.00072725 | 130.829 | 0.0002582 | 130.7205 | 0.00016667
|  | 5 | 10 | 10.0037 | 0.00059722 | 103.8355 | 0.00016667 | 103.7908 | 0.0002
|  | 6 | 80 | 79.9964 | 0.00083267 | 130.8383 | 0.00015275 | 130.7356 | 0.0004
Lab: | TUBITAK UME TEMPERATURE GROUP - HUMIDITY LABORATORY | 1 | 10 | 9.9933 | 0.00065064 | 103.8279 | 0.00023333 | 103.7894 | 0.00081921
|  | 2 | 80 | 79.9798 | 0.00072725 | 130.829 | 0.0002582 | 130.7205 | 0.00016667
|  | 5 | 10 | 10.0037 | 0.00059722 | 103.8355 | 0.00016667 | 103.7908 | 0.0002
|  | 6 | 80 | 79.9964 | 0.00083267 | 130.8383 | 0.00015275 | 130.7356 | 0.0004
Lab: | Physikalisch-Technische Bundesanstalt Braunschweig (PTB) | 1 | 10 | 9.9929 | 0.0025 | 103.8265 | 0.008 | 103.7878 | 0.008
|  | 2 | 80 | 80.0444 | 0.0025 | 130.8535 | 0.008 | 130.7575 | 0.008